## Slide 1: MQXFA06 SG First Thermal cycle data (In Process)
D. Cheng, L. Garcia-Fajardo, C. Myers, P. Ferracin
19-Jul-2021
LBNL

## Slide 2: MQXFA SG Gauges Layout
Only HBM gauges are installed.
Shell gauges --- three axial locations:
  Shell gauges (T & Z): on shell 2, 4, & 7
Coil gauges (T & Z): on the axial location of 3989 mm (RE) from LE.
Rod stations are located on the RE (not shown here)
  Each station is full bridge with four individual gauges.
LE
RE
MID
Shell 2
Shell 4
Shell 7
Coil Stations

## Slide 3: Data Processing Notes
Cooldown Data captured from initial cooldown during 9-Jun-2021 to 15-Jun-2021
Warmup Data captured during 2-Jul-2021 to 10-Jul-2021
Cooldown data was processed using offsets the “preload” SG files
  Observed swapped Coil 4 T&Z channels, so swapped back when processing

## Slide 4: Cooldown SG Readings
The stresses are average stress: (MPa)
The Targets set for MQXFA06
To get reasonable TF plot, the coil data pairs with the shell 7 data.
Average Coil stress is on target, the average stress of each shell is higher than target (similar to earlier magnets tested)
 |  | MQXFA06 (Target) | MQXFA06 (Meas.)
RT | Coil | -80±8 | -75
 | Shell | 58±6 | 59 (Shell 7)
CD | Coil | -103 | -68
 | Shell | 110 | 126 (Shell 7)
 | Shell 02 | Shell 04 | Shell 7
R.T. | 54 MPa | 48 MPa | 59 MPa
CD | 125 MPa | 119 MPa | 126 MPa
Measurements spread
Proposed target window
The TF plots of MQXFA06
(Preliminary, still processing
Quench data)

## Slide 5: SG Readings in the Cooldown and Warmup Cycles

## Slide 6: Shell Azimuthal Strain (Cooldown and Warmup)
Cooldown
Warmup
Target: 110MPa @4 K
Target: 110MPa @4 K
Target: 110MPa @4 K

## Slide 7: Shell Axial Strain (Cooldown and Warmup)
Cooldown
Warmup
Event ~@3700 sec
(see also coil axial and rod data)

## Slide 8: Shell Azimuthal Stress (Cooldown and Warmup)
Cooldown
Warmup
Target: 110MPa @4 K
Target: 110MPa @4 K
Target: 110MPa @4 K
Some relaxation (~10 MPa) was observed on return to RT

## Slide 9: Coil Strains (Cooldown and Warmup)
Cooldown
Warmup
Shift/slippage ~@3700 sec

## Slide 10: Coil Stress (Cooldown and Warmup)
Target: -103MPa @4 K
The average coil stress increased after cooldown (opposite expected, as in cooldown as well)
RT preload average is close to pre-test conditions
The average coil stress decreased after cooldown
Still under examination, but quench performance still showed ~nominal preload levels
Cooldown
Warmup
Shift/slippage ~@3700 sec

## Slide 11: Rods (Cooldown and Warmup)
Target: 1800 µε @4 K
Shift/slippage ~@3700 sec
Axial preload returned to ~nominal values at RT

## Slide 12: Additional Cooldown Notes
Some events and/or drift noted after magnet reached 4K
  Plotted zoomed timescale between 5500 min and 5750 minutes
Large signal may not be a real event—no other gauges repond
Rods see a small event
Shells 2 & 7 see drifts
Coils see several shifts then goes back

## Slide 13: SG Readings in Quenches 1 (4.5K) and 13 & 16 (1.9K) (on-going)

## Slide 14: Quench 1 @4.5K (Preliminary, and on-going)
Initial data (still to be examined)

## Slide 15: Quench 13 @1.9K (Preliminary, and on-going)
Initial data (still to be examined)

## Slide 16: Quench 16 @1.9K (Preliminary, and on-going)
Initial data (still to be examined)
Noted by Joe M. that this was a “noisy” quench with both precursors and ringing afterwards

## Slide 17: The Load Condition of MQXFA06 after CD (Preliminary)
  MQXFA06 (Will likely increase, based on Quench Data)
  MQXFA04/05