## Slide 1: MQXFA10 Coil Reception Electrical QC Measurements Coils 131, 132, 220, 221
MQXFA10 Coil Selection Review
Dan Cheng, Bob Memmo, Jordan Taylor
Sep 8, 2021
LBNL

## Slide 2: Overview
QXFA Coil 129 had already been approved for MQXFA09 (see MQXFA05 Coil Selection Review Report in https://indico.fnal.gov/event/49436/ )
LBNL EQC Acceptance Data is available
  Data has been uploaded;
  Summarized in following slides
  Now includes acceptance values reviewed by V. Marinozzi

## Slide 3: Coil Data
All coils are Series, with limited voltage taps
 |  |  | 131 | 132 | 221 | 220 | 129
See DocDB #956 / SU-1010-1903 |  |  |  |  |  |  | 
RLQ |  |  |  |  |  |  | 
 | R | (mΩ) | 607 | 608 | 606 | 606 | 606
 | Ls (mH) | 20 Hz | 10.05 | 10.07 | 10.04 | 10.04 | 10.02
 |  | 100 Hz | 6.42 | 6.42 | 6.42 | 6.42 | 6.41
 |  | 1 kHz | 5.82 | 5.82 | 5.83 | 5.83 | 5.81
 | Q | 20 Hz | 1.28 | 1.27 | 1.29 | 1.28 | 1.29
 |  | 100 Hz | 5.51 | 5.49 | 5.56 | 5.55 | 5.52
 |  | 1 kHz | 8.59 | 8.61 | 8.49 | 8.62 | 8.61
QH R (Ω) |  |  |  |  |  |  | 
 | NT | A01 | N/A | N/A | N/A | N/A | N/A
 | T | A02 | N/A | N/A | N/A | N/A | N/A
 | T-P | B01 | 2.4 | 2.1 | 2.2 | 2.3 | 2.3
 | NT-P | B02 | 2.4 | 2.2 | 2.2 | 2.3 | 2.3
 | T-M | B03 | 2.4 | 2.2 | 2.2 | 2.3 | 2.3
 | NT-M | B04 | 2.4 | 2.2 | 2.2 | 2.3 | 2.3
VT R (mΩ) Individual coils |  |  |  |  |  |  | 
@1 Amp |  | A1 | 0.04 | 0.03 | 0.02 | 0.02 | 0.02
 |  | A2 | 0.09 | 0.09 | 0.07 | 0.07 | 0.07
Note: |  | A3 |  |  |  |  | 
Some coils may |  | A4 |  |  |  |  | 
still have full |  | A5 |  |  |  |  | 
complement of VTs |  | A6 |  |  |  |  | 
but measure only |  | A7 |  |  |  |  | 
A1B1 and A2/B2 |  | A8 |  |  |  |  | 
 |  | B8 |  |  |  |  | 
 |  | B7 |  |  |  |  | 
 |  | B6 |  |  |  |  | 
 |  | B5 |  |  |  |  | 
 |  | B4 |  |  |  |  | 
 |  | B3 |  |  |  |  | 
 |  | B2 | 606 | 608 | 605 | 606 | 606
 |  | B1 | 607 | 608 | 606 | 606 | 606.2
HIPOT |  |  | OK | OK | OK | OK | OK
IMPULSE |  |  | OK | OK | OK | OK | OK

## Slide 4: Hipot and Impulse tests
Hipot tests passed for all coils
  See Coil EQC data workbook on Indico Page
Impulse tests 2000V data shown here
  See data in EQC data workbook

## Slide 5: Summary
All coil acceptance EQC tests were performed
Tests passed with all values within acceptance values
No issues observed

## Slide 6: Appendix