## Slide 2
In-situ Beam on target monitor system at ESS
Presented by Dr. Cyrille Thomas
2022-02-04

## Slide 3
Beam monitoring motivation
    Beam power and high potential of damage
    Assuring Nominal beam, control heat-load, …
Beam detection system overview
    Design overview and material selection
    Measure current density distribution, beam position: tuning and machine protection
Concluding remarks

## Slide 4: Beam Monitoring Motivation
1µs beamlet
2.8ms pulse
ESS beam properties:
Protons: 2 GeV
Peak Current: 62.5mA
Pulse: 2.8ms, 14Hz
Average current: 2.5mA
Lattice:
Beamlet r.m.s size: 13mm x 5mm
Raster system: spread the beam over target area: 160mmx64mm

## Slide 5
Ex. Power on Stainless Steel
Heat equation simple model
Power density
Radiative cooling
Melting condition region: 10-100 W/cm3
Core beam > kW/cm3
Outer beam zone X>100mm, Y> 50mm
  1 W/cm3
e emissivity
Stefan–Boltzmann constant: sB = 5.670… 10-12 W cm-2 K-4
z = V/A
r density
dE/dz stopping power
Ib current density
Cp specific heat capacity
r density
Conduction cooling

## Slide 6
Damage from the beam:
Melting materials
Thermomechanical stress
Fatigue from temperature cycle
But also …
Hardening of material with radiation dose
ESS-0028465 - ESS Materials Guide
Radiation Damage from proton beam and secondary particles:
Damage Types | Description
Neutron induced displacement damage | Change in material properties correlated with displacement damage caused by high neutron fluence
Proton induced displacement damage | Change in material properties correlated with displacement damage caused by high intensity proton beam
Thermal neutron capture | Change in material properties correlated with solid transmutations by thermal (E ≤ 0.625 eV) neutron capture
Fast neutron induced functional degradation | Degradation of material functionality correlated with high fast (E ≥ 0.1 MeV) neutron fluence
Helium embrittlement and swelling | Brittle behaviour and swelling of material due to high helium production rate caused by hadron irradiation
Radiation dose and dose rate | Change in material properties correlated with hadron and gamma induced dose and dose rate
Can be limited by beam control

## Slide 7: Beam on Target Instrumentation Suite
Requirements:
Measure all critical beam properties
  Support fast machine tuning
  Support machine protection
Lifetime as long as possible / or equal to lifetime of machine component supporting them
Complementarity and Redundancy (>95% reliability and availability of the machine)

## Slide 8: Instrumentation for Beam on Target: Overview
Accelerator region
Target region
TW
Current: 62.5mA
Energy: 1-2 GeV
Area = π 3x3 mm2
Peak power: 440 MW/cm2
Area = π 0.5x0.5 mm2
Peak Power: 16 GW/cm2
Area = 160x64 mm2
Peak Power: 1.3 MW/cm2
Raster system
PBIP

## Slide 9: Beam on Target Instrumentation Suite
Complementary diagnostics: permits rapid and trustable verification
Redundancy: permits high availability and reliability
Fast and accurate online measurement supporting tuning, production and protection
GRID PBIP: 1D profile H and V
IMG TW: 2D profile
IMG PBW: 2D profile
APTM PBW: Halo, beam loss
APTM PBIP: Halo, beam loss, scattering

## Slide 10: Instrumentation for Beam on Target: Overview
Beam position
Beam size
Edges positions
Uniformity
Current Density Distribution
Beamlet position
Beamlet size
Beam loss, halo, scattering
Raster pattern
IMG
APTM
GRID
Online systems
Supports:
Beam tuning: set beam properties
Beam operation: detect errant beam condition
BPM
descoped
indirect
direct

## Slide 11
Proton Beam Instrumentation PLUG: PBIP
A Slice per each instrument
Each slice can be installed independently
IMG
IMG
GRID
APTM

## Slide 12: IMG
Luminescent Coating
Optics
Image readout
Image processing
B18 X-ray beam inducing luminescence
Optics assembly in Oslo University Lab
ORCA-Flash4.0 V3
Kintex Ultrascale 40 (KU040)
FPGA algorithm -> FBIS interlock
Image correction (distortion, normalisation, etc.)
Beam properties
Coordinate system (position of fiducials
Peripherals
Illumination
Spectrograph
Filters
Etc.
8MeV proton beam inducing luminescence
TW mockup coating with Chromia Alumina

## Slide 13: Chromia Alumina Luminescent properties
Plasma and flame sprayed by HV (University West)
  Luminescent Yield: 104 photons / MeV
    Sufficient in fresh sample conditions
    Measurement performed at Oslo University
  Temperature studies – ESS BI lab and OCL
  Influence of temperature: reduced yield with increased temperature
    25C -> 230C , yield drops by 70%
    Quenching temperature not measured, expected to be around 300C
  Irradiation studies at DTU
  Impact of radiation on yield
    rapid decay to asymptotic value around 5%
  Luminescence lifetime (ESS BI lab)
    3ms (@ room temperature and non-irradiated sample)
    Not measured at high temperature
    No data or prediction after first irradiation
OCL

## Slide 14
PBW IMG
TW IMG
Beam | value | measured | Error (mm)
Size X (mm) | 36,4 | 36.4 | 0.1
Size Y (mm) | 14.7 | 14.7 | 0.1
Centre of mass x,y (mm) | 0, 0 | 0.07, 0.03 | 0.01
Edge position x left / right (mm) | -59 /59 | -59 / 59 | 0.01
Edge position y top / bottom (mm) | -24/24 | -24/24 | 0.01
Beam | value | measured | Error (mm)
Size X (mm) | 36,4 | 38.9 | 0.5
Size Y (mm) | 14.7 | 14.7 | 0.5
Centre of mass x,y (mm) | 0, 0 | 15.2, -0.03 | 0.05
Edge position x left / right (mm) | -59 /59 | -48.2 / 69.5 | +10.5
Edge position y top / bottom (mm) | -24/24 | -24/24 | 0.05
IMG PBW unique direct verification of the beam on target properties
IMG TW permit partial verification of the beam on target properties but with larger uncertainty in horizontal axis than allocated requirements
Coating:
Fresh
Temperature 40C (PBW) and 230C (TW)
Beam:
62mA
2GeV
2.86ms
No scattering considered on the TW image
Notes: First beam image scale
PBW: 0.5/16bit -> end of life: 5% 0.5 = 0.025/16bits = 1600counts
TW: 0.18/16bit -> end of life: 5% 0.18 = 0.025/16bits = 15counts

## Slide 15
Critical element: Mirror M1
MCNPX power input: 40mW/cm3
Radiation dose -> 20kGy/h
Note: PBW Al6061 T6 takes ~6 DPA in 5000h
ANSYS max temperature: 60C
Thermo-mechanical issues:
Heat-load: melting?
Thermal stress: change in optics properties?
Corrosion:
Estimation of nitric acid production in the target area*:
* ﻿Measurement of radiolytic yield of nitric acid in air, Y. Kanda et al.
Expected Lifetime of Al mirror: 2 years
Solution:
Keep as low pressure as possible
Si protective layer
Deformation not visible, and not measurable under tests
Pressure (mbar) | NHNO3 (moles/5000h) | Thickness (nm)
10-3 | 0.0026 | 3
10-1 | 0.26 | 300
1 | 2.6 | 3000
SNS M1 corroded (Courtesy W. Blokland)
J-PARC PBW corroded (Courtesy S. Meigo)

## Slide 16: IMG: Measurements
Core Beam properties
Peak current density < 1.4mA/cm2
Beam position: ±4mm (H) ±3 (V)
Beamlet profile and r.m.s sizes:
sx x sy >33 mm2
Fiducials
Retro-reflectors array
Distribution of reference points in global ESS coordinate system
  Dimension on the images: position and size of the beam
Position of the TW, and motion
Position: far from the core beam -> no significant heat
SS 316L material, same as TW beam entrance window
Resistant to corrosion, but may need to be protected
TW
Fiducials
Beam edges: ±2mm
TW
M1
FZ Juelich / ESS design

## Slide 17: IMG: Measurements
TW and PBW temperature
Luminescent spectra vs. Temperature
Dual or Single Gaussian fit:
Peak's centre(s) linearly depend on Temperature

## Slide 18
Challenges and possible issues
Uncontrolled / unpredicted scattering or halo:
  Larger heat load on M1, inducing image quality degradation
    Compensated by the fiducials
Fast degradation of the fiducial
Total loss of luminescence, due to contamination
Fast corrosion of M1, in spite of SiO protective layer
3ms laser rastered pulse through full IMG system,
… what can possibly go wrong …

## Slide 19: APTM
Plug slice
Ni Blades: read Proton induced current
TC: read Temperature – proxy for beam current density
Readout electronics
Current: pico4 AMC on Struck board, µ-TCA crate
Temperature readout: Standard ECAT modules control by µ-TCA platform
Fast interlock system to abort beam within 20µs

## Slide 20: APTM model
Current in blade #3 during nominal pulse
Model based in SEY theory
Confirm by experimental verification of APTM prototype (J-PARC collaboration)
  Found SEY larger than predicted
  Effectiveness of bias at -50V
  Correlation between current and temperature readout

## Slide 21: APTM: Critical elements
Blades and TC:
Heat up to 85C in nominal operation
  Possible fatigue from thermal cycles
Ceramic PCB:
Heat up to 45C in nominal operation
  Possible thermal stress at the bolts
Reach >>2000C in 1 non-rastered pulse
Connectors:
Rated for 1GGy: Dose 4kGy/h -> expected 40MGy in 2 years (104 h)
Single shot peak temperature
Steady state peak temperature
Non rastered beam
Thermal damages
Radiation damages
Blades and TC:
Expect He embrittlement and swelling
Ceramic PCB:
Dose: 4kGy / h
Rated for >100MGy at least (CERN-HS-RP-093): lifetime > 2years
Aluminum support
Dose: 4kGy / h << PBW dose
Expected lifetime > 2 years
Nominal beam: steady state regime
Corrosion Resistant materials
99.9% Nickel Blade
Alumina Ceramic
Gold PCB wires
SS304 TC sheath

## Slide 22: GRID
Plug slice
Readout electronics
Current: pico4 AMC on Struck board, µ-TCA crate
Struck with Kintex board as concentrator for FPGA processing
Fast interlock system to abort beam within 20µs
t= 100µs
Sampling 1µs
Beamlet centre and width

## Slide 23: GRID: Critical elements
Thermal damages
Radiation damages
Corrosion Resistant materials
SiC wire
Alumina Ceramic
Gold PCB wires
Macor
SS316L
Ceramic PCB:
Heat up to 75C in nominal operation
  Possible thermal stress at the bolts
Reach >>2000C in 1 non-rastered pulse
SiC wires:
Heat up to 300C in nominal operation
  Possible fatigue from thermal cycles
Heat to 1200C for non-rastered beam
Proven to be long-lived at J-PARC and ISIS
Ceramic PCB:
Dose: 4KGy/h
MCNPX + ANSYS, steady state nominal beam
SiC wires:
Dose: 5DPA / year
He-embrittlement and swelling more important at high temperature

## Slide 24
Challenges and possible issues
Uncontrolled / unpredicted scattering or halo:
  Larger heat load on ceramics and smaller parts, causing mechanical failure
Unpredicted Fast corrosion of components, causing failure in signal acquisition
Electromagnetic noise as seen at J-PARC but too large to prevent reliable measurements
J-Parc pulse on Dump, as seen by APTM Ni Blade.
Strong noise from turbo-pump

## Slide 25: Concluding Remarks
The ESS Beam carries an intrinsic damage potential to the target and target components, hence the beam delivery system and beam properties require a tight control.
The beam properties can be measured by a suite of instruments, which carry complement and redundant beam properties measurements.
The design of each instrument has to bring reliable measurement to permit fast tuning of the beam and machine protection. This remain a challenge due to the harsh environment associated with the high power beam and the proximity of the spallation target.
The lifetime of each instrument depends on the selection of the materials to survive heat deposition and radiation damage, as long as possible or at least for time defined for instance by cost consideration.

## Slide 26
ESS
Beam Diagnostics
Neutronics
Oslo University
Håvard Gjersdal
Erik Adli
University West
Shrikant Joshi
IMP
Haihua Niu
HongMing Xie
J-PARC
Shin Ishiro Meigo
Stony Broke
Sanjay Sampath
SNS
Vim Blockand
ESSB
Special Thanks to all our collaborators
FZ Juliech