## Slide 1: Residual-Resistivity-Ratio Measurement System for Rutherford Cable Extracted Strand
Algirdas Baskys, Ian Pong, Michael Naus, Charlie Sanabria, Jonathan Lee, Scott Myers, Heng Pan, Jordan Taylor, Marcos Turqueti, Li Wang, Xiaorong Wang
Lawrence Berkeley National Laboratory
Special thanks to Tom Lipton

## Slide 2: Outline
Motivation
System: Sample holder, cryostat, measurement
Measurements
Conclusions
This work was supported by the U.S. Department of Energy, Office of Science, Office of High Energy Physics, through the US LHC Accelerator Upgrade Project (AUP) and by the High Luminosity LHC project at CERN.

## Slide 3: Motivation
RRR (Residual Resistivity Ratio) of a cable’s matrix Cu plays an important role in magnet stability at low fields1
  Common Spec: RRR > 100
  RRR = rRoomTemp / rCold
In Rutherford cable, RRR varies along an individual cable strand due to large deformation at cable edges
  Exacerbated if the cable has a non-zero keystone angle (e.g., AUP = 0.4°)
Minor Edge
Major Edge
Cable Cross-Section
Rutherford Cable
[Colored strands indicate strand ID. Used at cable point & tail only]
[1B. Bordini et al., IEEE Trans. on Appl. Supercond., vol. 22, no. 3 June 2012]

## Slide 4: Motivation
Common practice for extracted strand RRR measurements:
  Straighten the extracted strand, leading to inaccurate RRR values
    2nd, extra deformation at edge locations  Lower RRR values
  Do not measure RRR at cable edges (“kinks”) & straight sections separately
System Goals
  Measure extracted strand RRR in as-used condition (i.e., kinks not-straightened)
  Measure RRR of individual strand sections: kinks & straights separately
  Improved repeatability through sample fixturing
  Measure many samples simultaneously and easily
Major Edge
Top Straight
Bottom Straight
Minor Edge
Minor Edge
Major Edge

## Slide 5: System: Sample Holder
Sample holder is printed circuit board (PCB)
PCB design must match cable design
  Current & voltage pin locations conform to cable width & twist pitch
Current & voltage pins are V-slotted
  Accommodates range of wire diameters
  Current taps are soldered
  Voltage taps are friction fit to strand

## Slide 6: System: Sample Holder (cont’d)
Dedicated location for reference sample
  Confirms measurement system health
  Measured with each new sample set
  Stays on PCB
Dedicated locations for straight-sample, heat treat witness wires
  Confirms HT health – Temp & Atmosphere
15 extracted sample strand locations
PCB mounted to 7mm thick Cu plate

## Slide 7: System: Cryostat
Cu plate design allows for stacking several together
Internal flex cable with 100 lines/conductors
  Printed on Kapton substrate for low heat loss
Outer SS vacuum jacket & inner Al radiation shield
Two cooling stages
  Two heaters on 2nd stage for temperature control
Cryomech PT415 cryocooler
  Cooling power: 1.5W @ 4.2K W, 40W @ 45K
  LHe-free operation
Cryocooler
PCB + Cu
Data Cables
SS Shell
Al Shell
Flex Cable

## Slide 8: System: Measurement
System is controlled with LabView software
Thermal EMF mitigated by averaging positive & negative polarity
  Typically, ±0.5 A for AUP wire
  Measurements are compliant to IEC 61788-04
Whole PCB (63 channels) scanned in ~1 s at each polarity
Voltages measured with nanovoltmeter
Temperature ramp rate is controlled through the SC transition
  Two heaters on 2nd stage (50 W max)
  Typically, 0.1 K/min to 20 K

## Slide 9: Example Resistive Traces
For AUP, cold & room temp resistances at SC transition and 20°C, respectively
Python script calculates resistances at cold and room temperatures
   Calculates RRR
Resistance Trace @ Cold
Resistance Trace @ 20°C

## Slide 10: Repeatability: Reference Sample
Reference sample RRR has been quite stable over past 40 cooldowns
  Python script analysis starting at Cycle #55; Manual analysis before.
  Sample had been moved on PCB prior to this, causing a shift in values
New Script Analysis
New PCB

## Slide 11: Repeatability: Same Samples, Different Slots
Samples were measured, then remeasured at a different location on PCB
  % Difference from 1st measure is shown
  Color-coding indicates change direction
  RRR values ~300 ± 100
Differences: Largest ~3% (~10), Median -0.7%
New (2nd) measurements tended to be lower
  Cold work from handling?
Matsushita2 found COV = 1.7% - 4.0%
  Straight samples stayed at PCB location (not moved)
  Round robin sending PCBs between 6 labs
Lower
Higher
RRR Remeasure
2Matsushita et al., IEEE Trans. Appl. Supercon., vol. 28, no. 2, pp. 1-5, March 2018

## Slide 12: Straightened versus Non-Straightened Strand
RRR was slightly lower in straightened strand as compared to non-straightened
  ~5% lower at all wire segment locations
2 cables, 5 strands per cable
  10 total samples compared thus far
  More samples to come
Both strand types of a given cable were heat treated together
  Measured separately due to form factors
RRR values ~300 ± 100
[ (Straightened - (Non-Straight)) / (Non-Straight) ]

## Slide 13: Correlation to Other Measurements
For AUP, 15% rolled (vendor-measured) and Minor Edge RRR (LBNL) are correlated
  15% rolled is more conservative. Vendor rCold at 20-22 K, not SC transition
  An indicator of future Minor Edge RRR in finished cable
Lance Cooley

## Slide 14: Repeatability: AUP + RRR Process
Samples extracted from same cabled strand but on opposite ends, in separate sessions, then HT’d independently & measured independently
RRR values are very consistent despite numerous variables at play
ITER benchmarking found sRRR ~20 (COV ~14%) for internal-Sn strand3
3I.Pong et al., IEEE Trans. Appl. Supercond., vol. 22, no. 3, pp. 4802606, June 2012

## Slide 15: AUP Extracted Strand Trend
Edges show RRR suppression at edges/kinks
  Minor Edge worse than Major Edge
5 extracted strands/cable & 3 cables per measurement

## Slide 16: Straight Sample PCB
45 samples per board
Quickly & efficiently measure many samples
  Heat treat studies
  Wire design studies
  Quality control
Current & voltage pins = screw posts

## Slide 17: Summary
Common RRR methods do not differentiate segments of extracted strands, rather it is ‘whole strand’
  Kinks/Edge locations versus straight sections
A novel extracted strand RRR measurement system has been built to independently & easily measure each strand segment
RRR measurements are internally and externally consistent
Data indicates straightened XS have lower RRR
System facilitates RRR cabling quality control, strand problem investigations, and/or other studies

## Slide 18
Fin

## Slide 19
Back-Up Slides

## Slide 20: Example: Test Facility Dipole (TFD)
Can quickly track RRR impact of cable mfg changes during cable development

## Slide 21: Example: Canted Cosine Theta SubScale (CCTSub)
Investigation of poorly performing strand for Ic showed low RRR
RRREdge ≈ RRRStraight

## Slide 22: Sheared Subelements

## Slide 23: RRR @ Tc vs. @ 20 K

## Slide 24: RRR versus Test Current

## Slide 25: AUP Cable Cross-Section