## Slide 1
Refrigerators and Cryogenic Engineering for Quantum Science
Matthew Hollister
QSC Cryogenic Engineering and Quantum Sensing 2022
September 15, 2022

## Slide 2: Course outline and objectives
Objective: To understand the physics behind the dilution cooling process, be familiar with the design features of modern refrigerators, and to have an introductory understanding of cryogenic engineering considerations at millikelvin temperatures.
Outline:
Physics of helium and the behavior of 3He/4He mixtures
Making a practical refrigerator
The development of dilution refrigerators from the 1960s to today
Modern designs and operation
Materials and design techniques at mK temperatures

## Slide 3: Why use dilution refrigerators?
Dilution refrigerators, based on the heat of mixing of Helium-3 and Helium-4 are the only viable cooling technology that can reach millikelvin temperatures and provide continuous cooling.
Other refrigerators such as adiabatic demagnetization of paramagnetic salts can cool to below 100 mK, but the process is discontinuous. Can be useful for screening of devices or material studies, but less so for device testing.
Helium evaporation fridges can reach 0.3 K (Helium 3) or 0.8 K (Helium 4), which can be useful for sensors such as SNSPDs

## Slide 4
Physics of Helium and Helium Mixtures

## Slide 5: Helium mixtures
Important to note that Helium-4 obeys Bose statistics. Superfluid transition occurs at 2.18 K. Helium-3 obeys Fermi statistics, obeys the Pauli principle and has a superfluid transition near 1 mK.
In the following discussions, the respective concentrations of the two helium isotopes are expressed as
where n3 and n4 are the number of atoms or the number of moles of 3He and 4He, respectively.

## Slide 6: Phase diagram and solubility
FORBIDDEN REGION
Finite solubility – on the 4He-rich side, the concentration does not go to zero at T=0
Diluting 3He into 4He depresses the superfluid transition temperature
Superfluid transition disappears completely for x > 67.5%
As a mixture with x > 6.6% is cooled, it will separate in to 2 phases.
One phase (the “concentrated phase”) will approach x=100%, the other (the “dilute phase”) will approach x = 6.6%.
This finite solubility and phase separation is key to the dilution refrigerator process

## Slide 7: Phase separation at low temperature
FORBIDDEN REGION
3He-rich phase (essentially pure 3He)
Lower density, floats on top of the 4He-rich phase
4He-rich phase
Higher density

## Slide 8: 3He-4He mixtures as Fermi liquids
As mentioned, 4He obeys Bose statistics and will condense into the quantum mechanical ground state at low enough temperature.
In a helium mixture, the condensed 4He acts as an “inert superfluid background” which contributes to the volume of the liquid and the effective mass of the 3He atoms but has zero heat capacity.
The 3He obeys Fermi statistics and will behave in a similar way to conduction electrons in a metal. However, whereas conduction electrons have a Fermi temperature of ~10000 K, the Fermi temperature of the 3He is ~1 K.

## Slide 9: Finite solubility
Let us consider whether a 3He atom at the phase separation line would “prefer” to be in the 3He phase or be in the 4He-rich phase. The 3He will go into the phase with the larger binding energy.
For 3He atoms in pure 3He, the binding energy is the latent heat of evaporation.
For 3He atoms in 4He, although the van der Waals forces between 4He and 3He are identical, the 3He atoms have larger zero-point motion due to their smaller mass.
  The 4He atoms occupy a smaller volume than the 3He atoms, so 3He atoms will feel a stronger attraction to a 4He atom.

## Slide 10: Finite solubility (continued)
As more 3He atoms enter the 4He phase, two effects occur:
  The larger zero-point motion of the 3He means that the liquid near the 3He atoms is more dilute than near 4He atoms, and the 3He atoms will tend to group. Binding energy of the 3He atoms will increase with increasing x3.
  The 3He atoms must obey the Pauli principle, so additional 3He atoms must occupy successively higher energy states. Binding energy of the 3He atoms will decrease with increasing x3.
The first effect is dominant – as x3 increases, the binding energy of 3He in 4He approaches the binding energy of 3He in 3He, and an equilibrium is reached.
This is true even at T = 0, hence the finite solubility effect.

## Slide 11: Cooling power of the dilution refrigerator
Measurements of the specific heat shows that the enthalpy of 3He in the dilute phase is larger than in the concentrated phase. Moving 3He from the concentrated to the dilute phase will result in cooling due to the heat of mixing
We won’t work through the full derivation here, but in summary
This result is correct to a few % for T<40 mK

## Slide 12: Cooling power – things to note
The equation giving the cooling power is for the liquid in the mixing chamber. It takes no account of the Kaptiza resistance between the liquid and the chamber itself, or the contact resistance between the chamber and any external equipment such as a cold plate.
Derivation assumes that the 3He in the concentrated and dilute phases are at the same temperature. In reality, the precooling of the inflowing 3He won’t be perfect, so we modify the derivation to account to TN, the temperature of the 3He leaving the last heat exchanger:

## Slide 13: Osmotic pressure
The osmotic pressure of the Helium mixture is a function of temperature and 3He concentration (see plot right from Betts, 1989).
We will discuss this in more detail in the next section, but if we pump 3He away from the dilute phase the concentration will be reduced, resulting in a difference in the osmotic pressure and driving the flow of 3He.

## Slide 14
Building a Refrigerator

## Slide 15: General layout of a dilution refrigerator
Simplest model for the refrigerator is a U-tube, with the concentrated phase on one side and the dilute side on the other.
As observed previously, the less dense concentrated phase floats on top of the dilute phase since the superfluid has much higher density.
Removing 3He from the dilute phase surface upsets the equilibrium.
From Richardson & Smith (1998)

## Slide 16: General refrigerator layout
Major components of a generalized dilution refrigerator shown in the diagram.
3He from the room temperature pump system is cooled and liquified above (or at) the still.
Cooled further on the way to the mixing chamber.
3He is pumped away from the still.

## Slide 17: Heat exchanger design above 50 mK
At higher temperatures, continuous heat exchangers work well. These are typically of the concentric (or coil-in-tube) configuration.
Heat is transferred from the dilute phase in the annular space and the concentrated phase in the coiled capillary.
Heat transfer along the fluid and the capillary are negligible.

## Slide 18: Heat exchanger design below 50 mK
Designs using continuous heat exchangers only can reach ~30 mK.
For lower temperatures or very high flow rates, larger surface area heat exchangers are needed to overcome the increasing Kaptiza resistance.
Simple form is a sintered metal block to give large surface area, with a machined channel to reduce flow impedance (and hence viscous heating).
To minimize axial heat flow, several discrete exchangers are connected in series – “step” heat exchangers at progressively lower temperature.
General discrete heat exchanger layout (from Pobell, 2007)
General Common design in commercial fridges similar to that shown right (from Pobell (2007), adapted from Frossati (1992)).

## Slide 19: Mixing chamber
The mixing chamber is a relatively simple. Major requirement is for a large contact area between the helium phases and the metal.
This is typically achieved through sintered structures, such as the silver sponge example shown right.
In this design, the diameter of the mixing chamber is approximately 100 mm, while the sponges have a surface area of tens of square meters.
Mixing chamber construction example from Uhlig (2015)

## Slide 20: Room temperature components - Pumps
Choice of pumps depend on circulation rate and desired base temperature.
Usually consist of a primary pump (scroll or roots).
Still pumping may be supplemented with one or more turbomolecular pumps for higher-performance systems.
For efficient condensation, it is often necessary to use a compressor to increase the inlet pressure (and decrease the outlet pressure of the primary pump).
Important that pumps are sealed for helium use and do not contaminate the process line.

## Slide 21: Room temperature components – Filtering and sealing
Majority of the gas handling at room temperature will be subatmospheric, so will tend to suck in air that could contaminate the process line.
  Fridges generally include cold filters with Pt or other metal sinters internally
  Many will also include an external activated charcoal trap in liquid nitrogen
  For very long duration runs, an LHe or 4-K trap can also be used.
Choice of flanges and seals is important. Metal seals generally much better than rubber, but harder to work with.
  Be wary of parts of the system that will see internal pressure. If using rubber O-rings, it may be necessary to add overpressure rings.

## Slide 22
History and Development of the Dilution Refrigerator

## Slide 23: Early development
The idea of using 3He/4He mixtures as a refrigeration method was first proposed by London in 1951, and further developed in 1962 by London, Clarke and Mendoza.
First demonstrations:
    Das, de Bruyn & Taconis (1965) reached 220 mK
    Hall, Ford and Thompson (1966) reached 80 mK
    Neganov, Borisov & Liburg (1966) reach 25 mK
The refrigerator of Hall et al. is shown on the right

## Slide 24: Early commercialization
The first commercial refrigerators were built in the mid-1960s by Oxford instruments, first in collaboration with Heinz London, and later based on Hall’s design licensed from the UK Atomic Energy Authority.
(Left) Early commercial refrigerator insert
(Right) Commercial refrigerator with copper heat exchangers, c.1975

## Slide 25: Eliminating the liquid helium
The primary feature that distinguishes modern refrigerators from the older style is the use of mechanical precoolers in place of the liquid helium bath.
This was made possible by an earlier innovation, the Joule-Thomson Heat Exchanger.
Described independently by Kraus at ZTB and by de Waele at Eindhoven in 1977, this device eliminates the 1-K Pot in a wet dilution refrigerator by using a long counterflow exchanger cool the Helium-3 stream with gas pumped from the still, followed by an expansion valve.
  This stage cools and partially liquifies the Helium-3, with the remaining liquification taken care of by the still.

## Slide 26: Eliminating the liquid helium (continued)
Removing the 1-K Pot served three important purposes:
    The pumped pot can be a significant source of vibration in sensitive experiments
    The pumped pot can dominate the Helium consumption of a system
    The pot is prone to blockage, making it a common source of fridge failure
Note that it would not be ideal to just cool and liquify the Helium-3 by running the gas directly to the still since the enthalpy of the gas is larger than that of the gas leaving the still at ~1K.
  This has been demonstrated, but only for very low flow rate fridges with efficient precooling, so not a very practical approach.

## Slide 27: Eliminating the liquid helium (continued)
With improvements in mechanical cryocoolers in the late 1980s, it was a natural step to couple one to a dilution refrigerator.
First example described by Pari in 1990, followed by the pioneering development work of Uhlig starting in 1993.
Pari’s machine (shown right) used a Stirling cycle cooler operating at 12 K combined with a Helium-4 J-T to produce a temperature at ~4 K.

## Slide 28: Eliminating the liquid helium (continued)
The development of Uhlig’s group eliminated the Helium-4 J-T stage, relying instead on the Joule Thompson Heat Exchangers used previously in wet refrigerators.
These fridges (example shown right) initially used 10 K Gifford-McMahon coolers, and later small 4 K pulse tube coolers.
Important to note that 4-K class cryocoolers were not available until 1999-2000. Much of the early work was carried out with 10-K coolers.

## Slide 29: “Dry” vs “Wet” Refrigerators
Cooled by Pulse Tube to ~50 K
Cooled by Pulse Tube to ~4 K
Cooled and partially liquified by JT
Cooled and liquified by Still
Mixing Chamber
Cooled by Liquid Helium to ~4 K
Cooled and liquified by 1-K Pot
Cooled by Still
Mixing Chamber
“Dry” Process
“Wet” Process

## Slide 30: Early commercial dry fridges
Early examples of commercial cryogen-free fridges came from Leiden Cryogenics and Air Liquide in the early 2000s
The Leiden refrigerator, built for the SCUBA-2 astronomical camera (right) is still operational at 14000 feet at the James Clerk Maxwell Telescope in Hawaii.
Limitations of these early designs were often poor heat exchange at 4 K on the second stage of the pulse tube cooler.

## Slide 31: Improvements to the precooler design
An interesting property of the pulse tube cooler was measured in the early 2000s in that excess cooling capacity on the regenerator between the two stages exists.
Some heat can be dissipated on the regenerator with little impact on the 4 K cooling performance.
Adding a coil to the regenerator provides a useful additional heat sink for the inflowing 3He, reducing the cooling load at the 4-K stage.
Additional heat exchanger (Uhlig, personal communication)

## Slide 32
Modern Dilution Refrigerator Design and Operation

## Slide 33: Modern refrigerator layout
Image: Oxford Instruments
2 stage pulse tube cooler
Additional heat exchanger

## Slide 34: Modern refrigerator layout – temperature stages and dilution unit
Images: Oxford Instruments
50 K (First stage pulse tube)
4 K (Second stage pulse tube)
1 K (Still)
0.1 K (Intermediate cold plate)
0.01 K (Mixing chamber)
Still
ICP
Mixing
Chamber

## Slide 35: Commercial fridge examples
Small fridges
Mixing chamber diameter ~150mm
Minimum temperature ~25 mK
Image: Oxford Instruments
Image: Oxford Instruments
Standard fridges
Mixing chamber diameter ~300mm
Minimum temperature ~10 mK

## Slide 36: Commercial fridge examples (continued)
Large frame fridges
Mixing chamber diameter ~500mm
Minimum temperature ~10 mK
Image: BlueFors
Image: Leiden Cryogenics
“XXL” fridges
Mixing chamber diameter ~1000mm
Minimum temperature ~10 mK

## Slide 37: Operational Considerations
Some general operational comments:
  Monitor the system behavior run-to-run. Differences in temperatures, pressures and performance can be useful indicators of problems.
  Monitor how much mixture you are recovering at the end of the cool down, useful to monitor for leaks.
  Be aware of how much contamination is accumulating in the cold traps between runs since this can indicate the presence of an air leak into the system.
  Be aware of recommended maintenance internals on pumps. Can be a source of performance issues or leaks.

## Slide 38
Materials at Low Temperature

## Slide 39: Thermal Conductivity
Increasing purity
Thermal conductivity includes contributions from electrons (in metals) and phonons (in metals and insulators).
As temperature decreases, the phonons will freeze out, strongly reducing the conductivity in insulators and leaving the electrons as the dominant conduction mechanism in metals.
In superconducting metals, conductions falls off below Tc since Cooper pairs have poor energy transport properties.
Pobell (2007)
Thermal conductivity in copper

## Slide 40: Thermal Isolation
Assuming a notional support with dimensions chosen to give the same stiffness/fundamental frequency (so we are comparing like-with-like), we can compare the conducted heat for different materials.
CFRP is an extremely amorphous material, so lots of scattering
Ti 15-3-3-3 is superconducting, with a Tc ~ 3 K
Material | Conducted Heat (1 K to 0.1 K)
G-10 CR | 11.5 mW
Stainless Steel 304 | 12.0 mW
Carbon Fiber Reinforced Polymer | 3.5 mW
Titanium 15-3-3-3 | 2.8 mW

## Slide 41: Contact Conductance
At interfaces, physical contact exists over microscopic asperities.
Area available for heat transfer is a tiny fraction of the macroscopic contact area. This can be maximized by using very low surface roughness (i.e. polished) surfaces.
Many of the “tricks” to improve conductance at high temperature don’t work well below 1 K.
From Thermal Contact Conductance (2014) C V Madhusudhana

## Slide 42: Permanent Joining Techniques
Less common to join components permanently than with a demountable joint, but still worth considering in some circumstances.
Permanent joining techniques include:
  Soldering/brazing (basically the same process at different temperatures), using a filler (solder or braze alloy). Be wary of superconductors in fillers.
  Welding, which melts the base metal.
Electron-beam welding is extremely effective at low temperatures, especially for copper. Can also be used to join dissimilar metals. E-Beam welding does not use a filler rod, unlike GTAW.

## Slide 43: Improving Contact Conductance
For a fixed joint pressure, thermal resistance at a joint can be lowered using an interposer.
Need to be wary that interposers such as N grease or indium are less effective below 1 K (and can degrade the joint) but can be useful in very low-pressure joints.
Gold (silver, or Pt) plating is very effective.
Surface treatments that prevent oxidation are also effective.
Au - Au
In-Cu
Cu - Cu
Adapted from Eiken (2006)

## Slide 44
Cryogenic Thermometry

## Slide 45: Definition, Primary and Secondary Thermometers
A thermometer is a device by which we can measure a property of matter which is connected via a physical law to the concept of “temperature”.
A primary thermometer is a device where the theoretical knowledge about the measured property is good enough for us to determine the temperature without a calibration (e.g. a gas pressure thermometer or thermal noise in an electrical resistor).
A secondary thermometer must be calibrated against a primary thermometry or fixed points, however since these devices are considerably easier to use, they are workhorse devices.

## Slide 46: Resistance thermometers
By far the most common temperature sensor you will encounter - any device that shows a change in resistance with temperature
Examples of common cryogenic resistance thermometers:
Element Material | Lower Temperature Limit (K) | Upper Temperature Limit (K) | Interchangable?
Ruthenium Oxide | 0.01 | ~40 | Yes, above 0.05 K
Zirconium-oxynitride (Cernox) | 0.1 | 420 | No, individual calibration
Germanium | 0.05 | 100 | No, individual calibration
Carbon-Glass | ~1 | 500 | No, individual calibration
Platinum | 14 | 900 | Yes
Rhondium-Iron | ~1 | 500 | No, individual calibration

## Slide 47: Resistance thermometers
Positive temperature coefficient RTDs
Negative temperature coefficient RTDs
Adapted from Eiken (2006)

## Slide 48: Considerations in temperature sensor selection
Temperature range
Is a full range sensor required, or can multiple sensors be used to cover the range (e.g. a Pt sensor to monitor the cool down and a RuOx for base temperature)?
Required accuracy (at the temperature of interest)
Required resolution (at the temperature of interest)
Available space and mounting scheme
Presence of magnetic field or ionizing radiation

## Slide 49: Measuring RTDs
Crudely, this measurement can be made with a digital multimeter. However, accuracy for small voltages is poor and the excitation current of the measurement is high
More common to use a balanced bridge, similar to the arrangement right.
Separates the voltage measurement from the bias current.
Adapted from Eiken (2006)

## Slide 50: Two-wire vs four-wire measurements
Simple measurement can be made with one wire in and one wire out.
The measurement will include the lead resistance since there is some loss of current on the leads – this will affect the measurement (lead resistance could be equal to the thermometer resistance in extreme cases).
Better accuracy is achieved by separating the I and the V leads in a four-wire configuration – no current on the V leads, so no resistance.
Downside is twice as many leads, so twice the conductive heat leak to the experiment (easily addressed using low conductivity material for wires).
Lake Shore Cryotronics

## Slide 51: Self-Heating in RTDs
In a resistor, P = I2R
For NTC devices, as R increases, I must decrease to reduce power dissipation in the sensor
Power dissipation will raise the temperature of the device relative to the environment
Also highlights the need to make sure the sensor is well connected to the sample
Credit: Marek Bartowiak, Paul Scherrer Institut

## Slide 52
Summary, Concluding Remarks, References

## Slide 53: Summary
We have reviewed:
  The physics or Helium-3/Helium-4 mixtures and the dilution process
  Practical aspects of constructing a refrigerator and the design of key components
  The history of development of modern, cryogen-free refrigerators
  Current state-of-the-art
  Materials and engineering approaches for low temperature experiments
  Cryogenic thermometry
This of course was a very quick tour through a complex set of topics – if anyone has detailed questions, please feel free to talk during the break, or you can contact me at mhollist@fnal.gov

## Slide 54: Useful References
“Matter and Methods at Low Temperatures”, Frank Pobell – Very good general textbook for low temperature physics
“Experimental Principles and Methods Below 1K”, O.V. Lounasmaa – Good description of the dilution refrigeration process, although was published in 1974 so is out of date on the technology
“Experimental Techniques for Low-Temperature Measurements”, Jack Ekin – Not so much about the cooling techniques themselves, but has a wealth of cryogenic reference data

## Slide 55: Acknowledgement
This manuscript has been authored by Fermi Research Alliance, LLC under Contract No. DE-AC02-07CH11359 with the U.S. Department of Energy, Office of Science, Office of High Energy Physics.