## Slide 1: DC-DC Topologies
Michael Miller
Jane Nachtman
Yasar Onel
The University of Iowa

## Slide 2: Cockroft Voltage Multiplier (LBL)
4X
Tested warm and cold (77K) no change of components or adjustments needed;
Regulated;
Capable of delivering up to 1mA;
Utilizing opamp proven to work 10 years in LAr (GERDA experiment)
Can perform with input voltage as low as 4.6V and as high as 6V
Top with shield
Top no shield
Bottom

## Slide 3: LBL Card at CERN

## Slide 4: Status of the Berkeley Circuit
Noise at Iowa measured to be 500 uV pk-pk
Noise at CERN measured at several mV
Have not drawn a conclusion why noise was higher at CERN
20 Boards are currently being produced at Berkeley

## Slide 5: Self-Excitation Isolated Transformer Circuit (Pico 5A48)

## Slide 6: Filter Circuit for the Pico 5A48S

## Slide 7: Unshielded Pico Card at CERN

## Slide 8: Test Results for the 5A48S
Noise at Iowa measured ~ 250 uV pk-pk without shielding
Shielding reduced noise to less than 200 uV pk-pk
Pico devices did not generate observable noise at CERN
Thermal start-ups did not reveal any anomalies
Awaiting 30 year approval of 5A48S component

## Slide 9: Berkeley and Pico Board
 | Berkeley Card | Pico Card
Bias Tuning | Adjustable by zener diode | Adjustable by input voltage, input resistor or zener diode
Circuit Complexity | Moderate to complex | Very simple – Active Components 2 transistors and 2 diodes
Noise Testing | Good result (500 uV at Iowa); problem at CERN | Better result at Iowa (< 200 uV at Iowa); Good result at CERN
Current Construction | 20 boards being assembled at Berkeley | 20 boards to be completed by December 1
Purchasing | Cost: ~ $32 ea. in quantity | Cost: ~ $85 ea. in quantity