## Slide 1
Laserwire Review Summary
Vic Scarpine
PIP-II Beam Instrumentation Laser Wire Design Follow-Up Review
Nov. 7, 2022

## Slide 2: Review Charge
The reviewers are asked to perform an ‘mini review’ of the laser wire profile monitor system in the context of a recommendation from the Instrumentation Comprehensive Preliminary Design Review.
This review did not address the laser wire design that didn’t change from the PDR.

## Slide 3: PDR Recommendation #1
The need to resolve issues with the noisy signal seen in the PIP2IT laser wire (lower dynamic range due to lower laser power).
The Introduction talk showed beam profile measurements from the fiber-based laser profile monitor at PIP2IT
  This profile monitor could measure core transverse and longitudinal beam profiles but had low signal-to-noise
Laser Wire Performance talk showed the simulation tools developed to understand signal-to-noise for different laserwire designs.
Performance studies shows benefits to using small number of large energy laser pulses
  Reduced integration time leads to lower self stripping background
  Lower gain needed and thus lower EMI noise
  The free-space transport laserwire system permits larger laser energy pulses
Charge #5

## Slide 4: PDR Recommendation #2
Do a performance and cost comparison between fiber transport line and the free space transport line. This could help make the decision between the two systems.
The Laserwire Cost Comparison talk gives a good estimate of the all-fiber laserwire versus the free-space transport laserwire
  Based on initial estimates the free-space system:
    Labor
      Adds 650h of installation
    M&S
      Initial estimates show a $178K savings
    Schedule
      The team has confidence that we can maintain the same schedule.
The cost impact is break even within uncertainty of design maturity.
  Allows laserwire selection to be based on performance
Charge #5

## Slide 5: Free-space Transport Laserwire
Optical Transport Line (OTL)
The free-space laser option requires a low vacuum OTL that is optically straight for its entire length.
The preliminary design of the OTL shows it can meet the design requirements.
Laser Design
Free-space optical transport decouples laser system and makes it easier for future laser upgrades
A newly designed laser system is more versatile
The new design will allow for higher signal-noise than old fiber-based design
Charge #5

## Slide 6: Free-space Transport Laserwire
Free-Space Alignment
We have a preliminary design for the transport optics
  Includes control of the laser and potential issues with power on vacuum windows
We have a procedure for alignment of laser transport and beamline optics
  Transport alignment is coincident with installation of transport line
Safety and Risk
Laserwire risks have been identified.
  Prevention through design to mitigate risks
Laserwire safety
  Interlock system to control uncontrolled release of laser
  Working with ESH and laser safety officer in design and operation of laserwire system
Charge #5

## Slide 7: Conclusion
A comparison of performance and cost of an all-fiber laser profile monitor system versus a free-space transport laser profile monitor system has been evaluated.
We conclude that the decoupling of the choice of laser from the transport of laser light into the PIP-II tunnel plus the ability to increase the peak laser pulse power with a shorter signal integration time strongly favors the selection of the free-space laserwire option. This option also allows for the greatest flexibility for the future.
Continued work going forward will include:
  Further studies to optimize laser system to improve signal-to-noise
  Lab testing of laser induced damage to optical vacuum windows
  Lab testing of laser point stability with feedback system
We ask for the committee's recommendation to select the free-space laserwire option and to continue toward the final design of the PIP-II profile monitoring system.
Charge #6