## Slide 1: Module post-assembly reception QC
D. Warner
11/15/22

## Slide 2: Module receipt and Initial Inspection
Modules assembled at UNICAMP. Shipped to CSU. No filter plates(?)
Initial receipt inspection to confirm safe transit
  Visual inspection of module with check-list (WLS plates, screws tight, etc.)
  Check of module fit into mechanical model (ensure fit into APS rails)
  Initial module scan in warm scanner
    DAPHNE readout
    Check against UNICAMP exit scan

## Slide 3: Module cryogenic test
Installation into large cryostat test stand (5-10 modules per test, to be confirmed)
Flood containment area with dry N2 gas, wait 15 minutes purging
Lower into LN2 bath (2 hour immersion time, 2cm/minute)
Tests:
  DCR test each supercell
  LED flasher test (367nm), check signal size at 3V OV (?)
Remove from LN2 bath into dry N2 gas. (2 hour extraction time, 2cm/minute)
Warm in dry LN2 volume (1 hour warm-up)
Remove from test stand

## Slide 4: Post cold test checkout, packing, shipping
Re-scan module in warm scanner-- look for changes
Final exit inspection (re-check visually with checklist. Tighten fasteners)
Install Dichroic filters (?)
Final pre-packing confirmation of all QC documentation
  Stop process if not complete
Pack module following procedure
Indicate QC data ready for SURF inspection. Await acceptance before shipping?
Ship to SURF

## Slide 5: Procedures required:
Module visual/reception inspection (check points, screw torque…) with traveler (Store in DB)
Warm scanner procedure with traveler (Store in DB)
Cryogenic test procedure with traveler (Store in DB)
Dichroic filter installation procedure with Traveler (Store in DB)
Shipping procedure/checklist with traveler (Store in DB)