## Slide 1: Test beam at Notre Dame Radiation Lab
Dan Levin, Yuxiang Guo
for UM Calvision team

## Slide 2
Beam line
electron source - injector

## Slide 3
Control room
pulse measurement

## Slide 4
Generates 8 MeV electrons
pulsed beam with injector: width = 1 ns to 1.6 us
Pulsed beam with injector : repetition rate. Was 30 Hz (normally can be 60 Hz)
We ramped down to very short ~ 1 ns (FWHM) pulse of many particles.
max current > 1 amp. We lowered the current to lowest measurable using beam transformer and Faraday Cup/
minimum measurable charge per pulse
   0.09 nC using in-line inductive pickup
defocus to achieve arbitrarily low rates.
  Cannot directly measure with in-house instrumentation !
  Gauge pulse intensity using the known average pulse height from CR muons in PWO
1 ns
rep rate = 30 Hz
What kind of pulse ?

## Slide 5
1 –1600 ns
rep rate = 30 Hz
5 mm bore
low particle count per pulse
We provided our own beam collimator: ¾ inch thick steel with 4.9 mm aperture

## Slide 6
Estimate the charge /pulse using PWO– normalized to CR muons as follows:
Average pulse amplitude = 100 mV from CR muon
Average energy loss in PWO for 8 MeV electrons dE/dX = 2.1 MeV cm2/g (from estar NIST database)
PWO density = 8.28 , approx. pathlength (avg =2.5 cm)
==> total muon energy loss = 44 MeV ==> this is equal to about 5 electrons
During beam set up we defocused to get abut 400 mV signal
   about 50 electrons/pulse.

## Slide 7
Rotatable platform
Use worm-gear drive, remote control stepper motor
beam pipe
collimator
Test crystal
cradle
SiPM readout
board
SiPM readout
board

## Slide 8
e11
s12
s11
s22
s21
e12
e21
e22
SiPM placement
25 mm
60 mm

## Slide 9
PMT
4 SiPMs
area of each = 6mm x6 mm
Beam angles: -60 < 𝜃 < 30
Beam max radius ~ 1 cm
𝜃
4 SiPMs
area of each = 6mm x6 mm
Signals available: total 9 channels (8 SiPMs, + 1 PMT)
DAQ channels = 4
All readout conducted with S11, S12 and E11, E12. (two side and two end rows)

## Slide 13
Test samples:
Chemical | Density g/cm3 | Index refraction | dE/dX 8 MeV ele. (range cm) | Cerenkov |  | Scintillation |  | 
 |  |  |  | Threshold [MeV] | Yield ph/MeV | Peak 𝜆 | Yield ph/MeV | Decay time [ns]
PbWO4 (PWO) | 8.28 | 2.16 | 2.1 (0.3 ) | 0.066 | tbd | 420 | 200 | 6-20
Bi4Ge3O12 (BGO) | 7.13 | 2.15 | 2.03 (.72) | 0.066 | 20 | 480 | 8K-10K | 300
PBF2 | 7.8 | 1.8 |  | 0.103 |  |  | ~500 | 
Optical glass Silica (59%) + PbO ? unknown | 2.2 | 1.5 | 2.05 (2.3) | 0.190 | tbd | N/A | N/A | N/A

## Slide 14: Experimental considerations:
Look for Cerenkov signal features:
in optical glass block as reference without scintillation
evaluate with SiPMs
  with/without the U330 notch filter
  With/without long wavelength pass filter
  Evaluate for BGO, PWO, Optical glass
Evaluate usefulness of NDRL for future experiments

## Slide 15
DAQ/readout
  Photodetectors are the 4 channel SiPM board.
  Use the fast Tektronix (1 GHz BW) digital scope
  Readout rate capability = 45 hz, four channels at 2000 points/scan

## Slide 16: Program
Dataset: 1 minute runs at 45 Hz: ~2K events/run
  2 samples: BGO, PWO, Optical Glass)
  each sample: data at. 60, 40, 20, 0 , -20 degrees
  each sample with and without filters on end SiPMs only.
Trigger:
  Option 1: two side SiPms in coincidence No filter.
Readout :
  two side SiPMs and two End SiPMs (same row).
  Many other possible combinations using linear analogue fan in- but we did not pursue this.

## Slide 17
Preliminary results
Total of 49 data runs
Evaluated PWO, BGO, Glass
Tested with and without long pass PWO and B˝O and notch filter on BGO only
Most runs were ~ 1 minute, about 500 events each
Conducted one long run of 15 minutes to check beam stability.
Data consists of full waveform in the DSO readout window
Beam diameter measured using a beam monitor imager.

## Slide 18: U330 notch filter
18
U330

## Slide 19
3.2 nC/pulse
FWHM = 8.7 mm
Image plane is at 9 cm from 5 mm aperture
Crystal surface is 20 cm from aperture
FWHM at crystal surface ~ 1.4 cm
beam used for crystal test ~0.3 nC/pulse
Image from the prototype Fast Scintillator Beam Monitor

## Slide 20
5 mm dia
r = 2.5 mm
FWHM @ 9 cm = 2(h+r) =9 mm  h= 2mm
H=h(X/x) = 4.4 mm
FWHM @ 20 cm = 2*(H+r) ~ 1.4 cm
h
H
x = 9 cm
X = 20 cm

## Slide 21
Data prep:
49 runs in hdf5 binary Digital Sampling Oscilloscope format
readout into temporary text files using python
all data files from all runs merged into a single root tree (with thanks to Claudio )
File is NDRL.root 12 GB
Analysis conducted on the root tree
Analysis considerations:
Due to:
Variations in beam over time
Beam intensity changes due to incidence angle
Large diameter of beam at crystal surface also introduces systematic change WRT angle
 develop a metric that is insensitive to these factors

## Slide 22
PWO
end
Average of 500 waveforms:
no filter
0 degree incidence angle
PWO
side

## Slide 23
BGO
end
Average of 500 waveforms:
no filter
0 degree incidence angle
BGO
side

## Slide 24
glass
end
Average of 500 waveforms:
no filter
0 degree incidence angle
glass
side
Use glass sample for Cerenkov only signal
Side vs End readout channels similar
  - same FWHM = 7 ns
Cerenkov signal is delta function:
  - FWHM and decay is feature of the SiPMs.

## Slide 25
BGO
Side
No filter
BGO
Side
UV330
400-600 nm notch filter
Removes most scintillation
FWHM= 40 ns
FWHM= 15 ns
Reduced signal

## Slide 26
Analysis method:
Define a metric that is insensitive to beam conditions:
Take ratio of the presumed
Cerenkov to scintillation components of the waveform:
i.e the ratio of leading integral to the trailing integral
Use a fixed limit of integration set by the glass sample
Example: Ilow = 1% above baseline to 1.5 ns
Ihigh = 1.5 ns to 100 ns

## Slide 27
PWO no filter
Towards SiPM end of crystal 

## Slide 30
Preliminary Conclusions
Conducted first beam test with at NDRL
There samples tested – glass, PWO, BGO
Developing a beam current insensitive metric to gauge Cerenkov vs scintillation components
Hints of angular effect from Cerenkov signal in scintillator, none in glass
UV filter affect pulse width – removes most scintillation light.
UV filter also removes most of the signal. (slide 20)