## Slide 1
GM and GM-JT cryocoolers (by Tomohiro Yamada)
300mm
550mm
1040mm
350mm
GM: RDE-418D4 (SHI)
GM-JT (SHI)
Above 100K
The GM doesn’t have thermal breakdown, whereas the GM-JT will be thermally broken above allowable heat input (9-10W).
The GM-JT takes long cooling time because of its poor cooling capacity in high temperature region.
GM: Less than 1 hour
GM-JT: Approx. 13 hours
Thermal breakdown
2023.12.07 Hot Topic session: Choice of cryocoolers, T. Yamada (KEK)
Cooling time w/o cold mass

## Slide 2
Cavity cooling (1.3GHz)
Max Min
2.735K 2.664K
Δ=70mK
100 mK
10 mK
Max Min
4.361K 4.359K
Δ=2mK
Cold heads’ temperature oscillation
GM: 12 h
(cavity: 2.7 K)
GM-JT: 63 h
(cavity: 4.3 K)
GM-JT takes almost 5 times long time to cool down the cavity.
Minimum temperatures for cases of GM and GM-JT were 2.7 K and 4.3 K, respectively.
GM-JT temperatures were quite stable.
We saw several hundreds Hz of frequency fluctuation in the GM case. <- Due to vibration or temperature?
Considering thermal resistance in the thermal link and RF heating at the cavity, the cold head temperature needs to be as low as possible to keep the cavity temperature near 4.2K.
2023.12.07 Hot Topic session: Choice of cryocoolers, T. Yamada (KEK)
Cooling time of 1.3 GHz single cell cavity