## Slide 1: Fabrication of 650 MHz Deflecting Cavity
Zhongyuan Yao
WG-2, TTC at FNAL
Dec. 5 – 8, 2023

## Slide 2
ARIEL e-Linac
100 kW CW electron beam for RIBs production in ARIEL
300 keV e-Gun
Injector and Accelerator CMs
  1.3 GHz, 9-cell, 10 MeV per cavity
to ARIEL

## Slide 3
e-Linac Upgrade

## Slide 4
Beams in ERL Mode
650MHz RF separator
ARIEL beam
Recirculating beam
Deacceleration beam

## Slide 5
650MHz Deflecting Cavity
3mm beam separation of ARIEL and recirculating beams in 3m
RF separator requirements
  Initiate separation
    ±4.4 mrad angular deflection
    0.3 MV for 50 MeV beam
  Peak fields not demanding
  Limited longitudinal space
  High shunt impedance required
3m
Beam Separator Scheme

## Slide 6
650MHz Deflecting Cavity
Baseline geometry from RF dipole design
Optimized for high R⊥/Q and short cavity length
 | Baseline | Post & Ridge
Frequency | 650 MHz | 650 MHz
Dimension | Φ 210 × 330 mm | Φ 204 × 175 mm
Nominal V⊥ | 0.3 MV | 0.3 MV
R⊥/Q | 411 Ω | 625 Ω
G | 117 Ω | 99 Ω
Epeak @ 0.3 MV | 4.5 MV/m | 9.5 MV/m
Bpeak @ 0.3 MV | 7.8 mT | 12 mT
P0 @ 4 K | < 1 W | < 1 W
73mm
50mm

## Slide 7
Low-Cost Fabrication
Reactor grade Nb (RRR-45)
Machined from solid block
  EDM to cutout beam pipes and RF ports
Copper cavity to prove machining

## Slide 8
Tungsten Inert Gas (TIG) Welding
TIG welding developed
Sample welded and RRR measured to qualify the process

## Slide 9
Completed Cavity
First SRF cavity fabricated in TRIUMF machine shop
1
2
3

## Slide 10
Preparations for Cold Test
aaa
bbb
  ccc
  ddd
BCP…
HPR…
Assembling…
Ready to go!
Into the cryostat...

## Slide 11
RF Surface Finish
BCP on reactor grade Niobium led to uneven surface finish
Smoother areas appear darker
Cracks with a series of pits on ridges

## Slide 12
Cold Test Results
Surpassed design specifications
  Peak fields 25 MV/m and 32 mT
  Cavity quench at 0.8 MV
  No measurable X-ray
TIG weld didn’t cause issues
No Q-disease from 100 K soak
Mechanical properties
  df/dp = +9.7 Hz/mbar
  LFD = -101 Hz/MV2

## Slide 13
Surface Resistance
RBCS and Rres extracted from cooldown data from 4 K to 2 K

## Slide 14
Summary
650 MHz deflecting cavity was built in-house with novel (low-cost) fabrication technologies
  Reactor grade niobium
  Machined from solid block
  TIG welding
Cavity surpassed design specifications for voltage and quality factor
  Cavity reached 32mT before quench
  No obvious effect from low temperature bake
  No Q-disease without hydrogen degassing
  High residual resistance and strong Q-slope caused by poor surface finish after BCP

## Slide 15: Acknowledgements
Many thanks to
  Entire SRF group
  Machine shop
  CMMS
  Cryogenics group
  Design office, Targets group, Detector group, and more
  Chris Compton (FRIB)

## Slide 16: Thank you Merci Question?
www.triumf.ca
Follow us @TRIUMFLab

## Slide 17
Thermal Effects
Solid ridge and RG Nb lead to ~100× thermal resistance from RF surface to helium interface comparing to typical SRF cavities
Measurement
  Suddenly increase VT
  Observed > 0.7 MV
  < 1.5% Q decrease
Reverse meas.
  Reduce to 0.2 MV
  ~ 4% Q increase

## Slide 18
Multipacting
A few MP barriers observed at 4K
Processed in a few minutes w/ < 10 W power
  Not re-appear after processing
Trajectory at 0.5 MV

## Slide 19
Solid Ridge
df/dp
  4mm shell – 80 Hz/mbar
  Solid ridge + 3mm shell – 1 Hz/mbar
    Test condition 13 Hz/mbar (9.7 Hz/mbar)

## Slide 20
Ti Flange to Nb Tube
EBW

## Slide 21
Cavity w/ RF Ancillaries