## Slide 1: Heat Load Measurement for 1.3 GHz CM at IHEP
Feisi He, Mei Li (IHEP, Beijing, China)
Presented by Han Li (IASF, Shenzhen, China)
TTC Meeting, FNAL, Dec. 6, 2023

## Slide 2: World’s First Mid-T 1.3 GHz Cryomodule by IHEP
For Dalian Advanced Light Source (DALS) R&D.
Refer to Jiyuan Zhai’s talk on Tuesday WG1: Performance of First Mid-T 1.3 GHz Module

## Slide 3: Module Cryogenic Control Interface
Endcap pre-cooling valve
Mass flow meter
J-T valve
Cooldown valve
Heater in liquid level tank
Heater in liquid level tank

## Slide 4: Heat Load Measurement: Uncertainty of Flow Rate
The flow rate reading from mass flow meter at room temperature have an uncertainty of , typically
The flow rate data is obtained during a period of time. The standard deviation of the data is the random error (typically with > 600 data in 5 mins)
The flow rate at static state may vary after several hours, which introduce a systematic error, and the uncertainty can be estimated by standard deviation of measured values. By measuring the flow rate at known heater power before and after Q-measurement, the could be calibrated.
Total uncertainty of flow rate is
Date of measurement | Static flow rate [g/s]
2023.6.26 10:19 | 1.211
2023.6.29 6:45 | 1.119
2023.7.3 1:44 | 1.093
2023.7.7 1:06 | 1.100
2023.7.7 8:16 | 1.070
Standard deviation | 0.055

## Slide 5: 2 K Static Heat Load Measurement (Mass Flow Method)
Close bottom cooldown valve (CV6218), endcap pre-cooling valve (CV6215) and J-T valve (CV6214).
Calibration of mass flow rate is done with a heater immersed in the helium liquid level tank.
In most cases, flow rate is measured from liquid level 80 % to 75 %.
when dynamic heat load is 0, then
where
 | Using latent heat | By calibration (2023.6.7-12) | By calibration (2023.6.26-29)
Module 2 K static heat load [W] | (23.04 J/g) * flow [g/s] | (24.0±0.2 J/g) * flow [g/s] | (23.8±0.4J/g) * flow [g/s]
Static heat load [W] | 25.8±1.3 | 28.6±0.4 | 27.5±0.8
Better fitting accuracy (important during single-cavity Q measurement) if closer to averaged flow rate during calibration.
Better fitting accuracy with more data (n).
*Note: standard deviation is used in this table

## Slide 6: Dynamic Heat Load of 8 Cavities at 16 MV/m and 21 MV/m
Close bottom cooldown valve (CV6218), endcap pre-cooling valve (CV6215) and J-T valve (CV6214).
Total uncertainty of dynamic heat load is
RF Voltage, cool down (date, rate) | Measurement Date | Mass flow [g/s] | Qdyn [W] (Cal. 6.7-12) | Qdyn [W] (Cal. 6.26-29) | Qdyn [W] (use latent heat)
133 MV, fast cool down (6.1, > 40 g/s) | 2023.6.2 | 3.75±0.05 | 61.3±1.3 | 61.6±1.4 | 60.5±1.9
133 MV, fast cool down (6.1, > 40 g/s) | 2023.6.2 | 3.82±0.05 | 63.0±1.3 | 63.2±1.4 | 62.3±1.9
133 MV, slow cool down (6.19, ~ 7 g/s) | 2023.7.3 | 3.60±0.05 | 57.7±1.2 | 58.0±1.4 | 57.2±1.9
133 MV, slow cool down (6.19, ~ 7 g/s) | 2023.7.4 | 3.64±0.05 | 58.7±1.2 | 59.0±1.4 | 58.4±1.9
174 MV, fast cool down (6.1, > 40 g/s) | 2023.6.2 | 5.62±0.05 | 106.2±1.4 | 106.0±1.9 | 103.8±1.9
*Note: standard deviation is used in this table

## Slide 7: Dynamic Heat Load of Single Cavity at 16 MV/m
The J-T valve CV6214 is fixed at 20 %.
Heater power 6-7 W keeps the liquid level constant at ~ 90 %.
The dynamic heat load is calibrated with a heater immersed in the liquid level tank; heater power 0-15 W is adopted to get best fitting accuracy around 7-8 W.
 | Cav1 | Cav2 | Cav3 | Cav4 | Cav5 | Cav6 | Cav7 | Cav8
Dynamic heat load at 16 MV/m [W] | 7.4±0.3 | 10.0±0.3 | 8.0±0.3 | 6.8±0.3 | 6.8±0.3 | 7.1±0.3 | 10.5±0.3 | 6.2±0.3
*Note: standard deviation is used in this table

## Slide 8: Back up

## Slide 9: 2 K Pumps and Mass Flow Meters