## Slide 1
Laser Transport Optics and Alignment
Randy Thurman-Keup
PIP2 Laserwire Final Design Review
May 2, 2024

## Slide 2: Laserwire Beamline
H- Beam
Laser
Mirror Box Optics
Beamline Box Optics
Vultures-eye View
Laser Room
12+1 Beamline Stations
Feedback
Station
Shield Wall
Vacuum Transport Line

## Slide 3: Feedback Box Optics
 2-inch cube beamsplitter
      Alternatively –
Pair of  3-inch mirrors
Optical BPM
Optical BPM
Thorlabs Lateral Effect Position Sensor
Not at focal point
Not 90 degrees to maximize mirror acceptance
1064 nm Laser
Alignment/Feedback Laser
Notes: Specify the IR and alignment laser (don’t quantify wavelength)

## Slide 4: Mirror Box Optics
Motorized stage to insert pickoff mirror
Manual stage to adjust longitudinal position of downward laser trajectory during alignment
Extra space for potential feedback or monitoring devices
Laser in Transport Line
Motorized stage
1
3
2
Manual Stage
2
3
Pickoff Mirror
1

## Slide 5: Beamline Box Optics
Laser from Mirror Box
Motorized stage
1
3
2
4
5
Focus Lens
6
Sampler (dielectric mirror)
Photodiode/BPM
Camera
H- beam into page
2
3
1
1
1
Mirror
3
3
3
4
4
4
5
6
7
Beamsplitter
7
Detectors may only be there for commissioning

## Slide 6: Dump Box
Two located on horizontal and vertical exit windows
Absorber with hole for photodiode
Dump aperture to be same as optical vacuum viewport
Dump interior painted black or black anodized
Interior angle ~ 30 degrees to minimize back reflections
Seal between dump box and optical viewport flange will be light-tight

## Slide 7: Laser Transport Optics
1064 nm
Laser System
Alignment/Feedback (AF)
Laser System
1064 nm
Transport Prep
1064/AF
Combiner
AF
Transport Prep
Mirror Expander
Feedback
Mirror
Bend
Mirror
Feedback
Detector
Transport Vacuum Window
Laser Room
Linac
12+1 Laserwire Stations
Feedback
Detector
Alignment
Path
End of Linac
Detector

## Slide 8: Laser Transport Optics
1064 nm
Laser System
1064 nm
Transport Prep
AF
Transport Prep
Laser Room
Linac
12+1 Laserwire Stations
Optics to adjust phase space parameters to those needed for transport to end of linac
  Adjust phase space parameters to match expander optics and produce desired lattice for linac transport
  Equalize phase space of transverse planes to eliminate astigmatism using natural astigmatism of lenses
Transport Prep
Alignment/Feedback (AF)
Laser System

## Slide 9: Laser Transport Optics
1064 nm
Laser System
1064 nm
Transport Prep
1064/AF Combiner
AF
Transport Prep
Laser Room
Linac
12+1 Laserwire Stations
Optics to combine the two lasers onto the same path
  Likely just a 1064 nm mirror
Combiner
Alignment/Feedback (AF)
Laser System

## Slide 10: Laser Transport Optics
1064 nm
Laser System
1064 nm
Transport Prep
1064/AF
Combiner
AF
Transport Prep
Laser Room
Linac
12+1 Laserwire Stations
Alignment
Path
Long reflective path on optical table in laser room to do initial alignment for transport
Alignment Path
Alignment/Feedback (AF)
Laser System

## Slide 11: Laser Transport Optics
1064 nm
Laser System
1064 nm
Transport Prep
1064/AF
Combiner
AF
Transport Prep
Mirror Expander
Laser Room
Linac
12+1 Laserwire Stations
Reflective expander to avoid chromatic effects
  e.g. Thorlabs BE06R
Mirror Expander
Alignment
Path
Alignment/Feedback (AF)
Laser System

## Slide 12: Laser Transport Optics
1064 nm
Laser System
1064 nm
Transport Prep
1064/AF
Combiner
AF
Transport Prep
Mirror Expander
Feedback
Mirror
Transport Vacuum Window
Laser Room
Linac
12+1 Laserwire Stations
Alignment
Path
Piezo mirror to control alignment with feedback from downstream detectors
  e.g. Thorlabs : PDP90A plus KPA101 plus POLARIS-K2S2P
Feedback Mirror
Alignment/Feedback (AF)
Laser System

## Slide 13: Laser Transport Optics
1064 nm
Laser System
1064 nm
Transport Prep
1064/AF
Combiner
AF
Transport Prep
Mirror Expander
Feedback
Mirror
Bend
Mirror
Transport Vacuum Window
Laser Room
Linac
12+1 Laserwire Stations
Alignment
Path
Mirror located in vacuum above laser room to direct the laser down the transport pipe
Part of initial alignment of transport line
Bend Mirror
Alignment/Feedback (AF)
Laser System

## Slide 14: Laser Transport Optics
1064 nm
Laser System
1064 nm
Transport Prep
1064/AF
Combiner
AF
Transport Prep
Mirror Expander
Feedback
Mirror
Bend
Mirror
Feedback
Detector
Transport Vacuum Window
Laser Room
Linac
12+1 Laserwire Stations
Feedback
Detector
Alignment
Path
Detectors used to feedback to the piezo mirror to control alignment
  Feedback uses the AF laser picked off from the transport line in a special box 30 m from the laser room
  A second detector will be placed near the middle of the linac to be used if the first one has insufficient sensitivity
Feedback Detectors
Alignment/Feedback (AF)
Laser System

## Slide 15: Laser Transport Optics
1064 nm
Laser System
1064 nm
Transport Prep
1064/AF
Combiner
AF
Transport Prep
Mirror Expander
Feedback
Mirror
Bend
Mirror
Feedback
Detector
Transport Vacuum Window
Laser Room
Linac
12+1 Laserwire Stations
Feedback
Detector
Alignment
Path
End of Linac
Detector
A detector will be placed at the end of the transport line to aid in verification of the alignment
  This detector will be sensitive to the 1064 nm and AF lasers
EOL Detector
Alignment/Feedback (AF)
Laser System

## Slide 16: Laser Transport Optics
Transport line optical arrangement features a single optical waist generated by an initial collimating section in the laser room
Each laserwire has an insertable pickoff mirror in the mirror box which leads to final focus lenses inside the beamline boxes
Insertable Pickoff Mirror
Final Focusing Lens
Not to scale
Laser Optical Waist
H- Beamline
Laser Room

## Slide 17: Laser Transport Optics
A single waist but many pickoff points requires an optical evaluation of each laserwire location to ensure that the laser radius has the proper size at all points of interest
  Laser size does not become too large (or too small) down the transport
  Does not become too small on vacuum windows and optical elements
  Is small enough at the H- intersection
  Plan is ~5 mm
Four different waist sizes
Solid is M2 = 1
Dashed is M2 = 2
Notes: Add that we’re looking at using 4 or 5 mm

## Slide 18: Laser Transport Optics
Test of sensitivity of optics to ideal collimating lens
  Control of optics via lens position, and stability of optics to small movements in lens position
Collimating Lens Position
Waist Position
Focus Length
Size on Lens
Focus Size
Waist Size
Waist Size
Waist Position
Size on Lens
Focus Size
Focus Length
Waist Position
Waist Position
Collimating Lens Position
Region to avoid
Small size on focus lens  small size on vacuum window
Focus Size (mm)
Focus Length (m)
Focus Length ~constant
Focus Size small

## Slide 19: Laser Transport Optics – Test
Collimated a HeNe laser over 130 m length
Expander Lens
Collimator Lens
Source Point
1 in
11 m
40 m
64 m
90 m
130 m
1 in
Lens tilted horizontally
Refocused after 117 m
Horizontal
Vertical
Second test

## Slide 20: Laser Transport Optics – Feedback
Feedback will be implemented with a detector before the shield wall
  9 mm active area; sub-micron sensitivity
  Approximately 35 m from laser room
  Laser partially focused onto detector
Ray tracing simulations
  Study steering errors corresponding to deflections between +/- 6 mm at the end of linac
Thorlabs Lateral Effect Position Sensor
Less than 1% loss of laser beam through components
Position at H- beam:
<2 μm variation
Size on feedback sensor
Position at feedback sensor

## Slide 21: Installation and Alignment – Laser Room
Install transport pipe and mirror housing in laser room structure
Align transport pipe to nominal location relative to beam (Alignment Crew)
Install target at end of linac and align to nominal location relative to beam (Alignment Crew)
Setup laser systems on optical table
Align and focus green laser through complete transport line
Wall
Wall
Ceiling
Ceiling
Optical Table
Transport Pipe
Mirror Housing
Mechanical Supports

## Slide 22: Installation and Alignment – Transport Line
Alignment Laser
X
Target
(Second one on other side)
Install shield wall support and align to green laser
Install feedback box and align to green laser
Install mirror boxes
  Align to nominal location relative to beam (Alignment Crew)
  Verify transverse alignment with green laser
Install pipe hangers and align to green laser
Alignment Laser
X
Mirror Box
Pipe Hanger

## Slide 23: Installation and Alignment – Beamline Box Mount
Install mirror box mirror with longitudinal position adjustment
  Align mirror with alignment laser
Install beamline alignment fixture
  Designed as a sparse replica of actual beamline box
  Allows alignment without bulk of actual box
Adjust mounting to align vacuum flanges
Adjust downward alignment laser trajectory to intercept both targets
Beamline Alignment Fixture
Laser Targets
Mirror Box

## Slide 24: Installation and Alignment – Beamline Box Optics
Alignment of beamline box optics will happen in the laser lab
Utilizes three elements
  An alignment plate that replicates the beamline box mounting plate
  A laser aligned with the alignment plate to replicate the trajectory from the mirror box
  A target also aligned with the alignment plate and positioned at the design location of the beamline
The aligned optics box can then be mounted on the beamline box mount
Alignment Laser
Laser Target
X
Alignment Plate

## Slide 25: Summary
Have a design for the transport optics
  Have considered control of the laser
  Have tested the transport over 130 m
Have a procedure for alignment of laser transport and beamline optics
  Transport alignment is coincident with installation of transport line
  Beamline box mounting alignment is done after installation of warm units
  Beamline box optics alignment is done in laser room

## Slide 26: Backup Slides

## Slide 27: Laserwire Locations
Twelve + one laserwire stations in the PIP2 Linac proper (last not shown below)
The laser room is upstream of the H- source which is at 0 m

## Slide 28: Laser Transport Optics – Feedback
Steering error in x

## Slide 29: Laser Transport Optics – Feedback
Geometric Acceptance
Loss of laser beam from steering errors