## Slide 1
Ben Russell, Radiation Safety Officer
DOE 420.2D Assist Visit to Fermilab
September, 2023
Radiation Generating Devices (RGDs), Equivalencies, and Accelerators

## Slide 2: RGD vs Accelerator
Radiation Generating Device (RGD) in DOE G 441.1-1C: Collective term for devices which produce ionizing radiation, including, certain sealed radioactive sources, small particle accelerators used for single purpose applications which produce ionizing radiation (e.g., radiography), and electron generating devices that produce X-rays incidentally.
  For sealed radioactive sources this only applies to a device with a source that produces radiation fields exceeding 100 millirem on out hour at one foot from the source.
Updated definition for accelerator in DOE O 420.2D: A device and its components employing electrostatic or electromagnetic fields to impart kinetic energy to molecular, atomic, or sub-atomic particles and capable of creating a radiological area as defined by 10 CFR Part 835, Occupational Radiation Protection. Accelerator components include injectors, targets, beam dumps, detectors, experimental enclosures, accelerator enclosures, experimental areas, and experimental apparatus utilizing the accelerator. The accelerator also includes associated support and test facilities, equipment, systems, and utilities necessary to operate the accelerator or utilize the accelerated beam.

## Slide 3: RGD vs Accelerator
Interpretation from DOE HQ Assist visit to JLab in June:
  If voltage or RF is applied, and there is a potential that a Radiation Area (≥ 5 mR / hr at 1 foot) may be created, that device is an accelerator
  If the intent is to test a component by simulating its function in the accelerator, and there is a potential that a Radiation Area may be created, that device is an accelerator
  If the device produces an accelerating gradient or capacity < 10 MeV, that device may fall into one of five equivalencies, and may be managed according to applicable invoked standards

## Slide 4: RGD vs Accelerator
All RGDs as determined by DOE O 420.2C were run through the definitions of DOE O 420.2D to determine their new designation. The possibilities are:
  Radiation Generating Device
  Exempt / Equivalent Accelerator
  Non-Exempt / Non-Equivalent Accelerator

## Slide 5: RGD vs Accelerator
RGDs and Exempt / Equivalent Accelerators (Devices) are handled under the new RGD Program based on 10 CFR 835, DOE G 441.1-1C, and the 5-equivalency category/standards listed in DOE O 420.2D:
  Devices operated in accordance with American National Standards Institute/Health Physics Society (ANSI/HPS) N43.2-2021, Radiation Safety for X-Ray Diffraction and Fluorescence Analysis Equipment;
  Devices operated in accordance with ANSI/HPS N43.3-2008, General Radiation Safety Standard for Installations Using Non-Medical X-Ray and Sealed Gamma-Ray Sources, Energies up to 10 MeV;
  Devices operated in accordance with ANSI/HPS N43.5-2005, Radiological Safety Standard for the Design of Radiographic and Radioscopic Non-Medical X-Ray Equipment Below 1 MeV (reaffirmed 2013);
  Neutron generators operated below 600 keV in accordance with National Council on Radiation Protection and Measurements (NCRP) Report 72-1983, Radiation Protection and Measurement of Low-Voltage Neutron Generators; and
  Devices operated at or below 10 MeV in accordance with NCRP Report 144-2003, Radiation Protection for Particle Accelerator Facilities and/or ANSI/HPS N43.1-2011, Radiation Safety for the Design and Operation of Particle Accelerators.

## Slide 6: RGD vs Accelerator
Fermilab Radiation Generating Device Working Group
  The Radiation Generating Device Program approved in August 2022 was developed based on the DOE Guide 441.1-1C Radiation Protection Program Guide
    This Guide incorporates the following documents:
      ANSI N43.3-2008, American National Standard for General Radiation Safety – Installations Using Non-Medical X-Ray and Sealed Gamma-Ray Sources, Energies Up to 10 MeV,
      ANSI/HPS N43.2-2021, Radiation Safety for X-Ray Diffraction and Fluorescence Analysis Equipment, and
      ANSI/HPS N43.5-2005, Radiological Safety Standard for the Design of Radiographic and Fluoroscopic Industrial X-Ray Equipment
    The Working Group performed a detailed Gap Analysis to include the two missing documents referenced in the DOE O 420.2D 3.c.(3) Equivalencies section:
      NCRP Report 72-1983, Radiation Protection and Measurement of Low-Voltage Neutron Generators, and
      NCRP Report 144-2003, Radiation Protection for Particle Accelerator Facilities and/or ANSI-HPS N43.1-2011, Radiation Safety for the Design and Operation of Particle Accelerators.

## Slide 7: RGD vs Accelerator
The RGD Procedure – ESH-RPO-RGD-01 Radiation Generating Devices was updated to clarify the new requirements and assist in the classification process, record consolidation, and review/approval process for the “Devices.”
RP Form 133 – Radiation Generating Device Checklist was updated to walk the RGD Custodian, the assigned RSO, and management through the classification of the Device and the requirements to meet the appropriate standards.
RP Form 113 – Preliminary Hazard Assessment Checklist for Deuterium-Based Neutron Generators was updated to include additional requirements not captured in RP Form 133, as well as remove duplicative data.
RP Form 108 – List of Radiation Generating Devices (FRCM 362) was updated to include the new necessary categories for classification and record management for each Device under the RGD program.

## Slide 8: RP Form 133 Overview
Section 1 – Appointment of a custodian by Management
Section 2 – Collection of information
  Device Operating Conditions,
  Training,
  Shielding Requirements,
  Interlocks and Device Controls, and
  Other important and relevant safety considerations
Section 3 – Official Device Classification (based on information from Section 2)
Section 4 – Additional Safety Approvals and Walkdowns necessary for Device Operation

## Slide 9: RGD vs Accelerator
Is voltage or RF applied?
Radiation Area possible?
≥10 MeV?
Device is an accelerator
ISM or may request equivalency
May NOT have an equivalency
Device is an RGD
yes
yes
yes
yes
no
Is ionizing radiation possible?
no
yes
no
Device not under this program
no
Device not under this program

## Slide 10: RGD vs Accelerator
RGD Requirements:
Follow RGD Program Requirements
Complete RP Form 133
RWP as Operation Authorization Document
SRSO and Div. Management approval prior to initial operations
Identified on RP Form 108 List of Radiation Generating Devices
RGDs:
JEOL Sem
Helios SEM
SPECS KERIOS
ToF SIMS
Device is an RGD

## Slide 11: RGD vs Accelerator
Equivalent Accelerator Requirements:
FSO approval of equivalency request
Follow RGD Program Requirements
Complete RP Form 133
RWP as Operation Authorization Document
SRSO and Div. Management approval prior to initial operations
Identified on RP Form 108 List of Radiation Generating Devices
Equivalent Accelerators:
DD Generator
DT Generator
A2D2
Brucker S1 Turbo
Spoke Test Cave (STC)
ISM or may request equivalency

## Slide 12: RGD vs Accelerator
Accelerator Requirements:
Follow DOE O 420.2D Requirements (FESHM 2010)
Have a SAD Chapter
Have an ASE
Conduct an ARR prior to commissioning and/or operations
Obtain FSO approval prior to commissioning and/or operations
Beam Permit/Running Condition (or equivalent) as Operation Authorization Document
Accelerators:
Main Fermilab Complex (Linac to BNB, NuMI, Muon Campus, SY)
  Now also includes Pre-Acc
FAST
VTS
CMTS1
PIP-II IT Cryo Cave
May NOT have an equivalency

## Slide 13: RGD vs Accelerator
Device | Potential Device Classification under 420.2D | Accelerating Potential | Greater than 5 mR/hr at 1 ft?
A2D2 | Equivalent Accelerator | 9 MeV electrons or gamma | Yes
Bruker S1 Turbo | Equivalent Accelerator | 40 kV and 20 uA (15kV and 60 uA without filter) | Yes
CMTS1 Cave | Non-Exempt Accelerator | Cavity vertical test gradient Eacc_VTS >18 MV/m Cavity horizontal test gradient Eacc_HTS > 18 MV/m | Yes
PIP2IT Cave | Non-Exempt Accelerator | Nominal = 18.8 MV/m; Operational maximum = 20.7 MV/m; Administrative limit = 21.6 MV/m | Yes
DD Generator | Equivalent Accelerator | 100 kV | Yes
DT Generator | Equivalent Accelerator | 120 kV | Yes
JEOL SEM | RGD | N/A | No
Helios SEM | RGD | N/A | No
SPECS KREIOS | RGD | N/A | No
STC | Non-Exempt Accelerator | A maximum gradient of 39 MV/m. | Yes
VTS (VCTF) | Non-Exempt Accelerator | Maximum reached so far is 47.7 MV/m | Yes

## Slide 14: CMTS1 Accelerator
Brief Overview
  Testing of 1.3 and 3.9 GHz LCLS-II HE cryomodules for SLAC
SAD Chapter: VI-02 Cryomodule Test Stand 1 (CMTS1)
ASE: VII.A-03 Cryomodule Test Stand 1 (CMTS1) Accelerator ASE
Operator training/qualification process
  Cryomodule test operators are restricted to personnel who have been appointed by LCLS-II-HE Senior Team Leader and authorized by SRF Systems Department Head or his/her designee. The prospective operators must follow the test procedure (ED0029229) and have demonstrated competence through 3 consecutive CM tests with qualified operators before they can be appointed.

## Slide 15: PIP2IT Accelerator
Brief Overview
  2020-21: PIP-II Injector Test area with H- beam up to 17 MeV
  2022: Area decommissioned and retrofit for cryomodule testing
  2023-28: Test area for PIP-II 650 MHz and 325 MHz cryomodules
SAD Chapter: VI-03 Proton Improvement Plan II Integration Test (PIP2IT)
ASE: VII-A.04 Proton Improvement Plan II Integration Test (PIP2IT) Accelerator ASE
Operator training/qualification process
  Interface Between PIP2IT Control Room and Main Control Room, PIP-II docDB 4872
  PIP2IT Access Modes, PIP-II docDB 4931
  PIP2IT Search and Secure Procedure, PIP-II docDB 4946
  Authorization record maintained in PIP2IT control room by PIP-II Operations Coordinator

## Slide 16: VTS Accelerator
Brief Overview
  The VTS Accelerator is built to test and characterize superconducting radiofrequency (SRF) cavities in liquid helium at temperatures between 1.3 K and 4.4 K. The facility is not designed to accelerate any beams or has no beam sources but some of SRF cavities tested in the facility can produce an accelerating gradient or capacity >10 MeV.
  The facility contributes to various projects (especially qualification of cavities for cryomodules) and R&D (e.g., testing new treatments, performing studies of how superconductors behave under different relevant conditions for accelerators and quantum systems, testing of new instrumentation at cryogenic temperatures).
  The facility has three test stands, each of which features a large cryogenic vessel called a “dewar” and can host and test multiple SRF cavities simultaneously depending on a cavity geometry. Recent statistics show that ~300 RF tests have been carried out in the last 12 month.
SAD Chapter: VI-04 Vertical Test Stand (VTS)
ASE: VII-A.05 Vertical Test Stand (VTS) Accelerator ASE
Operator training/qualification process
  Unsupervised operation of the VTS Accelerator is restricted to personnel who have been authorized by the SRF Measurements & Research Department Head or his/her designee. This authorization is contingent upon successful completion of general and facility-specific training as well as demonstrated competence in operating the VTS Accelerator. Details are described in “VCTF RF/Cavity Test Operations Authorized Operators List”.