## Slide 1: MQXFA08b, MQXFA13 and MQXFA14b Test Results
Anis Ben Yahia
Brookhaven National Laboratory

## Slide 2: Outline
Quench test results
  MQXFA08b
  MQXFA13
  MQXFA14b
Splice measurements
Magnetic measurement results

## Slide 3: Outline
Quench test results
  MQXFA08b
  MQXFA13
  MQXFA14b
Splice measurements
Magnetic measurement results

## Slide 4: MQXFA08b
All quenches occurred in coil 219.
After the thermal cycle 2 additional quenches above nominal current (at 16.51 kA and 16.48 kA)
First re-assembly: built by replacing the limiting coil (213) in MQXFA08 by a new coil (219) in quadrant 3.

## Slide 5: MQXFA08b – Quench localization from the Quench Antenna Array
First quench in the lead end of the coil.
Following quenches in the straight section (no recurring location).
Return end
Lead end

## Slide 6: Outline
Quench test results
  MQXFA08b
  MQXFA13
  MQXFA14b
Splice measurements
Magnetic measurement results

## Slide 7: MQXFA13
Training “typical” up to quench #9.
Drop in quench current (315 A) between quench 10 and 14.
Training continued from quench 15.
Reached 16.53 kA on the 25th ramp.

## Slide 8: MQXFA13
Following the ramp at +30/-100 A/s, there was a drop in quench current in the subsequent ramp at 20 A/s.
Magnet trains back again to 16.53 kA.
Behavior was repeatable 3 times.
Ramps at 100 A/s showed lower quench current (16.22 and 16.21 kA).
All these quenches occurred in coil 227 (Q1.).

## Slide 9: MQXFA13 – Quench localization from the Quench Antenna Array
Recurring quenches in the coil ends.
Return end
Lead end
Quench 10 to14

## Slide 10: MQXFA13 – Quench localization from the Quench Antenna Array
Recurring quenches in the coil ends.
Return end
Lead end
Quenches following the +30/-100 A/s ramp.

## Slide 11: MQXFA13 – Quench localization from the Quench Antenna Array
Recurring quenches in the coil ends.
Return end
Lead end
Quenches at 100 A/s.

## Slide 12: MQXFA13
Analysis of CMM data showed that all MQXFA13 coils had small arc-length in the ends, with coil 227 the smallest ever.
FEM analysis of MQXFA13 as built showed risk of high strain with moderate prestress, and this issue may be prevented with higher local prestress.
MQXFA13 test was stopped in order to replace coil 227 and avoid damage to the other coils.

## Slide 13: Outline
Quench test results
  MQXFA08b
  MQXFA13
  MQXFA14b
Splice measurements
Magnetic measurement results

## Slide 14: MQXFA14b
Achieved nominal + 300A after 5 quenches.
All quenches except the first were in quadrant 3 (coil 143).
After the thermal cycle
  One quench at 16.42 kA.
  7 ramps without quench (2 to nominal + 300 A).
Assembled using revised MQXFA specifications and targets.

## Slide 15: MQXFA14b – Quench localization from the Quench Antenna Array
Return end
Lead end
No recurring quench location.

## Slide 16: Outline
Quench test results
  MQXFA08b
  MQXFA13
  MQXFA14b
Splice measurements
Magnetic measurement results

## Slide 17: MQXFA08b – Splices measurements

## Slide 18: MQXFA13 – Splices measurements

## Slide 19: MQXFA14b – Splices measurements

## Slide 20: Outline
Quench test results
  MQXFA08b
  MQXFA13
  MQXFA14b
Splice measurements
Magnetic measurement results

## Slide 21: MQXFA08b
Measured integrated gradient at 16.23 kA: 562.41 T
Mix of previously tested coil and one new coil (coil 219 in Q3).
The field errors after loading were within the bounds, except for a5 and a7 that are on the boarder line.
No magnetic shims installed.
Change in the low order multipole from warm pre-test to cold to the final warm measurement.

## Slide 22: MQXFA08b
Variation of a3/b3 along the length warm and cold measurements at 10 and 16.3 kA.

## Slide 23: MQXFA08b
Variation of a3/b3 vs current.

## Slide 24: MQXFA13
At the end of the quench tests, two sets of magnetic measurements were taken.
The difference was in the pre-cycle:
  1st pre-cycle was at +/- 20 A/s.
  2nd pre-cycle was at +30/-100 A/s.
The measurements were taken at 13 kA.

## Slide 25: MQXFA13
The two measurements were identical.
(“fast” refers to the measurement with +30/-100 A/s pre-cycle).

## Slide 26: MQXFA13
Example of the variation along the length.

## Slide 27: MQXFA14b
Warm measurements and magnetic shim configuration.
G. L. Sabbi

## Slide 28: MQXFA14b
 | Integ harmonics variation (unit) |  | 
 | Measured | Expected | Difference
b3 | 1.83 | 1.57 | -0.26
a3 | 2.82 | 2.50 | -0.32
b4 | -0.06 | 0.04 | 0.10
a4 | 0.37 | -0.01 | -0.38
b5 | 0.38 | 0.08 | -0.30
a5 | -0.05 | -0.29 | -0.24
b6 | 0.77 | 0.11 | -0.66
a6 | 0.06 | -0.07 | -0.13
Measured integrated gradient at 16.23 kA: 563.14 T
Field quality at nominal current and change from the magnetic shims.

## Slide 29: Summary
MQXFA08b:
  First successful test of an MQXFA magnet after replacing a coil.
  All quenches occurred in the new coil.
MQXFA13:
  Test stopped for replacing the coil showing ramp-rate dependence.
MQXFA14b:
  Assembled with revised specs and pre-load targets.
  5 quenches to reach acceptance current.
Magnetic field measurements:
  Field quality generally within the expected bound.
  Correction from the magnetic shims in A14b in agreement with calculated values.
All cold and warm high voltage withstand tests were successful.
All magnet splices are less than 0.5 nΩ.

## Slide 30: THANK YOU

## Slide 31: Additional slides

## Slide 32: Integrated Gradient and Field Quality
Chart: Integrated Gradient Variations in MQXFA Models
| A03 | A04 | A05 | A06 | A10 | A11 | A08b | A14b
Corrections applied | -1.1399 | -0.28434 | 0.55046
As measured | -0.61254 | 0.33746 | 0.36045 | 0.030355 | 0.75807
 | -1.5 | -1.5 | -1.5 | -1.5 | -1.5 | -1.5 | -1.5 | -1.5
 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5
 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Average integrated gradient at 16.23 kA: 562.38 T
Correction factors applied to account for measurements taken at different currents as well as changes in the measurement setup and coil design.

## Slide 33: Field Quality – All tested magnets