## CSR Report
MQXFA12b |  |  | Coil Selection Review Electricals
Updated: 18-Jan-24 |  |  |  |  | Retest | Retest | MQXFA17 | Extra
|  |  | 158 | 159 | 224 | 225 | 150
See DocDB #956 / SU-1010-1903
RLQ | [8 empty] | Acceptance Range
| R | (mΩ) | 612 | 602 | 611 | 608 | 602 |  | 590-625
| Ls (mH) | 20 Hz | 10.088 | 9.966 | 10.077 | 10.043 | 9.97 |  | 9.9-10.3
|  | 100 Hz | 6.419 | 6.405 | 6.428 | 6.421 | 6.41 |  | 6.2-6.6
|  | 1 kHz | 5.814 | 5.804 | 5.864 | 5.861 | 5.81 |  | 5.5-6
| Q | 20 Hz | 1.275 | 1.294 | 1.279 | 1.283 | 1.3 |  | 1.2-1.4
|  | 100 Hz | 5.469 | 5.541 | 5.615 | 5.625 | 5.51 |  | 5.2-5.8
|  | 1 kHz | 8.536 | 8.474 | 8.493 | 8.551 | 8.6 |  | 5.8-10.7
QH R (Ω)
| NT | A01 | N/A | N/A | N/A | N/A | N/A
| T | A02 | N/A | N/A | N/A | N/A | N/A
| T-P | B01 | 2.2 | 2 | 2.5 | 2.3 | 1.9 |  | 1.4-2.7
| NT-P | B02 | 2.1 | 2.1 | 2.6 | 2.3 | 1.9 |  | 1.4-2.7
| T-M | B03 | 2.2 | 2.1 | 2.5 | 2.3 | 1.7 |  | 1.4-2.7
| NT-M | B04 | 2.2 | 2.1 | 2.5 | 2.3 | 1.7 |  | 1.4-2.7
VT R (mΩ) Individual coils
@1 Amp |  | A1 | 0.00022 | 0.00021 | 0.00018 | 0.00015 | 0.00019
|  | A2 | 0.00027 | 0.00026 | 0.00023 | 0.00019 | 0.00026
A3
A4
A5
A6
A7
A8
B8
B7
B6
B5
B4
B3
|  | B2 | 0.61149 | 0.6016 | 0.61018 | 0.60798 | 0.6018
|  | B1 | 0.61154 | 0.60164 | 0.61021 | 0.60803 | 0.6019
HIPOT |  |  | OK | OK | OK | OK | OK
IMPULSE |  |  | OK | OK | OK | OK | OK
X

## Step 2 Hipots all coils
MQXFA12b | [5 empty] | Date | 2023-12-19 |  |  |  | Notes:
Step 2: Coil Acceptance (Post-Shipping) Hipots
|  | Coil | [5 empty] | Day 1
| Coil 158 | V_Test/V_Target [V] | I_leak_Final [µA] | I_leak_Max [µA] |  |  | Test Time:
| PHB01 | 3680/3680 V | 0 | 1.3 |  |  | Room Temperature : | 20.6
| PHB02 | 3680/3680 V | 0 | 1.3 |  |  | Relative Humidity: | 31.7
| PHB03 | 3680/3680 V | 0 | 1.8
| PHB04 | 3680/3680 V | 0 | 2.5 | PHB01 |  |  | PHB02 |  |  | PHB03 |  |  | PHB04
| Pole | 100/100 V | 0 | 0.4 | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | LE IL Endshoe |  |  | RE IL Endshoe
| LE IL Endshoe | 1000/1000 | 0 | 0.1 | [12 empty] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA]
| LE OL Endshoe | 1000/1000 | 0 | 0.2 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.3 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.2 | 1000/1000 | 0 | 0.2
| RE IL Endshoe | 1000/1000 | 0 | NA
| RE OL Endshoe | 1000/1000 | 0 | NA | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0.1 | 2500/2500 | 0 | 0 |  |  |  | 1000/1000 | 0 | 0
Red text = test not passed
| Measure | Target |  |  | Actual / Target (Leakage) |  | Date | Hipot Checks
| QH to Coil | 3680 V |  |  | (< 10 μA leakage) |  | 2022-07-05
| Coil to Pole | 100 V
| Coil to Endshoe | 1000 V
| QH to Endshoe | 2500 V
| Endshoe to Endshoe | 1000 V
|  | Coil | [5 empty] | Day 1
| Coil 159 | V_Test/V_Target [V] | I_leak_Final [µA] | I_leak_Max [µA] |  |  | Test Time:
| PHB01 | 3680/3680 V | 0 | 1.3 |  |  | Room Temperature : | 22
| PHB02 | 3680/3680 V | 0 | 2.1 |  |  | Relative Humidity: | 46
| PHB03 | 3680/3680 V | 0.1 | 1.7
| PHB04 | 3680/3680 V | 0 | 1.8 | PHB01 |  |  | PHB02 |  |  | PHB03 |  |  | PHB04
| Pole | 100/100 V | 0.1 | 0.6 | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | LE IL Endshoe |  |  | RE IL Endshoe
| LE IL Endshoe | 1000/1000 | 0 | 0 | [12 empty] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA]
| LE OL Endshoe | 1000/1000 | 0 | 0.3 | 2500/2500 | 0 | 0.3 | 2500/2500 | 0 | 0.4 | 2500/2500 | 0 | 0.1 | 2500/2500 | 0 | 0.2 | 1000/1000 | 0 | 0.1
| RE IL Endshoe | 1000/1000 | 0 | NA
| RE OL Endshoe | 1000/1000 | 0 | NA | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 |  |  |  | 1000/1000 | 0 | 0
Red text = test not passed
| Measure | Target |  |  | Actual / Target (Leakage) |  | Date | Hipot Checks
| QH to Coil | 3680 V |  |  | (< 10 μA leakage) |  | 2022-07-05
| Coil to Pole | 100 V
| Coil to Endshoe | 1000 V
| QH to Endshoe | 2500 V
| Endshoe to Endshoe | 1000 V
|  | Coil | [5 empty] | Day 1
| Coil 224 | V_Test/V_Target [V] | I_leak_Final [µA] | I_leak_Max [µA] |  |  | Test Time:
| PHB01 | 3680/3680 V | 0 | 1.8 |  |  | Room Temperature : | 21.7
| PHB02 | 3680/3680 V | 0 | 2.5 |  |  | Relative Humidity: | 56
| PHB03 | 3680/3680 V | 0 | 1.7
| PHB04 | 3680/3680 V | 0.2 | 3.5 | PHB01 |  |  | PHB02 |  |  | PHB03 |  |  | PHB04
| Pole | 100/100 V | 0 | 0.4 | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | LE IL Endshoe |  |  | RE IL Endshoe
| LE IL Endshoe | 1000/1000 | 0 | 0.1 | [12 empty] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA]
| LE OL Endshoe | 1000/1000 | 0 | 0.20.0 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.2 | 1000/1000 | 0 | 0.1
| RE IL Endshoe | 1000/1000 | 0 | NA
| RE OL Endshoe | 1000/1000 | 0 | NA | 2500/2500 | 0 | 0.1 | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0.1 | 2500/2500 | 0 | 0 |  |  |  | 1000/1000 | 0 | 0
Red text = test not passed
| Measure | Target |  |  | Actual / Target (Leakage) |  | Date | Hipot Checks
| QH to Coil | 3680 V |  |  | (< 10 μA leakage) |  | 2022-12-19
| Coil to Pole | 100 V
| Coil to Endshoe | 1000 V
| QH to Endshoe | 2500 V
| Endshoe to Endshoe | 1000 V
|  | Coil | [5 empty] | Day 1
| Coil 225 | V_Test/V_Target [V] | I_leak_Final [µA] | I_leak_Max [µA] |  |  | Test Time:
| PHB01 | 3680/3680 V | 0 | 1.2 |  |  | Room Temperature : | 21
| PHB02 | 3680/3680 V | 0 | 1.3 |  |  | Relative Humidity: | 56
| PHB03 | 3680/3680 V | 0 | 2.4
| PHB04 | 3680/3680 V | 0.1 | 2.7 | PHB01 |  |  | PHB02 |  |  | PHB03 |  |  | PHB04
| Pole | 100/100 V | 0 | 0.4 | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | LE IL Endshoe |  |  | RE IL Endshoe
| LE IL Endshoe | 1000/1000 | 0 | 3.8 | [12 empty] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA]
| LE OL Endshoe | 1000/1000 | 0 | 0.1 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.1 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.2 | 1000/1000 | 0 | 0.1
| RE IL Endshoe | 1000/1000 | 0 | NA
| RE OL Endshoe | 1000/1000 | 0 | NA | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 |  |  |  | 1000/1000 | 0 | 0
Red text = test not passed
| Measure | Target |  |  | Actual / Target (Leakage) |  | Date | Hipot Checks
| QH to Coil | 3680 V |  |  | (< 10 μA leakage) |  | 2023-12-19
| Coil to Pole | 100 V
| Coil to Endshoe | 1000 V
| QH to Endshoe | 2500 V
| Endshoe to Endshoe | 1000 V
|  | Coil | [5 empty] | Day 1
| Coil 150 | V_Test/V_Target [V] | I_leak_Final [µA] | I_leak_Max [µA] |  |  | Test Time:
| PHB01 | 3680/3680 V | 0 | 2 |  |  | Room Temperature : | 21.4
| PHB02 | 3680/3680 V | 0 | 3.3 |  |  | Relative Humidity: | 45
| PHB03 | 3680/3680 V | 0 | 3.1
| PHB04 | 3680/3680 V | 0 | 3.1 | PHB01 |  |  | PHB02 |  |  | PHB03 |  |  | PHB04
| Pole | 100/100 V | 0 | 0.4 | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | LE IL Endshoe |  |  | RE IL Endshoe
| LE IL Endshoe | 1000/1000 | 0 | 0 | [12 empty] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA]
| LE OL Endshoe | 1000/1000 | 0 | 0.2 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.3 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.2 | 1000/1000 | 0 | 0.1
| RE IL Endshoe | 1000/1000 | 0 | NA
| RE OL Endshoe | 1000/1000 | 0 | NA | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 |  |  |  | 1000/1000 | 0 | 0
Red text = test not passed
| Measure | Target |  |  | Actual / Target (Leakage) |  | Date | Hipot Checks
| QH to Coil | 3680 V |  |  | (< 10 μA leakage) |  | 2024-01-16
| Coil to Pole | 100 V
| Coil to Endshoe | 1000 V
| QH to Endshoe | 2500 V
| Endshoe to Endshoe | 1000 V
|  | Coil | [5 empty] | Day 1
| COIL xxx | V_Test/V_Target [V] | I_leak_Final [µA] | I_leak_Max [µA] |  |  | Test Time:
| PHB01 | 3680/3680 V | 0 | 0.2 |  |  | Room Temperature : | 22
| PHB02 | 3680/3680 V | 0 | 0.3 |  |  | Relative Humidity: | 45
| PHB03 | 3680/3680 V | 0 | 0.3
| PHB04 | 3680/3680 V | 0 | 0.3 | PHB01 |  |  | PHB02 |  |  | PHB03 |  |  | PHB04
| Pole | 100/100 V | 0 | 0.3 | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | LE IL Endshoe |  |  | RE IL Endshoe
| LE IL Endshoe | 1000/1000 | 0 | 0.1 | [12 empty] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA] | V_Test\Break [V] | I_leak_Final [µA] | I_leak [µA]
| LE OL Endshoe | 1000/1000 | 0 | 0.3 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.2 | 2500/2500 | 0 | 0.3 | 1000/1000 | 0 | 0.1
| RE IL Endshoe | 1000/1000 | 0 | NA
| RE OL Endshoe | 1000/1000 | 0 | NA | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 | 2500/2500 | 0 | 0 |  |  |  | 1000/1000 | 0 | 0
Red text = test not passed
| Measure | Target |  |  | Actual / Target (Leakage) |  | Date | Hipot Checks
| QH to Coil | 3680 V |  |  | (< 10 μA leakage) |  | 2024-01-16
| Coil to Pole | 100 V

## Step 2 158 Impulse
Time(s) | 1.0 kV | 1.5 kV | 2.0 kV | 2.5 kV | 158 | Time (s) | 1.0 kV | 1.5 kV | 2.0 kV | 2.5 kV |  | Notes
0 | 0 | 0 | 0 | 0 | A12b Coil Pack QXFA 158 Outer to Inner | 0 | 0 | 0 | 0 | 0 | A12b Coil Pack QXFA 158 Inner to Outer
2e-07 | 0 | 0 | 0 | 0 |  | 2e-07 | 0 | 0 | 0 | 0
4e-07 | 0 | 0 | 0 | 0 |  | 4e-07 | 0 | 0 | 0 | 0
6e-07 | 154 | 308 | 450 | 592 |  | 6e-07 | 154 | 308 | 426 | 592
8e-07 | 626 | 1065 | 1514 | 2035 |  | 8e-07 | 614 | 1065 | 1514 | 2035
1e-06 | 980 | 1491 | 2011 | 2484 |  | 1e-06 | 980 | 1491 | 2011 | 2484
1.2e-06 | 945 | 1372 | 1775 | 2130 |  | 1.2e-06 | 945 | 1372 | 1775 | 2130
1.4e-06 | 850 | 1254 | 1633 | 2011 |  | 1.4e-06 | 850 | 1254 | 1633 | 2011
1.6e-06 | 850 | 1278 | 1680 | 2059 |  | 1.6e-06 | 850 | 1278 | 1656 | 2059
1.8e-06 | 850 | 1278 | 1656 | 2035 |  | 1.8e-06 | 850 | 1254 | 1656 | 2035
2e-06 | 827 | 1207 | 1609 | 1964 |  | 2e-06 | 827 | 1230 | 1585 | 1964
2.2e-06 | 815 | 1207 | 1585 | 1940 |  | 2.2e-06 | 815 | 1207 | 1562 | 1940
2.4e-06 | 803 | 1183 | 1562 | 1917 |  | 2.4e-06 | 803 | 1183 | 1562 | 1917
2.6e-06 | 791 | 1183 | 1538 | 1893 |  | 2.6e-06 | 791 | 1183 | 1538 | 1893
2.8e-06 | 780 | 1159 | 1514 | 1869 |  | 2.8e-06 | 780 | 1159 | 1514 | 1869
3e-06 | 768 | 1136 | 1491 | 1822 |  | 3e-06 | 768 | 1136 | 1491 | 1822
3.2e-06 | 744 | 1112 | 1443 | 1798 |  | 3.2e-06 | 744 | 1112 | 1443 | 1775
3.4e-06 | 732 | 1088 | 1420 | 1751 |  | 3.4e-06 | 732 | 1065 | 1396 | 1751
3.6e-06 | 709 | 1041 | 1372 | 1704 |  | 3.6e-06 | 709 | 1041 | 1372 | 1704
3.8e-06 | 697 | 1041 | 1325 | 1656 |  | 3.8e-06 | 697 | 1017 | 1325 | 1656
4e-06 | 673 | 994 | 1325 | 1633 |  | 4e-06 | 673 | 994 | 1301 | 1633
4.2e-06 | 661 | 994 | 1278 | 1585 |  | 4.2e-06 | 661 | 994 | 1301 | 1585
4.4e-06 | 650 | 970 | 1278 | 1562 |  | 4.4e-06 | 650 | 970 | 1254 | 1562
4.6e-06 | 638 | 946 | 1254 | 1538 |  | 4.6e-06 | 638 | 946 | 1254 | 1514
4.8e-06 | 638 | 923 | 1207 | 1491 |  | 4.8e-06 | 626 | 946 | 1207 | 1491
5e-06 | 626 | 923 | 1207 | 1467 |  | 5e-06 | 626 | 923 | 1183 | 1467
5.2e-06 | 602 | 899 | 1159 | 1443 |  | 5.2e-06 | 602 | 899 | 1159 | 1443
5.4e-06 | 591 | 875 | 1136 | 1396 |  | 5.4e-06 | 591 | 875 | 1136 | 1396
5.6e-06 | 579 | 852 | 1112 | 1372 |  | 5.6e-06 | 579 | 852 | 1112 | 1372
5.8e-06 | 555 | 828 | 1088 | 1325 |  | 5.8e-06 | 555 | 828 | 1088 | 1325
6e-06 | 543 | 804 | 1041 | 1278 |  | 6e-06 | 543 | 804 | 1041 | 1278
6.2e-06 | 520 | 781 | 1017 | 1254 |  | 6.2e-06 | 531 | 781 | 1017 | 1254
6.4e-06 | 508 | 734 | 970 | 1183 |  | 6.4e-06 | 508 | 757 | 970 | 1183
6.6e-06 | 484 | 710 | 946 | 1136 |  | 6.6e-06 | 484 | 710 | 923 | 1136
6.8e-06 | 472 | 686 | 899 | 1088 |  | 6.8e-06 | 472 | 686 | 899 | 1088
7e-06 | 449 | 639 | 852 | 1041 |  | 7e-06 | 449 | 639 | 852 | 1041
7.2e-06 | 425 | 615 | 804 | 994 |  | 7.2e-06 | 425 | 615 | 804 | 994
7.4e-06 | 402 | 592 | 781 | 946 |  | 7.4e-06 | 402 | 592 | 757 | 946
7.6e-06 | 378 | 544 | 734 | 899 |  | 7.6e-06 | 378 | 544 | 710 | 899
7.8e-06 | 354 | 521 | 686 | 828 |  | 7.8e-06 | 354 | 521 | 686 | 828
8e-06 | 331 | 473 | 639 | 781 |  | 8e-06 | 343 | 473 | 639 | 781
8.2e-06 | 307 | 450 | 592 | 710 |  | 8.2e-06 | 307 | 450 | 592 | 710
8.4e-06 | 283 | 402 | 544 | 663 |  | 8.4e-06 | 283 | 402 | 521 | 663
8.6e-06 | 260 | 379 | 497 | 592 |  | 8.6e-06 | 260 | 379 | 497 | 592
8.8e-06 | 236 | 331 | 426 | 544 |  | 8.8e-06 | 236 | 331 | 426 | 544
9e-06 | 213 | 308 | 379 | 473 |  | 9e-06 | 213 | 308 | 379 | 473
9.2e-06 | 177 | 260 | 331 | 426 |  | 9.2e-06 | 189 | 260 | 331 | 426
9.4e-06 | 154 | 237 | 284 | 355 |  | 9.4e-06 | 165 | 237 | 284 | 355
9.6e-06 | 130 | 189 | 237 | 284 |  | 9.6e-06 | 130 | 189 | 237 | 284
9.8e-06 | 106 | 166 | 189 | 237 |  | 9.8e-06 | 118 | 166 | 189 | 237
1e-05 | 94 | 118 | 166 | 189 |  | 1e-05 | 83 | 118 | 142 | 189
1.02e-05 | 59 | 95 | 118 | 118 |  | 1.02e-05 | 59 | 71 | 118 | 118
1.04e-05 | 47 | 47 | 47 | 71 |  | 1.04e-05 | 35 | 47 | 47 | 71
1.06e-05 | 12 | 24 | 0 | 0 |  | 1.06e-05 | 12 | 24 | 24 | 0
1.08e-05 | -11 | -23 | -46 | -46 |  | 1.08e-05 | 0 | -23 | -23 | -46
1.1e-05 | -34 | -46 | -70 | -94 |  | 1.1e-05 | -23 | -70 | -70 | -94
1.12e-05 | -58 | -94 | -117 | -165 |  | 1.12e-05 | -58 | -94 | -117 | -165
1.14e-05 | -70 | -117 | -165 | -212 |  | 1.14e-05 | -82 | -117 | -165 | -212
1.16e-05 | -93 | -165 | -212 | -259 |  | 1.16e-05 | -105 | -165 | -212 | -259
1.18e-05 | -117 | -188 | -259 | -307 |  | 1.18e-05 | -117 | -188 | -259 | -307
1.2e-05 | -141 | -236 | -283 | -378 |  | 1.2e-05 | -141 | -236 | -283 | -354
1.22e-05 | -164 | -259 | -330 | -425 |  | 1.22e-05 | -164 | -259 | -330 | -425
1.24e-05 | -188 | -283 | -378 | -472 |  | 1.24e-05 | -188 | -283 | -378 | -472
1.26e-05 | -212 | -307 | -425 | -520 |  | 1.26e-05 | -200 | -307 | -425 | -520
1.28e-05 | -223 | -330 | -472 | -567 |  | 1.28e-05 | -223 | -354 | -449 | -567
1.3e-05 | -247 | -378 | -496 | -614 |  | 1.3e-05 | -247 | -378 | -496 | -614
1.32e-05 | -271 | -401 | -543 | -662 |  | 1.32e-05 | -271 | -401 | -543 | -662
1.34e-05 | -294 | -425 | -591 | -709 |  | 1.34e-05 | -282 | -425 | -567 | -709
1.36e-05 | -306 | -472 | -614 | -756 |  | 1.36e-05 | -306 | -472 | -614 | -756
1.38e-05 | -330 | -496 | -662 | -803 |  | 1.38e-05 | -330 | -496 | -662 | -803
1.4e-05 | -342 | -520 | -685 | -851 |  | 1.4e-05 | -342 | -520 | -685 | -851
1.42e-05 | -365 | -543 | -733 | -898 |  | 1.42e-05 | -365 | -543 | -733 | -898
1.44e-05 | -389 | -591 | -756 | -945 |  | 1.44e-05 | -377 | -567 | -756 | -945
1.46e-05 | -401 | -614 | -803 | -993 |  | 1.46e-05 | -401 | -614 | -803 | -993
1.48e-05 | -424 | -638 | -827 | -1040 |  | 1.48e-05 | -424 | -638 | -827 | -1040
1.5e-05 | -436 | -662 | -874 | -1087 |  | 1.5e-05 | -436 | -662 | -874 | -1064
1.52e-05 | -460 | -685 | -898 | -1111 |  | 1.52e-05 | -448 | -685 | -898 | -1111
1.54e-05 | -471 | -709 | -945 | -1158 |  | 1.54e-05 | -471 | -709 | -945 | -1158
1.56e-05 | -483 | -733 | -969 | -1182 |  | 1.56e-05 | -483 | -733 | -969 | -1182
1.58e-05 | -507 | -756 | -993 | -1229 |  | 1.58e-05 | -507 | -756 | -993 | -1229
1.6e-05 | -519 | -780 | -1016 | -1253 |  | 1.6e-05 | -519 | -780 | -1016 | -1253
1.62e-05 | -530 | -803 | -1040 | -1300 |  | 1.62e-05 | -530 | -803 | -1040 | -1300
1.64e-05 | -542 | -827 | -1087 | -1324 |  | 1.64e-05 | -542 | -827 | -1064 | -1324
1.66e-05 | -554 | -827 | -1087 | -1348 |  | 1.66e-05 | -554 | -827 | -1087 | -1348
1.68e-05 | -566 | -851 | -1135 | -1371 |  | 1.68e-05 | -566 | -851 | -1111 | -1371
1.7e-05 | -578 | -874 | -1135 | -1419 |  | 1.7e-05 | -578 | -874 | -1135 | -1419
1.72e-05 | -590 | -898 | -1158 | -1442 |  | 1.72e-05 | -590 | -898 | -1158 | -1442
1.74e-05 | -601 | -898 | -1182 | -1466 |  | 1.74e-05 | -601 | -898 | -1182 | -1466
1.76e-05 | -613 | -922 | -1206 | -1490 |  | 1.76e-05 | -613 | -922 | -1206 | -1490
1.78e-05 | -625 | -945 | -1206 | -1513 |  | 1.78e-05 | -625 | -922 | -1206 | -1513
1.8e-05 | -637 | -945 | -1253 | -1537 |  | 1.8e-05 | -625 | -945 | -1229 | -1537
1.82e-05 | -637 | -969 | -1253 | -1537 |  | 1.82e-05 | -637 | -969 | -1253 | -1537
1.84e-05 | -649 | -969 | -1277 | -1561 |  | 1.84e-05 | -649 | -969 | -1277 | -1561
1.86e-05 | -660 | -969 | -1277 | -1584 |  | 1.86e-05 | -649 | -969 | -1277 | -1584
1.88e-05 | -660 | -993 | -1300 | -1608 |  | 1.88e-05 | -660 | -993 | -1300 | -1608
1.9e-05 | -660 | -993 | -1300 | -1608 |  | 1.9e-05 | -660 | -993 | -1300 | -1608
1.92e-05 | -672 | -993 | -1324 | -1632 |  | 1.92e-05 | -660 | -1016 | -1324 | -1632
1.94e-05 | -672 | -1016 | -1324 | -1632 |  | 1.94e-05 | -672 | -1016 | -1324 | -1632
1.96e-05 | -672 | -1016 | -1324 | -1632 |  | 1.96e-05 | -672 | -1016 | -1324 | -1655
1.98e-05 | -672 | -1016 | -1324 | -1632 |  | 1.98e-05 | -684 | -1016 | -1324 | -1655
2e-05 | -684 | -1016 | -1324 | -1655 |  | 2e-05 | -684 | -1016 | -1324 | -1655
2.02e-05 | -684 | -1016 | -1348 | -1655 |  | 2.02e-05 | -684 | -1016 | -1348 | -1655
2.04e-05 | -684 | -1016 | -1324 | -1655 |  | 2.04e-05 | -684 | -1016 | -1348 | -1655
2.06e-05 | -684 | -1016 | -1324 | -1655 |  | 2.06e-05 | -684 | -1016 | -1348 | -1655
2.08e-05 | -684 | -1016 | -1324 | -1655 |  | 2.08e-05 | -684 | -1016 | -1324 | -1655
2.1e-05 | -672 | -1016 | -1324 | -1632 |  | 2.1e-05 | -672 | -1016 | -1324 | -1632
2.12e-05 | -672 | -993 | -1324 | -1632 |  | 2.12e-05 | -672 | -1016 | -1324 | -1632
2.14e-05 | -660 | -993 | -1300 | -1608 |  | 2.14e-05 | -672 | -993 | -1300 | -1608
2.16e-05 | -660 | -993 | -1277 | -1584 |  | 2.16e-05 | -660 | -993 | -1300 | -1584
2.18e-05 | -649 | -969 | -1277 | -1561 |  | 2.18e-05 | -649 | -969 | -1277 | -1561
2.2e-05 | -637 | -945 | -1253 | -1537 |  | 2.2e-05 | -637 | -945 | -1253 | -1537
2.22e-05 | -625 | -945 | -1206 | -1490 |  | 2.22e-05 | -625 | -945 | -1229 | -1513
2.24e-05 | -613 | -922 | -1182 | -1466 |  | 2.24e-05 | -613 | -922 | -1206 | -1466
2.26e-05 | -601 | -898 | -1158 | -1442 |  | 2.26e-05 | -601 | -898 | -1158 | -1442
2.28e-05 | -590 | -874 | -1135 | -1395 |  | 2.28e-05 | -590 | -874 | -1135 | -1395
2.3e-05 | -566 | -851 | -1111 | -1348 |  | 2.3e-05 | -566 | -851 | -1111 | -1371
2.32e-05 | -554 | -827 | -1064 | -1324 |  | 2.32e-05 | -554 | -827 | -1087 | -1324
2.34e-05 | -530 | -803 | -1040 | -1277 |  | 2.34e-05 | -542 | -803 | -1040 | -1277
2.36e-05 | -519 | -756 | -993 | -1229 |  | 2.36e-05 | -519 | -780 | -993 | -1253
2.38e-05 | -495 | -733 | -969 | -1182 |  | 2.38e-05 | -495 | -733 | -969 | -1182
2.4e-05 | -471 | -709 | -922 | -1135 |  | 2.4e-05 | -483 | -709 | -922 | -1135
2.42e-05 | -460 | -662 | -874 | -1087 |  | 2.42e-05 | -460 | -685 | -898 | -1087
2.44e-05 | -436 | -638 | -827 | -1016 |  | 2.44e-05 | -436 | -638 | -827 | -1040
2.46e-05 | -412 | -614 | -780 | -969 |  | 2.46e-05 | -412 | -614 | -803 | -993
2.48e-05 | -389 | -567 | -733 | -922 |  | 2.48e-05 | -389 | -567 | -756 | -922
2.5e-05 | -353 | -543 | -709 | -851 |  | 2.5e-05 | -365 | -543 | -709 | -874
2.52e-05 | -330 | -496 | -638 | -780 |  | 2.52e-05 | -342 | -496 | -662 | -803
2.54e-05 | -306 | -449 | -591 | -733 |  | 2.54e-05 | -318 | -472 | -614 | -733
2.56e-05 | -282 | -425 | -543 | -662 |  | 2.56e-05 | -294 | -425 | -567 | -685
2.58e-05 | -259 | -378 | -496 | -614 |  | 2.58e-05 | -271 | -378 | -520 | -614
2.6e-05 | -235 | -330 | -449 | -543 |  | 2.6e-05 | -247 | -354 | -449 | -543
2.62e-05 | -212 | -307 | -401 | -472 |  | 2.62e-05 | -212 | -330 | -401 | -496
2.64e-05 | -188 | -283 | -330 | -425 |  | 2.64e-05 | -188 | -283 | -354 | -425
2.66e-05 | -153 | -236 | -307 | -354 |  | 2.66e-05 | -164 | -236 | -307 | -378
2.68e-05 | -129 | -188 | -236 | -283 |  | 2.68e-05 | -129 | -188 | -259 | -307
2.7e-05 | -105 | -141 | -188 | -236 |  | 2.7e-05 | -105 | -165 | -188 | -236
2.72e-05 | -82 | -117 | -141 | -165 |  | 2.72e-05 | -82 | -117 | -141 | -188
2.74e-05 | -46 | -70 | -94 | -94 |  | 2.74e-05 | -58 | -94 | -94 | -117
2.76e-05 | -23 | -23 | -46 | -46 |  | 2.76e-05 | -34 | -46 | -46 | -46
2.78e-05 | 0 | 0 | 24 | 24 |  | 2.78e-05 | -11 | 0 | 0 | 0
2.8e-05 | 24 | 47 | 71 | 95 |  | 2.8e-05 | 24 | 24 | 47 | 71
2.82e-05 | 47 | 71 | 118 | 142 |  | 2.82e-05 | 47 | 71 | 95 | 142
2.84e-05 | 83 | 118 | 166 | 213 |  | 2.84e-05 | 71 | 118 | 142 | 189
2.86e-05 | 106 | 166 | 213 | 260 |  | 2.86e-05 | 94 | 142 | 213 | 260
2.88e-05 | 118 | 189 | 260 | 308 |  | 2.88e-05 | 118 | 189 | 237 | 308
2.9e-05 | 142 | 237 | 284 | 379 |  | 2.9e-05 | 142 | 213 | 284 | 355
2.92e-05 | 165 | 260 | 331 | 426 |  | 2.92e-05 | 165 | 237 | 331 | 426
2.94e-05 | 189 | 284 | 379 | 473 |  | 2.94e-05 | 189 | 284 | 355 | 473
2.96e-05 | 213 | 308 | 426 | 544 |  | 2.96e-05 | 201 | 308 | 426 | 521
2.98e-05 | 236 | 355 | 473 | 568 |  | 2.98e-05 | 224 | 331 | 450 | 568
3e-05 | 248 | 379 | 497 | 639 |  | 3e-05 | 248 | 355 | 497 | 615
3.02e-05 | 272 | 402 | 544 | 686 |  | 3.02e-05 | 272 | 402 | 521 | 663
3.04e-05 | 283 | 426 | 568 | 710 |  | 3.04e-05 | 283 | 426 | 568 | 710
3.06e-05 | 307 | 450 | 592 | 757 |  | 3.06e-05 | 295 | 450 | 592 | 757
3.08e-05 | 319 | 473 | 639 | 781 |  | 3.08e-05 | 319 | 473 | 615 | 781
3.1e-05 | 331 | 497 | 663 | 828 |  | 3.1e-05 | 331 | 497 | 663 | 804
3.12e-05 | 354 | 521 | 686 | 875 |  | 3.12e-05 | 343 | 521 | 686 | 852
3.14e-05 | 366 | 544 | 710 | 899 |  | 3.14e-05 | 354 | 544 | 710 | 875
3.16e-05 | 378 | 544 | 734 | 923 |  | 3.16e-05 | 366 | 544 | 734 | 899
3.18e-05 | 390 | 568 | 757 | 946 |  | 3.18e-05 | 378 | 568 | 757 | 946
3.2e-05 | 390 | 592 | 781 | 970 |  | 3.2e-05 | 390 | 592 | 757 | 970
3.22e-05 | 402 | 592 | 804 | 994 |  | 3.22e-05 | 402 | 592 | 781 | 970
3.24e-05 | 413 | 615 | 804 | 994 |  | 3.24e-05 | 413 | 615 | 804 | 994
3.26e-05 | 425 | 615 | 828 | 1017 |  | 3.26e-05 | 413 | 615 | 828 | 1017
3.28e-05 | 425 | 639 | 828 | 1041 |  | 3.28e-05 | 425 | 615 | 828 | 1041
3.3e-05 | 437 | 639 | 852 | 1041 |  | 3.3e-05 | 425 | 639 | 852 | 1041
3.32e-05 | 437 | 639 | 852 | 1041 |  | 3.32e-05 | 437 | 639 | 852 | 1041
3.34e-05 | 437 | 639 | 852 | 1065 |  | 3.34e-05 | 437 | 639 | 852 | 1065
3.36e-05 | 449 | 663 | 875 | 1065 |  | 3.36e-05 | 437 | 639 | 852 | 1065
3.38e-05 | 437 | 663 | 875 | 1065 |  | 3.38e-05 | 437 | 663 | 852 | 1065
3.4e-05 | 437 | 639 | 852 | 1065 |  | 3.4e-05 | 437 | 639 | 852 | 1065
3.42e-05 | 437 | 639 | 852 | 1065 |  | 3.42e-05 | 437 | 663 | 852 | 1065
3.44e-05 | 437 | 639 | 852 | 1041 |  | 3.44e-05 | 437 | 639 | 852 | 1041
3.46e-05 | 437 | 639 | 852 | 1041 |  | 3.46e-05 | 437 | 639 | 852 | 1041
3.48e-05 | 425 | 639 | 828 | 1017 |  | 3.48e-05 | 437 | 639 | 828 | 1041
3.5e-05 | 425 | 615 | 828 | 1017 |  | 3.5e-05 | 425 | 615 | 828 | 1017
3.52e-05 | 413 | 615 | 804 | 994 |  | 3.52e-05 | 413 | 615 | 804 | 1017
3.54e-05 | 413 | 615 | 804 | 970 |  | 3.54e-05 | 413 | 615 | 804 | 994
3.56e-05 | 402 | 592 | 781 | 970 |  | 3.56e-05 | 402 | 592 | 781 | 970
3.58e-05 | 390 | 568 | 757 | 946 |  | 3.58e-05 | 402 | 592 | 757 | 946
3.6e-05 | 390 | 568 | 734 | 923 |  | 3.6e-05 | 390 | 568 | 757 | 923
3.62e-05 | 378 | 544 | 734 | 899 |  | 3.62e-05 | 378 | 544 | 734 | 899
3.64e-05 | 366 | 544 | 710 | 875 |  | 3.64e-05 | 366 | 544 | 710 | 875
3.66e-05 | 354 | 521 | 686 | 828 |  | 3.66e-05 | 354 | 521 | 686 | 852
3.68e-05 | 343 | 497 | 663 | 804 |  | 3.68e-05 | 343 | 497 | 663 | 828
3.7e-05 | 331 | 473 | 639 | 781 |  | 3.7e-05 | 331 | 497 | 639 | 781
3.72e-05 | 319 | 473 | 592 | 757 |  | 3.72e-05 | 319 | 473 | 615 | 757
3.74e-05 | 295 | 450 | 568 | 710 |  | 3.74e-05 | 307 | 450 | 592 | 710
3.76e-05 | 283 | 426 | 544 | 663 |  | 3.76e-05 | 283 | 426 | 544 | 686
3.78e-05 | 272 | 379 | 521 | 639 |  | 3.78e-05 | 272 | 402 | 521 | 639
3.8e-05 | 260 | 379 | 473 | 592 |  | 3.8e-05 | 260 | 379 | 497 | 615
3.82e-05 | 236 | 355 | 450 | 568 |  | 3.82e-05 | 248 | 355 | 450 | 568
3.84e-05 | 224 | 331 | 426 | 521 |  | 3.84e-05 | 224 | 331 | 426 | 544
3.86e-05 | 213 | 308 | 379 | 473 |  | 3.86e-05 | 213 | 308 | 402 | 497
3.88e-05 | 189 | 260 | 355 | 426 |  | 3.88e-05 | 189 | 284 | 355 | 450
3.9e-05 | 177 | 260 | 308 | 402 |  | 3.9e-05 | 177 | 260 | 331 | 402
3.92e-05 | 154 | 237 | 284 | 355 |  | 3.92e-05 | 154 | 237 | 308 | 379
3.94e-05 | 142 | 189 | 260 | 308 |  | 3.94e-05 | 142 | 213 | 260 | 331
3.96e-05 | 118 | 166 | 237 | 260 |  | 3.96e-05 | 130 | 189 | 237 | 284
3.98e-05 | 106 | 166 | 189 | 237 |  | 3.98e-05 | 106 | 166 | 213 | 260
4e-05 | 94 | 118 | 166 | 189 |  | 4e-05 | 94 | 142 | 166 | 213
4.02e-05 | 71 | 95 | 118 | 142 |  | 4.02e-05 | 71 | 118 | 142 | 166
4.04e-05 | 59 | 71 | 95 | 95 |  | 4.04e-05 | 59 | 71 | 95 | 118
4.06e-05 | 35 | 47 | 47 | 71 |  | 4.06e-05 | 35 | 47 | 71 | 71
4.08e-05 | 12 | 24 | 24 | 24 |  | 4.08e-05 | 24 | 24 | 24 | 24
4.1e-05 | 0 | -23 | 0 | -23 |  | 4.1e-05 | 12 | 0 | 0 | 0
4.12e-05 | -11 | -46 | -46 | -70 |  | 4.12e-05 | -11 | -23 | -46 | -46
4.14e-05 | -34 | -70 | -94 | -117 |  | 4.14e-05 | -23 | -46 | -70 | -94
4.16e-05 | -46 | -94 | -117 | -141 |  | 4.16e-05 | -46 | -70 | -94 | -141
4.18e-05 | -70 | -117 | -141 | -188 |  | 4.18e-05 | -58 | -94 | -141 | -165
4.2e-05 | -82 | -141 | -188 | -212 |  | 4.2e-05 | -82 | -141 | -165 | -212
4.22e-05 | -93 | -165 | -212 | -259 |  | 4.22e-05 | -93 | -141 | -188 | -236
4.24e-05 | -117 | -188 | -236 | -283 |  | 4.24e-05 | -105 | -165 | -236 | -283
4.26e-05 | -129 | -188 | -259 | -330 |  | 4.26e-05 | -129 | -188 | -259 | -307
4.28e-05 | -141 | -212 | -307 | -354 |  | 4.28e-05 | -141 | -212 | -283 | -354
4.3e-05 | -153 | -236 | -307 | -378 |  | 4.3e-05 | -153 | -236 | -307 | -378
4.32e-05 | -164 | -259 | -330 | -425 |  | 4.32e-05 | -164 | -259 | -330 | -401
4.34e-05 | -176 | -283 | -354 | -449 |  | 4.34e-05 | -176 | -259 | -354 | -449
4.36e-05 | -188 | -283 | -378 | -472 |  | 4.36e-05 | -188 | -283 | -378 | -472
4.38e-05 | -200 | -307 | -401 | -496 |  | 4.38e-05 | -200 | -307 | -401 | -496
4.4e-05 | -212 | -330 | -425 | -520 |  | 4.4e-05 | -212 | -330 | -425 | -520
4.42e-05 | -223 | -330 | -449 | -543 |  | 4.42e-05 | -223 | -330 | -449 | -543
4.44e-05 | -235 | -354 | -449 | -567 |  | 4.44e-05 | -223 | -354 | -449 | -567
4.46e-05 | -247 | -354 | -496 | -591 |  | 4.46e-05 | -247 | -354 | -472 | -591
4.48e-05 | -247 | -378 | -496 | -614 |  | 4.48e-05 | -247 | -378 | -496 | -614
4.5e-05 | -259 | -378 | -520 | -638 |  | 4.5e-05 | -259 | -378 | -496 | -614
4.52e-05 | -271 | -401 | -543 | -638 |  | 4.52e-05 | -259 | -401 | -520 | -638
4.54e-05 | -271 | -401 | -543 | -662 |  | 4.54e-05 | -271 | -401 | -543 | -662
4.56e-05 | -271 | -401 | -543 | -685 |  | 4.56e-05 | -282 | -425 | -543 | -685
4.58e-05 | -282 | -425 | -567 | -685 |  | 4.58e-05 | -282 | -425 | -567 | -685
4.6e-05 | -282 | -425 | -567 | -685 |  | 4.6e-05 | -282 | -425 | -567 | -685
4.62e-05 | -294 | -425 | -567 | -709 |  | 4.62e-05 | -282 | -425 | -567 | -709
4.64e-05 | -282 | -425 | -567 | -709 |  | 4.64e-05 | -294 | -425 | -567 | -709
4.66e-05 | -294 | -449 | -591 | -709 |  | 4.66e-05 | -294 | -449 | -567 | -709
4.68e-05 | -294 | -449 | -567 | -709 |  | 4.68e-05 | -294 | -449 | -567 | -733
4.7e-05 | -294 | -449 | -591 | -709 |  | 4.7e-05 | -294 | -449 | -591 | -709
4.72e-05 | -294 | -449 | -591 | -709 |  | 4.72e-05 | -294 | -449 | -591 | -733
4.74e-05 | -294 | -449 | -591 | -709 |  | 4.74e-05 | -294 | -449 | -591 | -733
4.76e-05 | -294 | -425 | -567 | -709 |  | 4.76e-05 | -294 | -449 | -567 | -709
4.78e-05 | -294 | -449 | -567 | -709 |  | 4.78e-05 | -294 | -449 | -591 | -709
4.8e-05 | -282 | -425 | -567 | -685 |  | 4.8e-05 | -294 | -449 | -567 | -709
4.82e-05 | -282 | -425 | -567 | -685 |  | 4.82e-05 | -282 | -425 | -567 | -709
4.84e-05 | -282 | -425 | -567 | -685 |  | 4.84e-05 | -282 | -425 | -567 | -685
4.86e-05 | -271 | -401 | -543 | -662 |  | 4.86e-05 | -282 | -425 | -543 | -685
4.88e-05 | -271 | -401 | -543 | -662 |  | 4.88e-05 | -282 | -425 | -543 | -662
4.9e-05 | -271 | -401 | -520 | -638 |  | 4.9e-05 | -271 | -401 | -543 | -662
4.92e-05 | -259 | -378 | -520 | -638 |  | 4.92e-05 | -271 | -401 | -520 | -638
4.94e-05 | -259 | -378 | -496 | -614 |  | 4.94e-05 | -259 | -401 | -520 | -614
4.96e-05 | -247 | -378 | -472 | -591 |  | 4.96e-05 | -259 | -378 | -496 | -614
4.98e-05 | -247 | -354 | -472 | -567 |  | 4.98e-05 | -247 | -378 | -472 | -591
5e-05 | -235 | -354 | -449 | -567 |  | 5e-05 | -235 | -354 | -472 | -567
5.02e-05 | -223 | -330 | -425 | -543 |  | 5.02e-05 | -235 | -354 | -449 | -543
5.04e-05 | -212 | -330 | -425 | -520 |  | 5.04e-05 | -223 | -330 | -425 | -543
5.06e-05 | -200 | -307 | -401 | -496 |  | 5.06e-05 | -212 | -307 | -401 | -520
5.08e-05 | -200 | -283 | -378 | -472 |  | 5.08e-05 | -200 | -307 | -401 | -496
5.1e-05 | -188 | -283 | -354 | -449 |  | 5.1e-05 | -200 | -307 | -378 | -449
5.12e-05 | -176 | -259 | -330 | -425 |  | 5.12e-05 | -188 | -283 | -354 | -449
5.14e-05 | -164 | -259 | -330 | -401 |  | 5.14e-05 | -176 | -259 | -330 | -425
5.16e-05 | -153 | -236 | -307 | -378 |  | 5.16e-05 | -164 | -236 | -307 | -378
5.18e-05 | -141 | -236 | -283 | -354 |  | 5.18e-05 | -153 | -236 | -307 | -354
5.2e-05 | -141 | -212 | -259 | -307 |  | 5.2e-05 | -141 | -212 | -283 | -330
5.22e-05 | -129 | -188 | -236 | -307 |  | 5.22e-05 | -129 | -188 | -259 | -307
5.24e-05 | -105 | -165 | -212 | -259 |  | 5.24e-05 | -117 | -188 | -236 | -283
5.26e-05 | -105 | -165 | -212 | -236 |  | 5.26e-05 | -105 | -165 | -212 | -259
5.28e-05 | -93 | -141 | -188 | -212 |  | 5.28e-05 | -93 | -141 | -188 | -236
5.3e-05 | -82 | -117 | -141 | -188 |  | 5.3e-05 | -82 | -141 | -165 | -188
5.32e-05 | -70 | -94 | -141 | -165 |  | 5.32e-05 | -82 | -117 | -141 | -165
5.34e-05 | -58 | -94 | -94 | -141 |  | 5.34e-05 | -58 | -94 | -141 | -141
5.36e-05 | -46 | -70 | -94 | -94 |  | 5.36e-05 | -58 | -70 | -94 | -117
5.38e-05 | -34 | -46 | -70 | -70 |  | 5.38e-05 | -34 | -70 | -70 | -94
5.4e-05 | -23 | -23 | -46 | -46 |  | 5.4e-05 | -23 | -46 | -46 | -70
5.42e-05 | -11 | -23 | -23 | -23 |  | 5.42e-05 | -11 | -23 | -23 | -23
5.44e-05 | 0 | 0 | 0 | 0 |  | 5.44e-05 | -11 | 0 | 0 | 0
5.46e-05 | 12 | 24 | 24 | 47 |  | 5.46e-05 | 0 | 0 | 24 | 24
5.48e-05 | 24 | 24 | 47 | 71 |  | 5.48e-05 | 12 | 24 | 47 | 47
5.5e-05 | 35 | 47 | 71 | 95 |  | 5.5e-05 | 24 | 47 | 47 | 71
5.52e-05 | 47 | 71 | 95 | 118 |  | 5.52e-05 | 35 | 47 | 71 | 95
5.54e-05 | 59 | 95 | 95 | 142 |  | 5.54e-05 | 47 | 71 | 95 | 118
5.56e-05 | 71 | 95 | 142 | 166 |  | 5.56e-05 | 59 | 95 | 118 | 142
5.58e-05 | 71 | 118 | 142 | 189 |  | 5.58e-05 | 71 | 95 | 142 | 166
5.6e-05 | 83 | 118 | 166 | 213 |  | 5.6e-05 | 83 | 118 | 166 | 189
5.62e-05 | 94 | 142 | 189 | 237 |  | 5.62e-05 | 83 | 142 | 166 | 213
5.64e-05 | 106 | 142 | 189 | 237 |  | 5.64e-05 | 94 | 142 | 189 | 237
5.66e-05 | 106 | 166 | 213 | 260 |  | 5.66e-05 | 106 | 166 | 213 | 260
5.68e-05 | 118 | 189 | 237 | 284 |  | 5.68e-05 | 106 | 166 | 237 | 260
5.7e-05 | 118 | 189 | 237 | 308 |  | 5.7e-05 | 118 | 189 | 237 | 284
5.72e-05 | 130 | 189 | 260 | 331 |  | 5.72e-05 | 130 | 189 | 237 | 308
5.74e-05 | 142 | 213 | 260 | 331 |  | 5.74e-05 | 130 | 213 | 260 | 331
5.76e-05 | 142 | 213 | 284 | 355 |  | 5.76e-05 | 142 | 213 | 260 | 331
5.78e-05 | 154 | 213 | 284 | 355 |  | 5.78e-05 | 142 | 213 | 284 | 355
5.8e-05 | 154 | 237 | 308 | 379 |  | 5.8e-05 | 154 | 237 | 308 | 379
5.82e-05 | 154 | 237 | 308 | 402 |  | 5.82e-05 | 154 | 237 | 308 | 379
5.84e-05 | 165 | 237 | 331 | 402 |  | 5.84e-05 | 165 | 237 | 308 | 402
5.86e-05 | 165 | 237 | 331 | 402 |  | 5.86e-05 | 165 | 237 | 331 | 402
5.88e-05 | 177 | 260 | 331 | 426 |  | 5.88e-05 | 177 | 260 | 331 | 426
5.9e-05 | 177 | 260 | 331 | 426 |  | 5.9e-05 | 177 | 260 | 355 | 426
5.92e-05 | 177 | 260 | 355 | 426 |  | 5.92e-05 | 177 | 260 | 355 | 426
5.94e-05 | 177 | 260 | 355 | 426 |  | 5.94e-05 | 177 | 260 | 355 | 426
5.96e-05 | 177 | 260 | 355 | 450 |  | 5.96e-05 | 177 | 260 | 355 | 450
5.98e-05 | 189 | 260 | 355 | 450 |  | 5.98e-05 | 189 | 284 | 355 | 450
6e-05 | 189 | 260 | 355 | 450 |  | 6e-05 | 189 | 260 | 355 | 450
6.02e-05 | 189 | 260 | 355 | 450 |  | 6.02e-05 | 189 | 284 | 355 | 450
6.04e-05 | 189 | 260 | 355 | 450 |  | 6.04e-05 | 189 | 284 | 355 | 450
6.06e-05 | 189 | 260 | 355 | 450 |  | 6.06e-05 | 189 | 284 | 379 | 450
6.08e-05 | 189 | 260 | 355 | 450 |  | 6.08e-05 | 189 | 284 | 355 | 450
6.1e-05 | 189 | 284 | 355 | 450 |  | 6.1e-05 | 189 | 260 | 355 | 450
6.12e-05 | 189 | 260 | 355 | 450 |  | 6.12e-05 | 189 | 284 | 355 | 450
6.14e-05 | 177 | 260 | 355 | 426 |  | 6.14e-05 | 189 | 260 | 355 | 450
6.16e-05 | 177 | 260 | 355 | 426 |  | 6.16e-05 | 189 | 260 | 355 | 426
6.18e-05 | 177 | 260 | 331 | 426 |  | 6.18e-05 | 177 | 260 | 355 | 426
6.2e-05 | 177 | 260 | 331 | 426 |  | 6.2e-05 | 177 | 260 | 331 | 426
6.22e-05 | 165 | 260 | 331 | 402 |  | 6.22e-05 | 177 | 260 | 331 | 426
6.24e-05 | 165 | 260 | 308 | 402 |  | 6.24e-05 | 177 | 260 | 331 | 402
6.26e-05 | 165 | 237 | 308 | 402 |  | 6.26e-05 | 165 | 260 | 331 | 402
6.28e-05 | 165 | 237 | 308 | 379 |  | 6.28e-05 | 165 | 237 | 308 | 379
6.3e-05 | 154 | 237 | 284 | 379 |  | 6.3e-05 | 165 | 237 | 308 | 379
6.32e-05 | 154 | 213 | 284 | 355 |  | 6.32e-05 | 154 | 237 | 284 | 379
6.34e-05 | 142 | 213 | 260 | 331 |  | 6.34e-05 | 154 | 213 | 284 | 355
6.36e-05 | 142 | 213 | 260 | 331 |  | 6.36e-05 | 142 | 213 | 284 | 331
6.38e-05 | 130 | 189 | 260 | 308 |  | 6.38e-05 | 130 | 213 | 260 | 331
6.4e-05 | 130 | 189 | 237 | 308 |  | 6.4e-05 | 130 | 189 | 260 | 308
6.42e-05 | 130 | 189 | 237 | 284 |  | 6.42e-05 | 130 | 189 | 237 | 308
6.44e-05 | 118 | 166 | 213 | 260 |  | 6.44e-05 | 118 | 189 | 237 | 284
6.46e-05 | 106 | 166 | 213 | 260 |  | 6.46e-05 | 118 | 166 | 213 | 260
6.48e-05 | 106 | 142 | 189 | 237 |  | 6.48e-05 | 106 | 166 | 213 | 260
6.5e-05 | 106 | 142 | 189 | 237 |  | 6.5e-05 | 106 | 142 | 189 | 237
6.52e-05 | 94 | 142 | 166 | 213 |  | 6.52e-05 | 94 | 142 | 189 | 213
6.54e-05 | 83 | 118 | 166 | 189 |  | 6.54e-05 | 94 | 118 | 166 | 213
6.56e-05 | 83 | 118 | 142 | 166 |  | 6.56e-05 | 83 | 118 | 166 | 189
6.58e-05 | 71 | 95 | 142 | 166 |  | 6.58e-05 | 71 | 118 | 142 | 166
6.6e-05 | 59 | 95 | 118 | 142 |  | 6.6e-05 | 71 | 95 | 118 | 142
6.62e-05 | 59 | 71 | 95 | 118 |  | 6.62e-05 | 59 | 95 | 118 | 142
6.64e-05 | 47 | 71 | 95 | 95 |  | 6.64e-05 | 59 | 71 | 95 | 118
6.66e-05 | 35 | 47 | 71 | 95 |  | 6.66e-05 | 35 | 47 | 71 | 95
6.68e-05 | 35 | 47 | 47 | 71 |  | 6.68e-05 | 35 | 47 | 71 | 71
6.7e-05 | 24 | 24 | 47 | 47 |  | 6.7e-05 | 35 | 47 | 47 | 71
6.72e-05 | 24 | 24 | 24 | 24 |  | 6.72e-05 | 24 | 24 | 24 | 47
6.74e-05 | 12 | 0 | 0 | 24 |  | 6.74e-05 | 12 | 24 | 24 | 24
6.76e-05 | 0 | 0 | 0 | 0 |  | 6.76e-05 | 0 | 0 | 0 | 0
6.78e-05 | 0 | 0 | -23 | -23 |  | 6.78e-05 | 0 | 0 | 0 | 0
6.8e-05 | -11 | -23 | -23 | -46 |  | 6.8e-05 | 0 | -23 | -23 | -23
6.82e-05 | -11 | -23 | -46 | -46 |  | 6.82e-05 | -11 | -23 | -23 | -46
6.84e-05 | -23 | -46 | -46 | -70 |  | 6.84e-05 | -23 | -46 | -46 | -46
6.86e-05 | -23 | -46 | -70 | -94 |  | 6.86e-05 | -34 | -46 | -70 | -70
6.88e-05 | -34 | -70 | -70 | -94 |  | 6.88e-05 | -34 | -46 | -70 | -94
6.9e-05 | -46 | -70 | -94 | -117 |  | 6.9e-05 | -34 | -70 | -94 | -117
6.92e-05 | -46 | -70 | -117 | -141 |  | 6.92e-05 | -46 | -70 | -94 | -117
6.94e-05 | -58 | -94 | -117 | -141 |  | 6.94e-05 | -58 | -94 | -117 | -141
6.96e-05 | -58 | -117 | -141 | -165 |  | 6.96e-05 | -58 | -94 | -117 | -165
6.98e-05 | -70 | -117 | -141 | -165 |  | 6.98e-05 | -70 | -94 | -141 | -165
7e-05 | -70 | -117 | -165 | -188 |  | 7e-05 | -70 | -117 | -141 | -188
7.02e-05 | -82 | -117 | -165 | -188 |  | 7.02e-05 | -82 | -117 | -165 | -188
7.04e-05 | -82 | -117 | -165 | -212 |  | 7.04e-05 | -82 | -117 | -165 | -212
7.06e-05 | -82 | -141 | -188 | -212 |  | 7.06e-05 | -82 | -141 | -188 | -212
7.08e-05 | -93 | -141 | -188 | -236 |  | 7.08e-05 | -93 | -141 | -188 | -236
7.1e-05 | -93 | -141 | -188 | -236 |  | 7.1e-05 | -93 | -141 | -188 | -236
7.12e-05 | -93 | -141 | -212 | -236 |  | 7.12e-05 | -93 | -141 | -188 | -236
7.14e-05 | -105 | -165 | -212 | -259 |  | 7.14e-05 | -105 | -165 | -188 | -259
7.16e-05 | -105 | -165 | -212 | -259 |  | 7.16e-05 | -105 | -165 | -212 | -259
7.18e-05 | -105 | -165 | -212 | -259 |  | 7.18e-05 | -105 | -165 | -212 | -259
7.2e-05 | -117 | -165 | -212 | -283 |  | 7.2e-05 | -117 | -165 | -212 | -283
7.22e-05 | -117 | -165 | -236 | -283 |  | 7.22e-05 | -117 | -165 | -236 | -283
7.24e-05 | -117 | -188 | -236 | -283 |  | 7.24e-05 | -117 | -188 | -236 | -283
7.26e-05 | -117 | -188 | -236 | -283 |  | 7.26e-05 | -117 | -188 | -236 | -283
7.28e-05 | -117 | -188 | -236 | -283 |  | 7.28e-05 | -117 | -188 | -236 | -307
7.3e-05 | -117 | -188 | -236 | -307 |  | 7.3e-05 | -117 | -188 | -236 | -307
7.32e-05 | -117 | -188 | -236 | -307 |  | 7.32e-05 | -129 | -188 | -259 | -307
7.34e-05 | -129 | -188 | -236 | -283 |  | 7.34e-05 | -129 | -188 | -259 | -307
7.36e-05 | -129 | -188 | -236 | -307 |  | 7.36e-05 | -129 | -188 | -236 | -307
7.38e-05 | -129 | -188 | -236 | -307 |  | 7.38e-05 | -129 | -188 | -259 | -307
7.4e-05 | -117 | -188 | -236 | -307 |  | 7.4e-05 | -129 | -188 | -236 | -307
7.42e-05 | -117 | -188 | -236 | -307 |  | 7.42e-05 | -129 | -188 | -259 | -307
7.44e-05 | -117 | -188 | -236 | -307 |  | 7.44e-05 | -129 | -188 | -236 | -307
7.46e-05 | -117 | -188 | -236 | -307 |  | 7.46e-05 | -117 | -188 | -236 | -307
7.48e-05 | -117 | -188 | -236 | -307 |  | 7.48e-05 | -129 | -188 | -236 | -283
7.5e-05 | -117 | -188 | -236 | -283 |  | 7.5e-05 | -117 | -188 | -236 | -283
7.52e-05 | -117 | -188 | -236 | -283 |  | 7.52e-05 | -117 | -188 | -236 | -307
7.54e-05 | -117 | -165 | -236 | -283 |  | 7.54e-05 | -117 | -188 | -236 | -283
7.56e-05 | -105 | -165 | -212 | -283 |  | 7.56e-05 | -117 | -165 | -236 | -283
7.58e-05 | -105 | -165 | -212 | -259 |  | 7.58e-05 | -117 | -165 | -236 | -283
7.6e-05 | -105 | -165 | -212 | -259 |  | 7.6e-05 | -105 | -165 | -212 | -259
7.62e-05 | -105 | -165 | -212 | -259 |  | 7.62e-05 | -105 | -165 | -212 | -259
7.64e-05 | -105 | -141 | -188 | -236 |  | 7.64e-05 | -105 | -165 | -212 | -259
7.66e-05 | -93 | -141 | -188 | -236 |  | 7.66e-05 | -105 | -165 | -188 | -259
7.68e-05 | -93 | -141 | -188 | -236 |  | 7.68e-05 | -93 | -141 | -188 | -236
7.7e-05 | -93 | -141 | -188 | -236 |  | 7.7e-05 | -93 | -141 | -188 | -236
7.72e-05 | -93 | -141 | -165 | -212 |  | 7.72e-05 | -93 | -141 | -188 | -212
7.74e-05 | -82 | -117 | -165 | -212 |  | 7.74e-05 | -93 | -141 | -165 | -212
7.76e-05 | -82 | -117 | -165 | -188 |  | 7.76e-05 | -82 | -141 | -165 | -212
7.78e-05 | -82 | -117 | -141 | -188 |  | 7.78e-05 | -82 | -117 | -165 | -188
7.8e-05 | -70 | -117 | -141 | -165 |  | 7.8e-05 | -82 | -117 | -141 | -188
7.82e-05 | -70 | -94 | -141 | -165 |  | 7.82e-05 | -70 | -117 | -141 | -165
7.84e-05 | -58 | -94 | -117 | -141 |  | 7.84e-05 | -70 | -117 | -141 | -165
7.86e-05 | -58 | -94 | -117 | -141 |  | 7.86e-05 | -58 | -94 | -117 | -165
7.88e-05 | -58 | -94 | -94 | -141 |  | 7.88e-05 | -58 | -94 | -117 | -141
7.9e-05 | -46 | -70 | -94 | -117 |  | 7.9e-05 | -58 | -94 | -117 | -141
7.92e-05 | -46 | -70 | -94 | -117 |  | 7.92e-05 | -46 | -70 | -94 | -117
7.94e-05 | -34 | -70 | -70 | -94 |  | 7.94e-05 | -46 | -70 | -94 | -117
7.96e-05 | -34 | -46 | -70 | -94 |  | 7.96e-05 | -46 | -70 | -70 | -94
7.98e-05 | -34 | -46 | -46 | -70 |  | 7.98e-05 | -34 | -46 | -70 | -94
8e-05 | -23 | -46 | -46 | -46 |  | 8e-05 | -34 | -46 | -70 | -70
8.02e-05 | -23 | -23 | -46 | -46 |  | 8.02e-05 | -23 | -46 | -46 | -70
8.04e-05 | -11 | -23 | -23 | -46 |  | 8.04e-05 | -23 | -23 | -46 | -46
8.06e-05 | -11 | -23 | -23 | -23 |  | 8.06e-05 | -11 | -23 | -23 | -46
8.08e-05 | -11 | 0 | -23 | -23 |  | 8.08e-05 | -11 | -23 | -23 | -23
8.1e-05 | 0 | 0 | 0 | 0 |  | 8.1e-05 | -11 | 0 | -23 | -23
8.12e-05 | 0 | 0 | 0 | 0 |  | 8.12e-05 | 0 | 0 | 0 | 0
8.14e-05 | 0 | 0 | 24 | 24 |  | 8.14e-05 | 0 | 0 | 0 | 0
8.16e-05 | 12 | 0 | 24 | 24 |  | 8.16e-05 | 0 | 0 | 0 | 0
8.18e-05 | 12 | 24 | 24 | 24 |  | 8.18e-05 | 12 | 24 | 24 | 24
8.2e-05 | 24 | 24 | 47 | 47 |  | 8.2e-05 | 12 | 24 | 24 | 47
8.22e-05 | 24 | 24 | 47 | 47 |  | 8.22e-05 | 12 | 24 | 24 | 47
8.24e-05 | 24 | 47 | 47 | 71 |  | 8.24e-05 | 24 | 24 | 47 | 47
8.26e-05 | 35 | 47 | 47 | 71 |  | 8.26e-05 | 24 | 47 | 47 | 71
8.28e-05 | 35 | 47 | 71 | 71 |  | 8.28e-05 | 24 | 47 | 71 | 71
8.3e-05 | 35 | 47 | 71 | 95 |  | 8.3e-05 | 35 | 47 | 71 | 71
8.32e-05 | 47 | 47 | 71 | 95 |  | 8.32e-05 | 35 | 47 | 71 | 95
8.34e-05 | 47 | 71 | 95 | 95 |  | 8.34e-05 | 35 | 71 | 71 | 118
8.36e-05 | 47 | 71 | 95 | 118 |  | 8.36e-05 | 47 | 71 | 95 | 118
8.38e-05 | 47 | 71 | 95 | 118 |  | 8.38e-05 | 47 | 71 | 95 | 118
8.4e-05 | 59 | 71 | 95 | 142 |  | 8.4e-05 | 47 | 71 | 95 | 118
8.42e-05 | 59 | 71 | 118 | 142 |  | 8.42e-05 | 59 | 71 | 118 | 142
8.44e-05 | 59 | 95 | 118 | 142 |  | 8.44e-05 | 59 | 71 | 95 | 142
8.46e-05 | 59 | 95 | 118 | 142 |  | 8.46e-05 | 59 | 95 | 118 | 142
8.48e-05 | 71 | 95 | 118 | 142 |  | 8.48e-05 | 59 | 95 | 118 | 142
8.5e-05 | 71 | 95 | 118 | 166 |  | 8.5e-05 | 71 | 95 | 118 | 142
8.52e-05 | 71 | 95 | 118 | 166 |  | 8.52e-05 | 71 | 95 | 118 | 166
8.54e-05 | 71 | 95 | 142 | 166 |  | 8.54e-05 | 71 | 95 | 118 | 166
8.56e-05 | 71 | 95 | 142 | 166 |  | 8.56e-05 | 71 | 95 | 142 | 166
8.58e-05 | 71 | 118 | 142 | 166 |  | 8.58e-05 | 71 | 95 | 142 | 166
8.6e-05 | 71 | 118 | 142 | 166 |  | 8.6e-05 | 71 | 95 | 142 | 166
8.62e-05 | 71 | 118 | 142 | 166 |  | 8.62e-05 | 71 | 118 | 142 | 189
8.64e-05 | 71 | 118 | 142 | 189 |  | 8.64e-05 | 71 | 95 | 142 | 189
8.66e-05 | 71 | 118 | 142 | 189 |  | 8.66e-05 | 71 | 118 | 142 | 189
8.68e-05 | 71 | 118 | 142 | 189 |  | 8.68e-05 | 83 | 95 | 142 | 189
8.7e-05 | 71 | 118 | 142 | 189 |  | 8.7e-05 | 71 | 118 | 142 | 189
8.72e-05 | 71 | 118 | 142 | 189 |  | 8.72e-05 | 71 | 118 | 142 | 189
8.74e-05 | 71 | 118 | 142 | 189 |  | 8.74e-05 | 71 | 118 | 142 | 189
8.76e-05 | 83 | 118 | 142 | 189 |  | 8.76e-05 | 71 | 118 | 142 | 189
8.78e-05 | 71 | 118 | 142 | 189 |  | 8.78e-05 | 71 | 118 | 142 | 189
8.8e-05 | 71 | 118 | 142 | 166 |  | 8.8e-05 | 71 | 118 | 142 | 189
8.82e-05 | 71 | 95 | 142 | 166 |  | 8.82e-05 | 71 | 118 | 142 | 166
8.84e-05 | 71 | 95 | 142 | 166 |  | 8.84e-05 | 71 | 118 | 142 | 189
8.86e-05 | 71 | 95 | 142 | 166 |  | 8.86e-05 | 71 | 118 | 142 | 166
8.88e-05 | 71 | 95 | 142 | 166 |  | 8.88e-05 | 71 | 118 | 142 | 166
8.9e-05 | 71 | 95 | 142 | 166 |  | 8.9e-05 | 71 | 95 | 142 | 166
8.92e-05 | 71 | 95 | 118 | 166 |  | 8.92e-05 | 71 | 95 | 118 | 166
8.94e-05 | 71 | 95 | 118 | 166 |  | 8.94e-05 | 71 | 95 | 142 | 166
8.96e-05 | 59 | 95 | 118 | 142 |  | 8.96e-05 | 71 | 95 | 142 | 166
8.98e-05 | 59 | 95 | 118 | 142 |  | 8.98e-05 | 71 | 95 | 118 | 166
9e-05 | 59 | 95 | 118 | 142 |  | 9e-05 | 59 | 95 | 118 | 142
9.02e-05 | 59 | 95 | 118 | 142 |  | 9.02e-05 | 59 | 95 | 118 | 142
9.04e-05 | 59 | 71 | 118 | 118 |  | 9.04e-05 | 59 | 71 | 118 | 142
9.06e-05 | 59 | 71 | 95 | 118 |  | 9.06e-05 | 59 | 71 | 118 | 142
9.08e-05 | 47 | 71 | 95 | 118 |  | 9.08e-05 | 47 | 71 | 118 | 118
9.1e-05 | 47 | 71 | 95 | 118 |  | 9.1e-05 | 59 | 71 | 95 | 118
9.12e-05 | 47 | 71 | 95 | 118 |  | 9.12e-05 | 47 | 71 | 95 | 118
9.14e-05 | 47 | 71 | 71 | 95 |  | 9.14e-05 | 47 | 71 | 95 | 118
9.16e-05 | 47 | 71 | 71 | 95 |  | 9.16e-05 | 47 | 71 | 95 | 118
9.18e-05 | 35 | 47 | 71 | 95 |  | 9.18e-05 | 47 | 47 | 71 | 95
9.2e-05 | 35 | 47 | 71 | 95 |  | 9.2e-05 | 35 | 47 | 71 | 95
9.22e-05 | 35 | 47 | 47 | 71 |  | 9.22e-05 | 35 | 47 | 71 | 95
9.24e-05 | 35 | 47 | 47 | 71 |  | 9.24e-05 | 35 | 47 | 71 | 71
9.26e-05 | 24 | 47 | 47 | 47 |  | 9.26e-05 | 24 | 47 | 47 | 71
9.28e-05 | 24 | 24 | 47 | 47 |  | 9.28e-05 | 24 | 47 | 47 | 71
9.3e-05 | 24 | 24 | 47 | 47 |  | 9.3e-05 | 24 | 24 | 47 | 47
9.32e-05 | 24 | 24 | 47 | 47 |  | 9.32e-05 | 24 | 24 | 47 | 47
9.34e-05 | 24 | 24 | 24 | 24 |  | 9.34e-05 | 24 | 24 | 24 | 47
9.36e-05 | 12 | 24 | 24 | 24 |  | 9.36e-05 | 12 | 24 | 24 | 47
9.38e-05 | 12 | 24 | 24 | 24 |  | 9.38e-05 | 12 | 24 | 24 | 24
9.4e-05 | 12 | 0 | 0 | 24 |  | 9.4e-05 | 12 | 0 | 24 | 24
9.42e-05 | 12 | 0 | 0 | 0 |  | 9.42e-05 | 12 | 0 | 0 | 0
9.44e-05 | 0 | 0 | 0 | 0 |  | 9.44e-05 | 0 | 0 | 0 | 0
9.46e-05 | 0 | 0 | 0 | 0 |  | 9.46e-05 | 0 | 0 | 0 | 0
9.48e-05 | 0 | 0 | 0 | -23 |  | 9.48e-05 | 0 | 0 | 0 | 0
9.5e-05 | 0 | 0 | -23 | -23 |  | 9.5e-05 | 0 | 0 | -23 | 0
9.52e-05 | -11 | -23 | -23 | -23 |  | 9.52e-05 | -11 | -23 | -23 | -23
9.54e-05 | -11 | -23 | -23 | -23 |  | 9.54e-05 | -11 | -23 | -23 | -23
9.56e-05 | -11 | -23 | -23 | -46 |  | 9.56e-05 | -11 | -23 | -23 | -46
9.58e-05 | -23 | -23 | -23 | -46 |  | 9.58e-05 | -11 | -23 | -23 | -46
9.6e-05 | -11 | -23 | -46 | -46 |  | 9.6e-05 | -11 | -23 | -46 | -46
9.62e-05 | -23 | -23 | -46 | -46 |  | 9.62e-05 | -23 | -46 | -46 | -46
9.64e-05 | -23 | -23 | -46 | -70 |  | 9.64e-05 | -23 | -46 | -46 | -46
9.66e-05 | -23 | -46 | -46 | -70 |  | 9.66e-05 | -23 | -46 | -46 | -70
9.68e-05 | -23 | -46 | -46 | -70 |  | 9.68e-05 | -23 | -46 | -46 | -70
9.7e-05 | -23 | -46 | -70 | -70 |  | 9.7e-05 | -23 | -46 | -70 | -70
9.72e-05 | -34 | -46 | -70 | -70 |  | 9.72e-05 | -34 | -46 | -70 | -70
9.74e-05 | -34 | -46 | -70 | -94 |  | 9.74e-05 | -34 | -46 | -70 | -70
9.76e-05 | -34 | -46 | -70 | -94 |  | 9.76e-05 | -34 | -46 | -70 | -94
9.78e-05 | -34 | -70 | -70 | -94 |  | 9.78e-05 | -34 | -70 | -70 | -94
9.8e-05 | -34 | -70 | -70 | -94 |  | 9.8e-05 | -34 | -70 | -70 | -94
9.82e-05 | -34 | -70 | -94 | -94 |  | 9.82e-05 | -34 | -70 | -70 | -94
9.84e-05 | -34 | -70 | -94 | -94 |  | 9.84e-05 | -46 | -70 | -94 | -94
9.86e-05 | -34 | -70 | -94 | -117 |  | 9.86e-05 | -46 | -70 | -94 | -117
9.88e-05 | -46 | -70 | -94 | -117 |  | 9.88e-05 | -46 | -70 | -94 | -117
9.9e-05 | -34 | -70 | -94 | -117 |  | 9.9e-05 | -46 | -70 | -94 | -117
9.92e-05 | -46 | -70 | -94 | -117 |  | 9.92e-05 | -46 | -70 | -94 | -117
9.94e-05 | -46 | -70 | -94 | -117 |  | 9.94e-05 | -46 | -70 | -94 | -117
9.96e-05 | -46 | -70 | -94 | -117 |  | 9.96e-05 | -46 | -70 | -94 | -117
9.98e-05 | -46 | -70 | -94 | -117 |  | 9.98e-05 | -46 | -70 | -94 | -117
0.0001 | -46 | -70 | -94 | -117 |  | 0.0001 | -46 | -70 | -94 | -117
0.0001002 | -46 | -70 | -94 | -117 |  | 0.0001002 | -46 | -70 | -94 | -117
0.0001004 | -46 | -70 | -94 | -117 |  | 0.0001004 | -46 | -70 | -94 | -117
0.0001006 | -46 | -70 | -94 | -117 |  | 0.0001006 | -46 | -70 | -94 | -117
0.0001008 | -46 | -70 | -94 | -117 |  | 0.0001008 | -46 | -70 | -94 | -117
0.000101 | -46 | -70 | -94 | -117 |  | 0.000101 | -46 | -70 | -94 | -117
0.0001012 | -46 | -70 | -94 | -117 |  | 0.0001012 | -46 | -70 | -94 | -117
0.0001014 | -46 | -70 | -94 | -117 |  | 0.0001014 | -46 | -70 | -94 | -117
0.0001016 | -46 | -70 | -94 | -117 |  | 0.0001016 | -46 | -70 | -94 | -117
0.0001018 | -46 | -70 | -94 | -117 |  | 0.0001018 | -46 | -70 | -94 | -117
0.000102 | -46 | -70 | -94 | -117 |  | 0.000102 | -46 | -70 | -94 | -117
0.0001022 | -46 | -70 | -94 | -94 |  | 0.0001022 | -46 | -70 | -94 | -117
| 2 | 2 | 2 | 2 |  |  | 2 | 2 | 2 | 2
| 1000 | 1500 | 2000 | 2500 |  |  | 1000 | 1500 | 2000 | 2500
| 2 | 2 | 2 | 2 |  |  | 2 | 2 | 2 | 2

## Step 2 159 Impulse
Time (s) | 1.0 kV | 1.5 kV | 2.0 kV | 2.5 kV | 159 | Time (s) | 1.0 kV | 1.5 kV | 2.0 kV | 2.5 kV |  | Notes
0 | 0 | 0 | 0 | 0 | A12b Coil Pack QXFA 159 Outer to Inner | 0 | 0 | 0 | 0 | 0 | A12b Coil Pack QXFA 159 Inner to Outer
2e-07 | 0 | 0 | 0 | 0 |  | 2e-07 | 0 | 0 | 0 | 0
4e-07 | 0 | 0 | 0 | 0 |  | 4e-07 | 0 | 0 | 0 | 0
6e-07 | 165 | 308 | 426 | 592 |  | 6e-07 | 154 | 308 | 450 | 592
8e-07 | 626 | 1065 | 1514 | 2035 |  | 8e-07 | 626 | 1065 | 1514 | 2035
1e-06 | 980 | 1514 | 2011 | 2484 |  | 1e-06 | 980 | 1514 | 2011 | 2484
1.2e-06 | 945 | 1372 | 1751 | 2130 |  | 1.2e-06 | 945 | 1372 | 1775 | 2130
1.4e-06 | 850 | 1254 | 1633 | 2011 |  | 1.4e-06 | 850 | 1254 | 1633 | 2011
1.6e-06 | 850 | 1278 | 1680 | 2059 |  | 1.6e-06 | 850 | 1278 | 1680 | 2059
1.8e-06 | 850 | 1278 | 1656 | 2035 |  | 1.8e-06 | 850 | 1278 | 1656 | 2035
2e-06 | 827 | 1230 | 1609 | 1964 |  | 2e-06 | 827 | 1230 | 1609 | 1964
2.2e-06 | 815 | 1207 | 1585 | 1940 |  | 2.2e-06 | 815 | 1207 | 1585 | 1940
2.4e-06 | 803 | 1183 | 1562 | 1917 |  | 2.4e-06 | 803 | 1183 | 1562 | 1917
2.6e-06 | 791 | 1183 | 1538 | 1893 |  | 2.6e-06 | 791 | 1183 | 1538 | 1893
2.8e-06 | 780 | 1159 | 1514 | 1846 |  | 2.8e-06 | 780 | 1159 | 1514 | 1869
3e-06 | 768 | 1136 | 1491 | 1822 |  | 3e-06 | 768 | 1136 | 1491 | 1822
3.2e-06 | 744 | 1112 | 1443 | 1775 |  | 3.2e-06 | 756 | 1112 | 1443 | 1775
3.4e-06 | 732 | 1088 | 1396 | 1727 |  | 3.4e-06 | 732 | 1088 | 1396 | 1751
3.6e-06 | 709 | 1041 | 1372 | 1680 |  | 3.6e-06 | 709 | 1065 | 1372 | 1680
3.8e-06 | 697 | 1017 | 1349 | 1656 |  | 3.8e-06 | 685 | 1017 | 1349 | 1656
4e-06 | 673 | 994 | 1301 | 1609 |  | 4e-06 | 673 | 994 | 1301 | 1609
4.2e-06 | 661 | 970 | 1278 | 1585 |  | 4.2e-06 | 661 | 994 | 1278 | 1585
4.4e-06 | 650 | 970 | 1254 | 1538 |  | 4.4e-06 | 650 | 970 | 1254 | 1562
4.6e-06 | 650 | 946 | 1230 | 1514 |  | 4.6e-06 | 638 | 946 | 1254 | 1538
4.8e-06 | 626 | 923 | 1207 | 1491 |  | 4.8e-06 | 638 | 923 | 1207 | 1514
5e-06 | 626 | 923 | 1183 | 1467 |  | 5e-06 | 614 | 923 | 1183 | 1467
5.2e-06 | 602 | 899 | 1159 | 1443 |  | 5.2e-06 | 602 | 899 | 1159 | 1443
5.4e-06 | 591 | 875 | 1136 | 1396 |  | 5.4e-06 | 591 | 875 | 1136 | 1420
5.6e-06 | 579 | 852 | 1112 | 1349 |  | 5.6e-06 | 579 | 852 | 1112 | 1372
5.8e-06 | 555 | 828 | 1065 | 1325 |  | 5.8e-06 | 555 | 828 | 1088 | 1325
6e-06 | 543 | 804 | 1041 | 1278 |  | 6e-06 | 543 | 804 | 1041 | 1278
6.2e-06 | 531 | 781 | 1017 | 1254 |  | 6.2e-06 | 531 | 781 | 1017 | 1254
6.4e-06 | 508 | 757 | 970 | 1183 |  | 6.4e-06 | 508 | 757 | 970 | 1207
6.6e-06 | 484 | 710 | 946 | 1136 |  | 6.6e-06 | 484 | 710 | 946 | 1159
6.8e-06 | 472 | 686 | 899 | 1088 |  | 6.8e-06 | 472 | 686 | 899 | 1112
7e-06 | 449 | 663 | 852 | 1041 |  | 7e-06 | 449 | 663 | 852 | 1065
7.2e-06 | 425 | 615 | 828 | 994 |  | 7.2e-06 | 425 | 615 | 828 | 994
7.4e-06 | 402 | 592 | 781 | 946 |  | 7.4e-06 | 402 | 592 | 781 | 946
7.6e-06 | 378 | 568 | 734 | 899 |  | 7.6e-06 | 378 | 568 | 734 | 899
7.8e-06 | 366 | 521 | 686 | 828 |  | 7.8e-06 | 366 | 521 | 686 | 852
8e-06 | 331 | 497 | 639 | 781 |  | 8e-06 | 343 | 497 | 639 | 781
8.2e-06 | 307 | 450 | 592 | 710 |  | 8.2e-06 | 307 | 450 | 592 | 734
8.4e-06 | 283 | 426 | 544 | 663 |  | 8.4e-06 | 283 | 426 | 544 | 663
8.6e-06 | 260 | 379 | 497 | 592 |  | 8.6e-06 | 260 | 379 | 497 | 592
8.8e-06 | 236 | 331 | 450 | 544 |  | 8.8e-06 | 236 | 331 | 450 | 544
9e-06 | 213 | 308 | 402 | 497 |  | 9e-06 | 213 | 308 | 402 | 497
9.2e-06 | 189 | 260 | 331 | 426 |  | 9.2e-06 | 189 | 260 | 331 | 426
9.4e-06 | 165 | 237 | 308 | 355 |  | 9.4e-06 | 165 | 237 | 308 | 355
9.6e-06 | 130 | 189 | 260 | 308 |  | 9.6e-06 | 142 | 189 | 237 | 308
9.8e-06 | 118 | 166 | 213 | 260 |  | 9.8e-06 | 118 | 166 | 213 | 237
1e-05 | 83 | 118 | 166 | 189 |  | 1e-05 | 94 | 142 | 166 | 189
1.02e-05 | 71 | 71 | 95 | 142 |  | 1.02e-05 | 71 | 71 | 118 | 142
1.04e-05 | 35 | 47 | 71 | 71 |  | 1.04e-05 | 35 | 47 | 71 | 71
1.06e-05 | 12 | 24 | 0 | 24 |  | 1.06e-05 | 12 | 24 | 24 | 24
1.08e-05 | -11 | -23 | -23 | -46 |  | 1.08e-05 | 0 | -23 | -23 | -46
1.1e-05 | -34 | -46 | -70 | -94 |  | 1.1e-05 | -23 | -46 | -70 | -94
1.12e-05 | -58 | -94 | -117 | -141 |  | 1.12e-05 | -46 | -70 | -117 | -165
1.14e-05 | -82 | -117 | -165 | -212 |  | 1.14e-05 | -82 | -117 | -165 | -212
1.16e-05 | -93 | -165 | -212 | -259 |  | 1.16e-05 | -93 | -165 | -212 | -259
1.18e-05 | -117 | -188 | -259 | -307 |  | 1.18e-05 | -117 | -188 | -259 | -307
1.2e-05 | -141 | -236 | -283 | -378 |  | 1.2e-05 | -141 | -212 | -307 | -354
1.22e-05 | -164 | -259 | -330 | -425 |  | 1.22e-05 | -164 | -259 | -330 | -425
1.24e-05 | -188 | -283 | -378 | -472 |  | 1.24e-05 | -188 | -283 | -378 | -472
1.26e-05 | -200 | -307 | -425 | -520 |  | 1.26e-05 | -200 | -307 | -425 | -520
1.28e-05 | -223 | -354 | -449 | -567 |  | 1.28e-05 | -223 | -354 | -449 | -567
1.3e-05 | -247 | -378 | -496 | -614 |  | 1.3e-05 | -247 | -378 | -496 | -614
1.32e-05 | -271 | -401 | -543 | -662 |  | 1.32e-05 | -271 | -401 | -543 | -662
1.34e-05 | -282 | -425 | -567 | -709 |  | 1.34e-05 | -282 | -425 | -567 | -709
1.36e-05 | -306 | -449 | -614 | -756 |  | 1.36e-05 | -306 | -472 | -614 | -756
1.38e-05 | -330 | -496 | -662 | -803 |  | 1.38e-05 | -330 | -496 | -638 | -803
1.4e-05 | -342 | -520 | -685 | -851 |  | 1.4e-05 | -342 | -520 | -685 | -851
1.42e-05 | -365 | -543 | -733 | -898 |  | 1.42e-05 | -365 | -543 | -733 | -898
1.44e-05 | -377 | -567 | -756 | -945 |  | 1.44e-05 | -389 | -567 | -756 | -945
1.46e-05 | -401 | -614 | -803 | -993 |  | 1.46e-05 | -401 | -614 | -803 | -993
1.48e-05 | -424 | -638 | -827 | -1016 |  | 1.48e-05 | -412 | -638 | -827 | -1040
1.5e-05 | -436 | -662 | -874 | -1064 |  | 1.5e-05 | -436 | -662 | -874 | -1064
1.52e-05 | -448 | -685 | -898 | -1111 |  | 1.52e-05 | -448 | -685 | -898 | -1111
1.54e-05 | -471 | -709 | -922 | -1135 |  | 1.54e-05 | -471 | -709 | -922 | -1158
1.56e-05 | -483 | -733 | -969 | -1182 |  | 1.56e-05 | -483 | -733 | -969 | -1182
1.58e-05 | -507 | -756 | -993 | -1229 |  | 1.58e-05 | -495 | -756 | -993 | -1229
1.6e-05 | -519 | -780 | -1016 | -1253 |  | 1.6e-05 | -507 | -780 | -1016 | -1253
1.62e-05 | -530 | -803 | -1040 | -1300 |  | 1.62e-05 | -530 | -803 | -1040 | -1300
1.64e-05 | -542 | -827 | -1064 | -1324 |  | 1.64e-05 | -542 | -827 | -1064 | -1324
1.66e-05 | -554 | -827 | -1111 | -1348 |  | 1.66e-05 | -554 | -827 | -1087 | -1348
1.68e-05 | -566 | -874 | -1135 | -1371 |  | 1.68e-05 | -566 | -851 | -1111 | -1395
1.7e-05 | -578 | -874 | -1135 | -1395 |  | 1.7e-05 | -578 | -874 | -1135 | -1419
1.72e-05 | -590 | -898 | -1158 | -1442 |  | 1.72e-05 | -590 | -898 | -1158 | -1442
1.74e-05 | -601 | -898 | -1182 | -1466 |  | 1.74e-05 | -601 | -898 | -1182 | -1466
1.76e-05 | -613 | -922 | -1206 | -1490 |  | 1.76e-05 | -613 | -922 | -1206 | -1490
1.78e-05 | -625 | -945 | -1206 | -1513 |  | 1.78e-05 | -625 | -945 | -1229 | -1513
1.8e-05 | -637 | -945 | -1253 | -1537 |  | 1.8e-05 | -637 | -945 | -1253 | -1537
1.82e-05 | -637 | -969 | -1253 | -1537 |  | 1.82e-05 | -637 | -969 | -1253 | -1537
1.84e-05 | -649 | -969 | -1277 | -1561 |  | 1.84e-05 | -649 | -969 | -1277 | -1561
1.86e-05 | -660 | -969 | -1300 | -1584 |  | 1.86e-05 | -660 | -993 | -1277 | -1584
1.88e-05 | -660 | -993 | -1300 | -1584 |  | 1.88e-05 | -660 | -993 | -1300 | -1584
1.9e-05 | -660 | -993 | -1300 | -1608 |  | 1.9e-05 | -660 | -993 | -1300 | -1608
1.92e-05 | -672 | -993 | -1324 | -1632 |  | 1.92e-05 | -672 | -1016 | -1324 | -1632
1.94e-05 | -672 | -1016 | -1324 | -1632 |  | 1.94e-05 | -672 | -1016 | -1324 | -1632
1.96e-05 | -672 | -1016 | -1324 | -1632 |  | 1.96e-05 | -672 | -1016 | -1324 | -1655
1.98e-05 | -684 | -1016 | -1324 | -1632 |  | 1.98e-05 | -684 | -1016 | -1324 | -1655
2e-05 | -684 | -1016 | -1348 | -1655 |  | 2e-05 | -684 | -1016 | -1348 | -1655
2.02e-05 | -684 | -1016 | -1348 | -1655 |  | 2.02e-05 | -684 | -1016 | -1348 | -1655
2.04e-05 | -684 | -1016 | -1348 | -1655 |  | 2.04e-05 | -684 | -1040 | -1348 | -1655
2.06e-05 | -684 | -1040 | -1348 | -1655 |  | 2.06e-05 | -684 | -1016 | -1348 | -1655
2.08e-05 | -684 | -1016 | -1324 | -1655 |  | 2.08e-05 | -684 | -1016 | -1348 | -1655
2.1e-05 | -684 | -1016 | -1324 | -1632 |  | 2.1e-05 | -684 | -1016 | -1324 | -1655
2.12e-05 | -672 | -1016 | -1324 | -1632 |  | 2.12e-05 | -672 | -1016 | -1324 | -1632
2.14e-05 | -672 | -993 | -1324 | -1608 |  | 2.14e-05 | -660 | -993 | -1324 | -1632
2.16e-05 | -660 | -993 | -1300 | -1584 |  | 2.16e-05 | -660 | -993 | -1300 | -1608
2.18e-05 | -660 | -993 | -1277 | -1561 |  | 2.18e-05 | -660 | -993 | -1277 | -1584
2.2e-05 | -649 | -969 | -1253 | -1537 |  | 2.2e-05 | -649 | -969 | -1253 | -1561
2.22e-05 | -637 | -945 | -1229 | -1513 |  | 2.22e-05 | -637 | -945 | -1229 | -1513
2.24e-05 | -625 | -922 | -1206 | -1490 |  | 2.24e-05 | -625 | -922 | -1206 | -1490
2.26e-05 | -601 | -898 | -1182 | -1442 |  | 2.26e-05 | -613 | -922 | -1182 | -1466
2.28e-05 | -590 | -874 | -1158 | -1395 |  | 2.28e-05 | -590 | -898 | -1158 | -1419
2.3e-05 | -578 | -874 | -1111 | -1371 |  | 2.3e-05 | -578 | -874 | -1135 | -1371
2.32e-05 | -566 | -827 | -1087 | -1324 |  | 2.32e-05 | -566 | -827 | -1087 | -1348
2.34e-05 | -542 | -803 | -1040 | -1300 |  | 2.34e-05 | -542 | -803 | -1064 | -1300
2.36e-05 | -519 | -780 | -1016 | -1253 |  | 2.36e-05 | -530 | -780 | -1016 | -1253
2.38e-05 | -507 | -756 | -969 | -1182 |  | 2.38e-05 | -507 | -756 | -993 | -1206
2.4e-05 | -483 | -709 | -945 | -1135 |  | 2.4e-05 | -483 | -733 | -945 | -1158
2.42e-05 | -460 | -685 | -898 | -1087 |  | 2.42e-05 | -471 | -685 | -898 | -1111
2.44e-05 | -448 | -662 | -851 | -1040 |  | 2.44e-05 | -448 | -662 | -874 | -1064
2.46e-05 | -412 | -614 | -803 | -993 |  | 2.46e-05 | -424 | -614 | -803 | -993
2.48e-05 | -401 | -591 | -756 | -945 |  | 2.48e-05 | -389 | -591 | -780 | -945
2.5e-05 | -377 | -543 | -709 | -874 |  | 2.5e-05 | -377 | -567 | -709 | -898
2.52e-05 | -353 | -520 | -662 | -803 |  | 2.52e-05 | -353 | -520 | -685 | -827
2.54e-05 | -318 | -496 | -614 | -756 |  | 2.54e-05 | -330 | -496 | -614 | -756
2.56e-05 | -294 | -449 | -567 | -685 |  | 2.56e-05 | -306 | -449 | -567 | -709
2.58e-05 | -271 | -401 | -520 | -638 |  | 2.58e-05 | -271 | -401 | -520 | -638
2.6e-05 | -247 | -354 | -472 | -567 |  | 2.6e-05 | -247 | -378 | -472 | -591
2.62e-05 | -223 | -330 | -425 | -496 |  | 2.62e-05 | -223 | -330 | -425 | -520
2.64e-05 | -188 | -283 | -354 | -449 |  | 2.64e-05 | -200 | -283 | -378 | -449
2.66e-05 | -164 | -236 | -307 | -378 |  | 2.66e-05 | -176 | -259 | -330 | -378
2.68e-05 | -141 | -212 | -259 | -307 |  | 2.68e-05 | -153 | -212 | -283 | -330
2.7e-05 | -117 | -165 | -212 | -259 |  | 2.7e-05 | -117 | -165 | -212 | -259
2.72e-05 | -93 | -141 | -165 | -188 |  | 2.72e-05 | -93 | -141 | -165 | -188
2.74e-05 | -58 | -94 | -117 | -117 |  | 2.74e-05 | -70 | -94 | -117 | -141
2.76e-05 | -34 | -46 | -70 | -70 |  | 2.76e-05 | -46 | -46 | -70 | -70
2.78e-05 | -11 | 0 | 0 | 0 |  | 2.78e-05 | -11 | -23 | -23 | 0
2.8e-05 | 12 | 24 | 47 | 71 |  | 2.8e-05 | 12 | 24 | 24 | 47
2.82e-05 | 47 | 71 | 95 | 118 |  | 2.82e-05 | 35 | 71 | 95 | 118
2.84e-05 | 71 | 118 | 142 | 189 |  | 2.84e-05 | 59 | 95 | 142 | 166
2.86e-05 | 94 | 142 | 189 | 237 |  | 2.86e-05 | 83 | 142 | 189 | 237
2.88e-05 | 118 | 166 | 237 | 308 |  | 2.88e-05 | 106 | 166 | 237 | 284
2.9e-05 | 142 | 213 | 284 | 355 |  | 2.9e-05 | 130 | 213 | 260 | 355
2.92e-05 | 165 | 237 | 331 | 402 |  | 2.92e-05 | 165 | 237 | 331 | 402
2.94e-05 | 189 | 284 | 379 | 450 |  | 2.94e-05 | 177 | 260 | 355 | 450
2.96e-05 | 213 | 308 | 402 | 521 |  | 2.96e-05 | 201 | 308 | 402 | 521
2.98e-05 | 224 | 331 | 450 | 568 |  | 2.98e-05 | 213 | 331 | 450 | 544
3e-05 | 248 | 379 | 497 | 615 |  | 3e-05 | 248 | 355 | 497 | 615
3.02e-05 | 272 | 402 | 521 | 663 |  | 3.02e-05 | 260 | 402 | 521 | 639
3.04e-05 | 283 | 426 | 568 | 710 |  | 3.04e-05 | 283 | 426 | 544 | 686
3.06e-05 | 307 | 450 | 592 | 734 |  | 3.06e-05 | 295 | 450 | 592 | 734
3.08e-05 | 319 | 473 | 615 | 781 |  | 3.08e-05 | 319 | 473 | 615 | 781
3.1e-05 | 331 | 497 | 663 | 804 |  | 3.1e-05 | 331 | 497 | 639 | 804
3.12e-05 | 343 | 521 | 686 | 852 |  | 3.12e-05 | 343 | 521 | 686 | 852
3.14e-05 | 366 | 544 | 710 | 875 |  | 3.14e-05 | 354 | 521 | 710 | 875
3.16e-05 | 378 | 568 | 734 | 899 |  | 3.16e-05 | 366 | 544 | 734 | 923
3.18e-05 | 390 | 568 | 757 | 946 |  | 3.18e-05 | 378 | 568 | 757 | 946
3.2e-05 | 402 | 592 | 781 | 946 |  | 3.2e-05 | 390 | 592 | 781 | 970
3.22e-05 | 413 | 615 | 804 | 994 |  | 3.22e-05 | 402 | 615 | 804 | 994
3.24e-05 | 413 | 615 | 804 | 994 |  | 3.24e-05 | 413 | 615 | 804 | 1017
3.26e-05 | 425 | 639 | 828 | 1017 |  | 3.26e-05 | 413 | 615 | 828 | 1017
3.28e-05 | 437 | 639 | 828 | 1041 |  | 3.28e-05 | 425 | 639 | 852 | 1041
3.3e-05 | 437 | 639 | 852 | 1041 |  | 3.3e-05 | 437 | 639 | 852 | 1065
3.32e-05 | 437 | 663 | 875 | 1065 |  | 3.32e-05 | 437 | 663 | 875 | 1065
3.34e-05 | 449 | 663 | 875 | 1065 |  | 3.34e-05 | 449 | 663 | 875 | 1065
3.36e-05 | 449 | 663 | 875 | 1065 |  | 3.36e-05 | 449 | 663 | 875 | 1088
3.38e-05 | 449 | 663 | 875 | 1065 |  | 3.38e-05 | 449 | 663 | 875 | 1088
3.4e-05 | 449 | 663 | 875 | 1065 |  | 3.4e-05 | 449 | 663 | 875 | 1088
3.42e-05 | 449 | 663 | 875 | 1065 |  | 3.42e-05 | 449 | 663 | 875 | 1088
3.44e-05 | 449 | 663 | 875 | 1065 |  | 3.44e-05 | 449 | 663 | 875 | 1065
3.46e-05 | 449 | 663 | 875 | 1065 |  | 3.46e-05 | 449 | 663 | 875 | 1065
3.48e-05 | 437 | 639 | 852 | 1041 |  | 3.48e-05 | 437 | 663 | 852 | 1065
3.5e-05 | 437 | 639 | 828 | 1041 |  | 3.5e-05 | 437 | 639 | 852 | 1041
3.52e-05 | 437 | 639 | 828 | 1017 |  | 3.52e-05 | 437 | 639 | 828 | 1041
3.54e-05 | 425 | 615 | 828 | 994 |  | 3.54e-05 | 425 | 615 | 828 | 1017
3.56e-05 | 413 | 615 | 804 | 994 |  | 3.56e-05 | 413 | 615 | 804 | 994
3.58e-05 | 402 | 592 | 781 | 970 |  | 3.58e-05 | 413 | 615 | 804 | 994
3.6e-05 | 402 | 592 | 757 | 946 |  | 3.6e-05 | 402 | 592 | 781 | 970
3.62e-05 | 390 | 568 | 757 | 923 |  | 3.62e-05 | 390 | 592 | 757 | 946
3.64e-05 | 378 | 544 | 734 | 899 |  | 3.64e-05 | 378 | 568 | 734 | 923
3.66e-05 | 366 | 544 | 710 | 875 |  | 3.66e-05 | 366 | 544 | 710 | 875
3.68e-05 | 354 | 521 | 686 | 828 |  | 3.68e-05 | 354 | 521 | 686 | 852
3.7e-05 | 343 | 497 | 663 | 804 |  | 3.7e-05 | 343 | 521 | 663 | 828
3.72e-05 | 331 | 497 | 639 | 781 |  | 3.72e-05 | 331 | 497 | 639 | 804
3.74e-05 | 319 | 450 | 615 | 734 |  | 3.74e-05 | 319 | 473 | 615 | 757
3.76e-05 | 295 | 450 | 568 | 710 |  | 3.76e-05 | 307 | 450 | 592 | 710
3.78e-05 | 283 | 426 | 544 | 663 |  | 3.78e-05 | 283 | 426 | 544 | 686
3.8e-05 | 272 | 402 | 521 | 639 |  | 3.8e-05 | 272 | 402 | 521 | 639
3.82e-05 | 260 | 379 | 497 | 592 |  | 3.82e-05 | 260 | 379 | 497 | 615
3.84e-05 | 236 | 331 | 450 | 544 |  | 3.84e-05 | 248 | 355 | 450 | 568
3.86e-05 | 224 | 331 | 426 | 521 |  | 3.86e-05 | 224 | 331 | 426 | 521
3.88e-05 | 201 | 284 | 379 | 473 |  | 3.88e-05 | 213 | 308 | 402 | 497
3.9e-05 | 189 | 260 | 355 | 426 |  | 3.9e-05 | 189 | 284 | 355 | 450
3.92e-05 | 165 | 237 | 308 | 402 |  | 3.92e-05 | 177 | 260 | 331 | 402
3.94e-05 | 154 | 213 | 284 | 331 |  | 3.94e-05 | 154 | 213 | 284 | 355
3.96e-05 | 130 | 189 | 260 | 308 |  | 3.96e-05 | 142 | 213 | 260 | 308
3.98e-05 | 118 | 166 | 213 | 260 |  | 3.98e-05 | 118 | 166 | 237 | 284
4e-05 | 94 | 142 | 189 | 213 |  | 4e-05 | 106 | 166 | 189 | 237
4.02e-05 | 83 | 118 | 142 | 189 |  | 4.02e-05 | 94 | 118 | 166 | 189
4.04e-05 | 59 | 71 | 118 | 142 |  | 4.04e-05 | 71 | 95 | 118 | 166
4.06e-05 | 47 | 71 | 71 | 95 |  | 4.06e-05 | 47 | 71 | 95 | 118
4.08e-05 | 24 | 24 | 47 | 47 |  | 4.08e-05 | 35 | 47 | 47 | 71
4.1e-05 | 12 | 0 | 0 | 0 |  | 4.1e-05 | 12 | 24 | 24 | 24
4.12e-05 | -11 | -23 | -23 | -23 |  | 4.12e-05 | 0 | 0 | -23 | -23
4.14e-05 | -23 | -46 | -70 | -70 |  | 4.14e-05 | -11 | -46 | -46 | -70
4.16e-05 | -34 | -70 | -94 | -117 |  | 4.16e-05 | -34 | -70 | -94 | -117
4.18e-05 | -58 | -94 | -141 | -165 |  | 4.18e-05 | -58 | -94 | -117 | -141
4.2e-05 | -70 | -117 | -165 | -212 |  | 4.2e-05 | -70 | -117 | -165 | -188
4.22e-05 | -93 | -141 | -212 | -236 |  | 4.22e-05 | -82 | -141 | -188 | -236
4.24e-05 | -105 | -165 | -236 | -283 |  | 4.24e-05 | -105 | -165 | -212 | -259
4.26e-05 | -117 | -188 | -259 | -307 |  | 4.26e-05 | -117 | -188 | -236 | -307
4.28e-05 | -141 | -212 | -283 | -354 |  | 4.28e-05 | -129 | -212 | -259 | -330
4.3e-05 | -153 | -236 | -307 | -378 |  | 4.3e-05 | -141 | -236 | -307 | -378
4.32e-05 | -164 | -259 | -330 | -401 |  | 4.32e-05 | -164 | -259 | -330 | -401
4.34e-05 | -176 | -283 | -354 | -449 |  | 4.34e-05 | -176 | -259 | -354 | -425
4.36e-05 | -188 | -307 | -378 | -472 |  | 4.36e-05 | -188 | -283 | -378 | -472
4.38e-05 | -200 | -307 | -401 | -496 |  | 4.38e-05 | -200 | -307 | -401 | -496
4.4e-05 | -212 | -330 | -425 | -520 |  | 4.4e-05 | -212 | -330 | -425 | -520
4.42e-05 | -223 | -330 | -449 | -543 |  | 4.42e-05 | -212 | -330 | -449 | -543
4.44e-05 | -235 | -354 | -472 | -591 |  | 4.44e-05 | -235 | -354 | -472 | -567
4.46e-05 | -247 | -378 | -496 | -614 |  | 4.46e-05 | -247 | -354 | -496 | -591
4.48e-05 | -259 | -378 | -520 | -614 |  | 4.48e-05 | -247 | -378 | -496 | -614
4.5e-05 | -259 | -401 | -520 | -638 |  | 4.5e-05 | -259 | -401 | -520 | -638
4.52e-05 | -271 | -401 | -543 | -662 |  | 4.52e-05 | -271 | -401 | -543 | -662
4.54e-05 | -282 | -425 | -543 | -685 |  | 4.54e-05 | -271 | -425 | -543 | -685
4.56e-05 | -282 | -425 | -567 | -685 |  | 4.56e-05 | -282 | -425 | -567 | -685
4.58e-05 | -282 | -425 | -567 | -709 |  | 4.58e-05 | -282 | -425 | -567 | -709
4.6e-05 | -294 | -449 | -591 | -709 |  | 4.6e-05 | -294 | -449 | -567 | -733
4.62e-05 | -294 | -449 | -591 | -733 |  | 4.62e-05 | -294 | -449 | -591 | -733
4.64e-05 | -306 | -449 | -591 | -733 |  | 4.64e-05 | -306 | -449 | -591 | -733
4.66e-05 | -306 | -449 | -591 | -733 |  | 4.66e-05 | -306 | -449 | -591 | -756
4.68e-05 | -306 | -449 | -614 | -756 |  | 4.68e-05 | -306 | -472 | -591 | -756
4.7e-05 | -306 | -472 | -614 | -756 |  | 4.7e-05 | -306 | -472 | -614 | -756
4.72e-05 | -306 | -472 | -614 | -756 |  | 4.72e-05 | -306 | -472 | -614 | -756
4.74e-05 | -306 | -449 | -614 | -756 |  | 4.74e-05 | -306 | -472 | -614 | -756
4.76e-05 | -306 | -449 | -614 | -756 |  | 4.76e-05 | -306 | -472 | -614 | -756
4.78e-05 | -306 | -472 | -614 | -733 |  | 4.78e-05 | -306 | -472 | -614 | -756
4.8e-05 | -306 | -449 | -591 | -733 |  | 4.8e-05 | -306 | -472 | -614 | -756
4.82e-05 | -306 | -449 | -591 | -733 |  | 4.82e-05 | -306 | -449 | -591 | -733
4.84e-05 | -294 | -449 | -591 | -709 |  | 4.84e-05 | -306 | -449 | -591 | -733
4.86e-05 | -294 | -449 | -567 | -709 |  | 4.86e-05 | -294 | -449 | -591 | -709
4.88e-05 | -282 | -449 | -567 | -709 |  | 4.88e-05 | -294 | -449 | -567 | -709
4.9e-05 | -282 | -425 | -567 | -685 |  | 4.9e-05 | -294 | -425 | -567 | -709
4.92e-05 | -282 | -425 | -543 | -662 |  | 4.92e-05 | -282 | -425 | -543 | -685
4.94e-05 | -271 | -401 | -543 | -662 |  | 4.94e-05 | -282 | -425 | -543 | -662
4.96e-05 | -259 | -401 | -520 | -638 |  | 4.96e-05 | -271 | -401 | -520 | -662
4.98e-05 | -259 | -401 | -520 | -614 |  | 4.98e-05 | -271 | -378 | -520 | -638
5e-05 | -247 | -378 | -496 | -614 |  | 5e-05 | -259 | -378 | -496 | -614
5.02e-05 | -247 | -378 | -472 | -591 |  | 5.02e-05 | -247 | -378 | -496 | -591
5.04e-05 | -235 | -354 | -449 | -567 |  | 5.04e-05 | -247 | -354 | -472 | -567
5.06e-05 | -223 | -330 | -449 | -543 |  | 5.06e-05 | -235 | -354 | -449 | -567
5.08e-05 | -212 | -330 | -425 | -520 |  | 5.08e-05 | -223 | -330 | -425 | -520
5.1e-05 | -200 | -307 | -401 | -496 |  | 5.1e-05 | -212 | -330 | -401 | -520
5.12e-05 | -200 | -307 | -378 | -472 |  | 5.12e-05 | -200 | -307 | -378 | -496
5.14e-05 | -188 | -283 | -354 | -425 |  | 5.14e-05 | -188 | -307 | -378 | -449
5.16e-05 | -176 | -259 | -330 | -425 |  | 5.16e-05 | -176 | -283 | -354 | -425
5.18e-05 | -164 | -236 | -330 | -378 |  | 5.18e-05 | -164 | -259 | -330 | -401
5.2e-05 | -153 | -236 | -307 | -354 |  | 5.2e-05 | -164 | -236 | -307 | -378
5.22e-05 | -141 | -212 | -283 | -330 |  | 5.22e-05 | -153 | -236 | -283 | -354
5.24e-05 | -129 | -188 | -259 | -307 |  | 5.24e-05 | -141 | -212 | -283 | -330
5.26e-05 | -117 | -188 | -236 | -283 |  | 5.26e-05 | -129 | -188 | -236 | -307
5.28e-05 | -105 | -165 | -212 | -259 |  | 5.28e-05 | -117 | -165 | -212 | -259
5.3e-05 | -93 | -141 | -188 | -212 |  | 5.3e-05 | -105 | -165 | -188 | -236
5.32e-05 | -82 | -117 | -165 | -188 |  | 5.32e-05 | -93 | -141 | -165 | -212
5.34e-05 | -70 | -117 | -141 | -165 |  | 5.34e-05 | -82 | -117 | -141 | -188
5.36e-05 | -58 | -94 | -117 | -141 |  | 5.36e-05 | -58 | -94 | -117 | -141
5.38e-05 | -46 | -70 | -94 | -117 |  | 5.38e-05 | -58 | -70 | -94 | -117
5.4e-05 | -34 | -46 | -70 | -70 |  | 5.4e-05 | -46 | -70 | -70 | -94
5.42e-05 | -23 | -23 | -46 | -46 |  | 5.42e-05 | -23 | -46 | -46 | -46
5.44e-05 | -11 | -23 | -23 | -23 |  | 5.44e-05 | -11 | -23 | -23 | -23
5.46e-05 | 0 | 0 | 0 | 0 |  | 5.46e-05 | 0 | 0 | 0 | -23
5.48e-05 | 12 | 24 | 24 | 47 |  | 5.48e-05 | 12 | 0 | 24 | 24
5.5e-05 | 24 | 24 | 47 | 71 |  | 5.5e-05 | 24 | 24 | 47 | 47
5.52e-05 | 35 | 47 | 71 | 95 |  | 5.52e-05 | 24 | 47 | 71 | 71
5.54e-05 | 47 | 71 | 95 | 118 |  | 5.54e-05 | 35 | 47 | 71 | 118
5.56e-05 | 59 | 71 | 118 | 142 |  | 5.56e-05 | 59 | 71 | 95 | 142
5.58e-05 | 71 | 95 | 142 | 166 |  | 5.58e-05 | 59 | 95 | 142 | 166
5.6e-05 | 83 | 118 | 166 | 189 |  | 5.6e-05 | 71 | 118 | 142 | 189
5.62e-05 | 94 | 142 | 166 | 237 |  | 5.62e-05 | 83 | 118 | 166 | 213
5.64e-05 | 94 | 142 | 189 | 237 |  | 5.64e-05 | 94 | 142 | 189 | 237
5.66e-05 | 106 | 166 | 213 | 260 |  | 5.66e-05 | 106 | 166 | 213 | 260
5.68e-05 | 118 | 166 | 237 | 284 |  | 5.68e-05 | 118 | 166 | 213 | 284
5.7e-05 | 118 | 189 | 237 | 308 |  | 5.7e-05 | 118 | 189 | 237 | 284
5.72e-05 | 130 | 189 | 260 | 331 |  | 5.72e-05 | 130 | 189 | 260 | 308
5.74e-05 | 130 | 213 | 260 | 331 |  | 5.74e-05 | 130 | 213 | 260 | 331
5.76e-05 | 142 | 213 | 284 | 355 |  | 5.76e-05 | 142 | 213 | 284 | 355
5.78e-05 | 154 | 237 | 308 | 379 |  | 5.78e-05 | 154 | 213 | 308 | 379
5.8e-05 | 154 | 237 | 308 | 379 |  | 5.8e-05 | 154 | 237 | 308 | 379
5.82e-05 | 165 | 237 | 331 | 402 |  | 5.82e-05 | 165 | 237 | 331 | 402
5.84e-05 | 165 | 260 | 331 | 402 |  | 5.84e-05 | 165 | 237 | 331 | 402
5.86e-05 | 177 | 260 | 355 | 426 |  | 5.86e-05 | 177 | 260 | 355 | 426
5.88e-05 | 177 | 260 | 355 | 426 |  | 5.88e-05 | 177 | 260 | 355 | 450
5.9e-05 | 177 | 260 | 355 | 450 |  | 5.9e-05 | 177 | 284 | 355 | 450
5.92e-05 | 189 | 284 | 379 | 450 |  | 5.92e-05 | 189 | 284 | 379 | 450
5.94e-05 | 189 | 284 | 379 | 450 |  | 5.94e-05 | 189 | 284 | 379 | 473
5.96e-05 | 201 | 284 | 379 | 473 |  | 5.96e-05 | 189 | 284 | 379 | 473
5.98e-05 | 201 | 284 | 379 | 473 |  | 5.98e-05 | 201 | 284 | 379 | 473
6e-05 | 201 | 284 | 379 | 473 |  | 6e-05 | 201 | 308 | 379 | 473
6.02e-05 | 201 | 308 | 379 | 497 |  | 6.02e-05 | 201 | 308 | 402 | 497
6.04e-05 | 201 | 308 | 379 | 497 |  | 6.04e-05 | 201 | 284 | 402 | 497
6.06e-05 | 201 | 308 | 379 | 497 |  | 6.06e-05 | 201 | 308 | 402 | 497
6.08e-05 | 201 | 308 | 379 | 497 |  | 6.08e-05 | 201 | 284 | 402 | 497
6.1e-05 | 201 | 308 | 379 | 497 |  | 6.1e-05 | 201 | 308 | 379 | 497
6.12e-05 | 201 | 308 | 402 | 497 |  | 6.12e-05 | 201 | 308 | 402 | 497
6.14e-05 | 201 | 284 | 379 | 473 |  | 6.14e-05 | 201 | 284 | 379 | 497
6.16e-05 | 201 | 284 | 379 | 473 |  | 6.16e-05 | 201 | 308 | 379 | 473
6.18e-05 | 201 | 284 | 379 | 473 |  | 6.18e-05 | 201 | 284 | 379 | 473
6.2e-05 | 189 | 284 | 379 | 450 |  | 6.2e-05 | 189 | 284 | 379 | 473
6.22e-05 | 189 | 284 | 355 | 450 |  | 6.22e-05 | 189 | 284 | 379 | 450
6.24e-05 | 189 | 260 | 355 | 450 |  | 6.24e-05 | 189 | 284 | 355 | 450
6.26e-05 | 189 | 260 | 355 | 426 |  | 6.26e-05 | 189 | 284 | 355 | 450
6.28e-05 | 177 | 260 | 331 | 426 |  | 6.28e-05 | 189 | 260 | 355 | 426
6.3e-05 | 177 | 260 | 331 | 402 |  | 6.3e-05 | 177 | 260 | 355 | 426
6.32e-05 | 165 | 260 | 331 | 402 |  | 6.32e-05 | 177 | 260 | 331 | 402
6.34e-05 | 165 | 237 | 308 | 379 |  | 6.34e-05 | 165 | 260 | 331 | 402
6.36e-05 | 165 | 237 | 308 | 355 |  | 6.36e-05 | 165 | 237 | 308 | 379
6.38e-05 | 154 | 237 | 284 | 355 |  | 6.38e-05 | 154 | 237 | 308 | 379
6.4e-05 | 142 | 213 | 284 | 355 |  | 6.4e-05 | 154 | 237 | 284 | 355
6.42e-05 | 142 | 213 | 260 | 331 |  | 6.42e-05 | 142 | 213 | 284 | 331
6.44e-05 | 130 | 189 | 260 | 308 |  | 6.44e-05 | 142 | 213 | 260 | 331
6.46e-05 | 130 | 189 | 237 | 284 |  | 6.46e-05 | 130 | 189 | 260 | 308
6.48e-05 | 118 | 189 | 237 | 284 |  | 6.48e-05 | 130 | 189 | 237 | 284
6.5e-05 | 118 | 166 | 213 | 260 |  | 6.5e-05 | 118 | 166 | 237 | 284
6.52e-05 | 106 | 166 | 213 | 260 |  | 6.52e-05 | 106 | 166 | 213 | 260
6.54e-05 | 94 | 142 | 189 | 237 |  | 6.54e-05 | 106 | 142 | 189 | 237
6.56e-05 | 94 | 142 | 166 | 213 |  | 6.56e-05 | 94 | 142 | 189 | 213
6.58e-05 | 83 | 118 | 166 | 189 |  | 6.58e-05 | 83 | 118 | 166 | 213
6.6e-05 | 83 | 118 | 142 | 166 |  | 6.6e-05 | 83 | 118 | 142 | 189
6.62e-05 | 71 | 95 | 118 | 142 |  | 6.62e-05 | 71 | 118 | 142 | 166
6.64e-05 | 59 | 95 | 118 | 142 |  | 6.64e-05 | 71 | 95 | 118 | 142
6.66e-05 | 59 | 71 | 95 | 118 |  | 6.66e-05 | 59 | 95 | 118 | 142
6.68e-05 | 47 | 71 | 71 | 95 |  | 6.68e-05 | 47 | 71 | 95 | 118
6.7e-05 | 35 | 47 | 71 | 71 |  | 6.7e-05 | 35 | 47 | 71 | 95
6.72e-05 | 24 | 47 | 47 | 47 |  | 6.72e-05 | 35 | 47 | 71 | 71
6.74e-05 | 24 | 24 | 24 | 47 |  | 6.74e-05 | 24 | 24 | 47 | 47
6.76e-05 | 12 | 24 | 0 | 24 |  | 6.76e-05 | 24 | 24 | 24 | 24
6.78e-05 | 0 | 0 | 0 | 0 |  | 6.78e-05 | 12 | 0 | 24 | 24
6.8e-05 | 0 | 0 | -23 | -23 |  | 6.8e-05 | 0 | 0 | 0 | 0
6.82e-05 | -11 | -23 | -23 | -46 |  | 6.82e-05 | -11 | -23 | -23 | -23
6.84e-05 | -23 | -23 | -46 | -46 |  | 6.84e-05 | -11 | -23 | -46 | -46
6.86e-05 | -23 | -46 | -46 | -70 |  | 6.86e-05 | -23 | -23 | -46 | -70
6.88e-05 | -34 | -46 | -70 | -94 |  | 6.88e-05 | -34 | -46 | -70 | -94
6.9e-05 | -34 | -70 | -94 | -117 |  | 6.9e-05 | -34 | -70 | -70 | -117
6.92e-05 | -46 | -70 | -117 | -141 |  | 6.92e-05 | -46 | -70 | -94 | -117
6.94e-05 | -58 | -94 | -117 | -141 |  | 6.94e-05 | -58 | -70 | -117 | -141
6.96e-05 | -58 | -94 | -141 | -165 |  | 6.96e-05 | -58 | -94 | -117 | -165
6.98e-05 | -70 | -117 | -141 | -188 |  | 6.98e-05 | -70 | -94 | -141 | -165
7e-05 | -82 | -117 | -165 | -188 |  | 7e-05 | -70 | -117 | -141 | -188
7.02e-05 | -82 | -117 | -165 | -188 |  | 7.02e-05 | -82 | -117 | -165 | -188
7.04e-05 | -82 | -141 | -188 | -212 |  | 7.04e-05 | -82 | -117 | -165 | -212
7.06e-05 | -93 | -141 | -188 | -236 |  | 7.06e-05 | -93 | -141 | -188 | -212
7.08e-05 | -93 | -141 | -188 | -236 |  | 7.08e-05 | -93 | -165 | -188 | -236
7.1e-05 | -105 | -141 | -212 | -259 |  | 7.1e-05 | -105 | -165 | -212 | -236
7.12e-05 | -105 | -165 | -212 | -259 |  | 7.12e-05 | -105 | -165 | -212 | -259
7.14e-05 | -105 | -165 | -236 | -283 |  | 7.14e-05 | -105 | -165 | -212 | -259
7.16e-05 | -117 | -165 | -236 | -283 |  | 7.16e-05 | -117 | -165 | -236 | -283
7.18e-05 | -117 | -188 | -236 | -307 |  | 7.18e-05 | -117 | -188 | -236 | -283
7.2e-05 | -129 | -188 | -259 | -307 |  | 7.2e-05 | -117 | -188 | -236 | -283
7.22e-05 | -129 | -188 | -259 | -307 |  | 7.22e-05 | -117 | -188 | -259 | -307
7.24e-05 | -129 | -212 | -259 | -307 |  | 7.24e-05 | -129 | -188 | -259 | -307
7.26e-05 | -129 | -212 | -259 | -307 |  | 7.26e-05 | -129 | -212 | -259 | -330
7.28e-05 | -129 | -212 | -259 | -330 |  | 7.28e-05 | -129 | -188 | -259 | -330
7.3e-05 | -141 | -212 | -283 | -330 |  | 7.3e-05 | -129 | -188 | -259 | -330
7.32e-05 | -141 | -212 | -283 | -330 |  | 7.32e-05 | -141 | -212 | -283 | -330
7.34e-05 | -141 | -212 | -283 | -330 |  | 7.34e-05 | -141 | -212 | -283 | -330
7.36e-05 | -141 | -212 | -283 | -330 |  | 7.36e-05 | -141 | -212 | -283 | -354
7.38e-05 | -141 | -212 | -283 | -330 |  | 7.38e-05 | -141 | -212 | -283 | -330
7.4e-05 | -141 | -212 | -283 | -330 |  | 7.4e-05 | -141 | -212 | -283 | -354
7.42e-05 | -141 | -212 | -283 | -330 |  | 7.42e-05 | -141 | -212 | -283 | -330
7.44e-05 | -141 | -212 | -283 | -330 |  | 7.44e-05 | -141 | -212 | -283 | -330
7.46e-05 | -141 | -212 | -283 | -330 |  | 7.46e-05 | -141 | -212 | -283 | -330
7.48e-05 | -141 | -212 | -283 | -330 |  | 7.48e-05 | -141 | -212 | -283 | -330
7.5e-05 | -141 | -212 | -259 | -330 |  | 7.5e-05 | -141 | -212 | -259 | -330
7.52e-05 | -141 | -212 | -259 | -330 |  | 7.52e-05 | -141 | -212 | -283 | -330
7.54e-05 | -141 | -212 | -259 | -330 |  | 7.54e-05 | -129 | -212 | -259 | -330
7.56e-05 | -129 | -188 | -259 | -307 |  | 7.56e-05 | -129 | -212 | -259 | -330
7.58e-05 | -129 | -188 | -259 | -307 |  | 7.58e-05 | -129 | -212 | -259 | -307
7.6e-05 | -129 | -188 | -259 | -307 |  | 7.6e-05 | -129 | -188 | -259 | -307
7.62e-05 | -129 | -188 | -236 | -283 |  | 7.62e-05 | -129 | -188 | -259 | -307
7.64e-05 | -117 | -188 | -236 | -283 |  | 7.64e-05 | -117 | -188 | -236 | -307
7.66e-05 | -117 | -165 | -236 | -283 |  | 7.66e-05 | -117 | -188 | -236 | -307
7.68e-05 | -117 | -165 | -212 | -283 |  | 7.68e-05 | -117 | -188 | -236 | -283
7.7e-05 | -105 | -165 | -212 | -259 |  | 7.7e-05 | -105 | -165 | -236 | -283
7.72e-05 | -105 | -165 | -212 | -259 |  | 7.72e-05 | -105 | -165 | -212 | -259
7.74e-05 | -105 | -165 | -188 | -236 |  | 7.74e-05 | -105 | -165 | -212 | -259
7.76e-05 | -93 | -141 | -188 | -236 |  | 7.76e-05 | -105 | -165 | -188 | -259
7.78e-05 | -93 | -141 | -188 | -212 |  | 7.78e-05 | -93 | -141 | -188 | -236
7.8e-05 | -93 | -141 | -188 | -212 |  | 7.8e-05 | -93 | -141 | -188 | -236
7.82e-05 | -82 | -141 | -165 | -212 |  | 7.82e-05 | -93 | -141 | -188 | -212
7.84e-05 | -82 | -117 | -165 | -188 |  | 7.84e-05 | -82 | -117 | -165 | -212
7.86e-05 | -70 | -117 | -141 | -188 |  | 7.86e-05 | -82 | -117 | -165 | -188
7.88e-05 | -70 | -117 | -141 | -165 |  | 7.88e-05 | -70 | -117 | -141 | -188
7.9e-05 | -70 | -94 | -117 | -165 |  | 7.9e-05 | -70 | -117 | -141 | -165
7.92e-05 | -58 | -94 | -117 | -141 |  | 7.92e-05 | -70 | -94 | -117 | -165
7.94e-05 | -58 | -94 | -117 | -141 |  | 7.94e-05 | -58 | -94 | -117 | -141
7.96e-05 | -46 | -70 | -94 | -117 |  | 7.96e-05 | -58 | -70 | -117 | -141
7.98e-05 | -46 | -70 | -94 | -94 |  | 7.98e-05 | -46 | -70 | -94 | -117
8e-05 | -34 | -46 | -70 | -94 |  | 8e-05 | -46 | -70 | -94 | -94
8.02e-05 | -34 | -46 | -70 | -70 |  | 8.02e-05 | -34 | -46 | -70 | -94
8.04e-05 | -23 | -46 | -46 | -70 |  | 8.04e-05 | -34 | -46 | -70 | -70
8.06e-05 | -23 | -23 | -46 | -46 |  | 8.06e-05 | -23 | -46 | -46 | -70
8.08e-05 | -11 | -23 | -23 | -46 |  | 8.08e-05 | -23 | -46 | -46 | -46
8.1e-05 | -11 | -23 | -23 | -23 |  | 8.1e-05 | -11 | -23 | -23 | -46
8.12e-05 | -11 | 0 | 0 | -23 |  | 8.12e-05 | -11 | -23 | -23 | -23
8.14e-05 | 0 | 0 | 0 | 0 |  | 8.14e-05 | 0 | 0 | 0 | 0
8.16e-05 | 0 | 0 | 0 | 24 |  | 8.16e-05 | 0 | 0 | 0 | 0
8.18e-05 | 12 | 0 | 24 | 24 |  | 8.18e-05 | 0 | 0 | 24 | 24
8.2e-05 | 12 | 24 | 24 | 47 |  | 8.2e-05 | 12 | 24 | 24 | 24
8.22e-05 | 24 | 24 | 47 | 47 |  | 8.22e-05 | 12 | 24 | 24 | 47
8.24e-05 | 24 | 24 | 47 | 71 |  | 8.24e-05 | 24 | 24 | 47 | 47
8.26e-05 | 24 | 47 | 47 | 71 |  | 8.26e-05 | 24 | 47 | 47 | 71
8.28e-05 | 35 | 47 | 71 | 71 |  | 8.28e-05 | 24 | 47 | 71 | 71
8.3e-05 | 35 | 47 | 71 | 95 |  | 8.3e-05 | 35 | 47 | 71 | 71
8.32e-05 | 47 | 47 | 71 | 95 |  | 8.32e-05 | 35 | 47 | 71 | 95
8.34e-05 | 47 | 71 | 95 | 118 |  | 8.34e-05 | 47 | 71 | 95 | 118
8.36e-05 | 59 | 95 | 95 | 118 |  | 8.36e-05 | 47 | 71 | 95 | 118
8.38e-05 | 59 | 71 | 118 | 142 |  | 8.38e-05 | 59 | 71 | 95 | 142
8.4e-05 | 59 | 95 | 118 | 142 |  | 8.4e-05 | 59 | 95 | 118 | 142
8.42e-05 | 59 | 95 | 118 | 142 |  | 8.42e-05 | 59 | 95 | 118 | 142
8.44e-05 | 59 | 95 | 118 | 166 |  | 8.44e-05 | 59 | 95 | 118 | 166
8.46e-05 | 71 | 95 | 142 | 166 |  | 8.46e-05 | 71 | 95 | 118 | 166
8.48e-05 | 71 | 95 | 142 | 166 |  | 8.48e-05 | 71 | 95 | 142 | 166
8.5e-05 | 71 | 118 | 142 | 166 |  | 8.5e-05 | 71 | 118 | 142 | 166
8.52e-05 | 83 | 118 | 142 | 189 |  | 8.52e-05 | 71 | 118 | 142 | 189
8.54e-05 | 83 | 118 | 142 | 189 |  | 8.54e-05 | 71 | 118 | 142 | 189
8.56e-05 | 83 | 118 | 166 | 189 |  | 8.56e-05 | 83 | 118 | 166 | 189
8.58e-05 | 83 | 118 | 166 | 189 |  | 8.58e-05 | 83 | 118 | 166 | 189
8.6e-05 | 83 | 118 | 166 | 189 |  | 8.6e-05 | 83 | 118 | 166 | 213
8.62e-05 | 83 | 118 | 166 | 213 |  | 8.62e-05 | 83 | 118 | 166 | 213
8.64e-05 | 94 | 118 | 166 | 213 |  | 8.64e-05 | 94 | 118 | 166 | 213
8.66e-05 | 94 | 142 | 166 | 213 |  | 8.66e-05 | 94 | 118 | 166 | 213
8.68e-05 | 94 | 142 | 166 | 213 |  | 8.68e-05 | 83 | 142 | 166 | 213
8.7e-05 | 94 | 142 | 166 | 213 |  | 8.7e-05 | 94 | 142 | 166 | 213
8.72e-05 | 94 | 142 | 166 | 213 |  | 8.72e-05 | 94 | 142 | 189 | 213
8.74e-05 | 94 | 142 | 166 | 213 |  | 8.74e-05 | 94 | 142 | 166 | 213
8.76e-05 | 94 | 142 | 166 | 213 |  | 8.76e-05 | 94 | 142 | 166 | 213
8.78e-05 | 94 | 142 | 166 | 213 |  | 8.78e-05 | 94 | 142 | 189 | 213
8.8e-05 | 94 | 142 | 166 | 213 |  | 8.8e-05 | 94 | 142 | 166 | 213
8.82e-05 | 83 | 142 | 166 | 213 |  | 8.82e-05 | 94 | 142 | 189 | 213
8.84e-05 | 83 | 118 | 166 | 213 |  | 8.84e-05 | 94 | 142 | 166 | 213
8.86e-05 | 94 | 142 | 166 | 213 |  | 8.86e-05 | 83 | 142 | 166 | 213
8.88e-05 | 94 | 118 | 166 | 213 |  | 8.88e-05 | 83 | 142 | 166 | 213
8.9e-05 | 83 | 118 | 166 | 189 |  | 8.9e-05 | 94 | 118 | 166 | 213
8.92e-05 | 83 | 118 | 166 | 189 |  | 8.92e-05 | 83 | 118 | 166 | 213
8.94e-05 | 83 | 118 | 166 | 189 |  | 8.94e-05 | 83 | 118 | 166 | 189
8.96e-05 | 83 | 118 | 142 | 189 |  | 8.96e-05 | 83 | 118 | 166 | 189
8.98e-05 | 83 | 118 | 142 | 189 |  | 8.98e-05 | 83 | 118 | 166 | 189
9e-05 | 71 | 118 | 142 | 166 |  | 9e-05 | 83 | 118 | 142 | 189
9.02e-05 | 71 | 118 | 142 | 166 |  | 9.02e-05 | 71 | 118 | 142 | 189
9.04e-05 | 71 | 95 | 142 | 166 |  | 9.04e-05 | 71 | 118 | 142 | 166
9.06e-05 | 71 | 95 | 118 | 166 |  | 9.06e-05 | 71 | 118 | 142 | 166
9.08e-05 | 71 | 95 | 118 | 142 |  | 9.08e-05 | 71 | 95 | 142 | 166
9.1e-05 | 59 | 95 | 118 | 142 |  | 9.1e-05 | 71 | 95 | 118 | 166
9.12e-05 | 59 | 95 | 118 | 142 |  | 9.12e-05 | 71 | 95 | 118 | 142
9.14e-05 | 59 | 71 | 118 | 142 |  | 9.14e-05 | 59 | 95 | 118 | 142
9.16e-05 | 59 | 71 | 95 | 118 |  | 9.16e-05 | 59 | 71 | 118 | 142
9.18e-05 | 59 | 71 | 95 | 118 |  | 9.18e-05 | 59 | 71 | 118 | 142
9.2e-05 | 47 | 71 | 95 | 118 |  | 9.2e-05 | 47 | 71 | 95 | 118
9.22e-05 | 47 | 71 | 95 | 95 |  | 9.22e-05 | 47 | 71 | 95 | 118
9.24e-05 | 35 | 71 | 71 | 95 |  | 9.24e-05 | 47 | 71 | 71 | 118
9.26e-05 | 35 | 47 | 71 | 95 |  | 9.26e-05 | 47 | 47 | 71 | 95
9.28e-05 | 35 | 47 | 71 | 71 |  | 9.28e-05 | 35 | 47 | 71 | 95
9.3e-05 | 35 | 47 | 47 | 71 |  | 9.3e-05 | 35 | 47 | 71 | 71
9.32e-05 | 24 | 47 | 47 | 71 |  | 9.32e-05 | 24 | 47 | 47 | 71
9.34e-05 | 24 | 24 | 47 | 47 |  | 9.34e-05 | 24 | 47 | 47 | 47
9.36e-05 | 24 | 24 | 24 | 47 |  | 9.36e-05 | 24 | 24 | 47 | 47
9.38e-05 | 12 | 24 | 24 | 24 |  | 9.38e-05 | 24 | 24 | 24 | 47
9.4e-05 | 12 | 24 | 24 | 24 |  | 9.4e-05 | 24 | 24 | 24 | 47
9.42e-05 | 12 | 24 | 24 | 24 |  | 9.42e-05 | 12 | 0 | 24 | 24
9.44e-05 | 12 | 0 | 0 | 0 |  | 9.44e-05 | 12 | 0 | 0 | 24
9.46e-05 | 0 | 0 | 0 | 0 |  | 9.46e-05 | 12 | 0 | 0 | 0
9.48e-05 | 0 | 0 | 0 | 0 |  | 9.48e-05 | 0 | 0 | 0 | 0
9.5e-05 | 0 | -23 | -23 | -23 |  | 9.5e-05 | 0 | 0 | 0 | 0
9.52e-05 | -11 | 0 | -23 | -23 |  | 9.52e-05 | -11 | 0 | 0 | -23
9.54e-05 | -11 | -23 | -23 | -23 |  | 9.54e-05 | 0 | 0 | -23 | -23
9.56e-05 | -11 | -23 | -23 | -46 |  | 9.56e-05 | -11 | -23 | -23 | -23
9.58e-05 | -11 | -46 | -46 | -46 |  | 9.58e-05 | -11 | -23 | -23 | -46
9.6e-05 | -23 | -46 | -46 | -46 |  | 9.6e-05 | -11 | -23 | -46 | -46
9.62e-05 | -23 | -46 | -46 | -70 |  | 9.62e-05 | -23 | -23 | -46 | -46
9.64e-05 | -23 | -46 | -46 | -70 |  | 9.64e-05 | -23 | -46 | -46 | -70
9.66e-05 | -23 | -46 | -70 | -70 |  | 9.66e-05 | -23 | -46 | -46 | -70
9.68e-05 | -23 | -46 | -70 | -70 |  | 9.68e-05 | -34 | -46 | -46 | -70
9.7e-05 | -34 | -46 | -70 | -94 |  | 9.7e-05 | -34 | -46 | -70 | -70
9.72e-05 | -34 | -70 | -70 | -94 |  | 9.72e-05 | -34 | -70 | -70 | -94
9.74e-05 | -34 | -70 | -94 | -94 |  | 9.74e-05 | -34 | -70 | -70 | -94
9.76e-05 | -46 | -70 | -94 | -117 |  | 9.76e-05 | -34 | -70 | -94 | -94
9.78e-05 | -46 | -70 | -94 | -117 |  | 9.78e-05 | -46 | -70 | -94 | -117
9.8e-05 | -46 | -70 | -94 | -117 |  | 9.8e-05 | -46 | -70 | -94 | -117
9.82e-05 | -46 | -70 | -94 | -117 |  | 9.82e-05 | -46 | -70 | -94 | -117
9.84e-05 | -58 | -70 | -117 | -117 |  | 9.84e-05 | -46 | -70 | -117 | -117
9.86e-05 | -46 | -70 | -117 | -141 |  | 9.86e-05 | -46 | -70 | -117 | -117
9.88e-05 | -58 | -70 | -117 | -141 |  | 9.88e-05 | -58 | -94 | -117 | -141
9.9e-05 | -58 | -70 | -117 | -141 |  | 9.9e-05 | -58 | -94 | -117 | -141
9.92e-05 | -58 | -94 | -117 | -141 |  | 9.92e-05 | -58 | -94 | -117 | -141
9.94e-05 | -58 | -94 | -117 | -141 |  | 9.94e-05 | -58 | -94 | -117 | -141
9.96e-05 | -58 | -94 | -117 | -141 |  | 9.96e-05 | -58 | -94 | -117 | -141
9.98e-05 | -58 | -94 | -117 | -141 |  | 9.98e-05 | -58 | -94 | -117 | -141
0.0001 | -58 | -94 | -117 | -141 |  | 0.0001 | -58 | -94 | -117 | -141
0.0001002 | -58 | -94 | -117 | -141 |  | 0.0001002 | -58 | -94 | -117 | -141
0.0001004 | -58 | -94 | -117 | -165 |  | 0.0001004 | -58 | -94 | -117 | -141
0.0001006 | -58 | -94 | -117 | -141 |  | 0.0001006 | -58 | -94 | -141 | -165
0.0001008 | -58 | -94 | -141 | -165 |  | 0.0001008 | -58 | -94 | -117 | -165
0.000101 | -58 | -94 | -117 | -165 |  | 0.000101 | -58 | -94 | -117 | -165
0.0001012 | -58 | -94 | -117 | -165 |  | 0.0001012 | -58 | -94 | -117 | -165
0.0001014 | -58 | -94 | -117 | -141 |  | 0.0001014 | -58 | -94 | -117 | -165
0.0001016 | -58 | -94 | -117 | -165 |  | 0.0001016 | -58 | -94 | -141 | -141
0.0001018 | -58 | -94 | -117 | -141 |  | 0.0001018 | -70 | -94 | -117 | -165
0.000102 | -58 | -94 | -117 | -165 |  | 0.000102 | -58 | -94 | -117 | -165
0.0001022 | -58 | -94 | -117 | -141 |  | 0.0001022 | -58 | -94 | -117 | -141
| 2 | 2 | 2 | 2 |  |  | 2 | 2 | 2 | 2
| 1000 | 1500 | 2000 | 2500 |  |  | 1000 | 1500 | 2000 | 2500
| 2 | 2 | 2 | 2 |  |  | 2 | 2 | 2 | 2

## Step 2 224 Impulse
Time (s) | 1.0 kV | 1.5 kV | 2.0 kV | 2.5 kV | 224 | Time (s) | 1.0 kV | 1.5 kV | 2.0 kV | 2.5 kV |  | Notes
0 | 0 | 0 | 0 | 0 | A17 Coil Pack QXFA 224 Outer to Inner | 0 | 0 | 0 | 0 | 0 | A17 Coil Pack QXFA 224 Outer to Inner
2e-07 | 0 | 0 | 0 | 0 | Retest for 12b | 2e-07 | 0 | 0 | 0 | 0
4e-07 | 0 | 0 | 0 | 0 |  | 4e-07 | 0 | 0 | 0 | 0
6e-07 | 189 | 355 | 473 | 639 |  | 6e-07 | 177 | 355 | 473 | 639
8e-07 | 555 | 946 | 1372 | 1846 |  | 8e-07 | 543 | 946 | 1372 | 1846
1e-06 | 933 | 1443 | 1964 | 2508 |  | 1e-06 | 933 | 1443 | 1964 | 2484
1.2e-06 | 980 | 1443 | 1869 | 2271 |  | 1.2e-06 | 980 | 1443 | 1869 | 2271
1.4e-06 | 886 | 1278 | 1656 | 2011 |  | 1.4e-06 | 874 | 1278 | 1656 | 1988
1.6e-06 | 827 | 1230 | 1633 | 2011 |  | 1.6e-06 | 827 | 1230 | 1609 | 2011
1.8e-06 | 839 | 1254 | 1656 | 2035 |  | 1.8e-06 | 839 | 1254 | 1656 | 2035
2e-06 | 839 | 1230 | 1633 | 1988 |  | 2e-06 | 839 | 1230 | 1609 | 1988
2.2e-06 | 815 | 1207 | 1562 | 1917 |  | 2.2e-06 | 815 | 1183 | 1562 | 1940
2.4e-06 | 791 | 1159 | 1538 | 1893 |  | 2.4e-06 | 791 | 1159 | 1538 | 1893
2.6e-06 | 780 | 1159 | 1514 | 1869 |  | 2.6e-06 | 780 | 1159 | 1514 | 1869
2.8e-06 | 768 | 1136 | 1491 | 1846 |  | 2.8e-06 | 768 | 1136 | 1491 | 1846
3e-06 | 756 | 1112 | 1467 | 1822 |  | 3e-06 | 756 | 1136 | 1467 | 1822
3.2e-06 | 744 | 1088 | 1443 | 1775 |  | 3.2e-06 | 744 | 1088 | 1443 | 1775
3.4e-06 | 732 | 1065 | 1396 | 1727 |  | 3.4e-06 | 720 | 1065 | 1396 | 1727
3.6e-06 | 709 | 1041 | 1372 | 1680 |  | 3.6e-06 | 709 | 1041 | 1372 | 1680
3.8e-06 | 685 | 1017 | 1349 | 1656 |  | 3.8e-06 | 685 | 1017 | 1325 | 1656
4e-06 | 673 | 994 | 1301 | 1609 |  | 4e-06 | 673 | 994 | 1301 | 1633
4.2e-06 | 661 | 970 | 1278 | 1562 |  | 4.2e-06 | 661 | 970 | 1278 | 1585
4.4e-06 | 650 | 970 | 1254 | 1538 |  | 4.4e-06 | 650 | 970 | 1254 | 1538
4.6e-06 | 638 | 946 | 1230 | 1514 |  | 4.6e-06 | 638 | 946 | 1230 | 1514
4.8e-06 | 626 | 923 | 1207 | 1491 |  | 4.8e-06 | 626 | 923 | 1207 | 1491
5e-06 | 614 | 923 | 1183 | 1467 |  | 5e-06 | 614 | 923 | 1183 | 1467
5.2e-06 | 602 | 899 | 1159 | 1443 |  | 5.2e-06 | 602 | 899 | 1159 | 1443
5.4e-06 | 591 | 875 | 1136 | 1420 |  | 5.4e-06 | 591 | 875 | 1136 | 1420
5.6e-06 | 579 | 852 | 1112 | 1372 |  | 5.6e-06 | 579 | 875 | 1112 | 1396
5.8e-06 | 567 | 828 | 1088 | 1349 |  | 5.8e-06 | 567 | 828 | 1088 | 1349
6e-06 | 555 | 804 | 1065 | 1325 |  | 6e-06 | 555 | 828 | 1065 | 1301
6.2e-06 | 543 | 781 | 1041 | 1278 |  | 6.2e-06 | 531 | 804 | 1041 | 1278
6.4e-06 | 520 | 757 | 994 | 1230 |  | 6.4e-06 | 520 | 781 | 994 | 1230
6.6e-06 | 508 | 734 | 970 | 1183 |  | 6.6e-06 | 508 | 734 | 970 | 1183
6.8e-06 | 484 | 710 | 946 | 1159 |  | 6.8e-06 | 484 | 710 | 946 | 1159
7e-06 | 472 | 686 | 899 | 1112 |  | 7e-06 | 472 | 686 | 899 | 1112
7.2e-06 | 449 | 663 | 875 | 1065 |  | 7.2e-06 | 449 | 663 | 875 | 1065
7.4e-06 | 437 | 639 | 828 | 1017 |  | 7.4e-06 | 425 | 639 | 828 | 1017
7.6e-06 | 413 | 592 | 781 | 970 |  | 7.6e-06 | 413 | 592 | 781 | 970
7.8e-06 | 390 | 568 | 734 | 923 |  | 7.8e-06 | 390 | 568 | 734 | 923
8e-06 | 366 | 544 | 686 | 852 |  | 8e-06 | 366 | 544 | 710 | 852
8.2e-06 | 343 | 497 | 663 | 804 |  | 8.2e-06 | 343 | 497 | 663 | 804
8.4e-06 | 319 | 473 | 615 | 757 |  | 8.4e-06 | 319 | 473 | 615 | 757
8.6e-06 | 295 | 426 | 568 | 686 |  | 8.6e-06 | 295 | 426 | 568 | 686
8.8e-06 | 272 | 402 | 521 | 639 |  | 8.8e-06 | 272 | 402 | 521 | 639
9e-06 | 248 | 355 | 473 | 568 |  | 9e-06 | 248 | 355 | 473 | 592
9.2e-06 | 224 | 331 | 426 | 521 |  | 9.2e-06 | 224 | 331 | 426 | 521
9.4e-06 | 201 | 284 | 379 | 473 |  | 9.4e-06 | 201 | 284 | 379 | 473
9.6e-06 | 177 | 260 | 331 | 402 |  | 9.6e-06 | 177 | 260 | 331 | 402
9.8e-06 | 154 | 213 | 284 | 355 |  | 9.8e-06 | 154 | 237 | 284 | 355
1e-05 | 130 | 189 | 237 | 284 |  | 1e-05 | 130 | 189 | 237 | 308
1.02e-05 | 106 | 166 | 189 | 237 |  | 1.02e-05 | 106 | 142 | 189 | 237
1.04e-05 | 94 | 118 | 142 | 189 |  | 1.04e-05 | 94 | 118 | 142 | 189
1.06e-05 | 71 | 95 | 118 | 118 |  | 1.06e-05 | 71 | 71 | 118 | 118
1.08e-05 | 47 | 47 | 71 | 71 |  | 1.08e-05 | 35 | 47 | 71 | 71
1.1e-05 | 24 | 24 | 24 | 24 |  | 1.1e-05 | 24 | 24 | 24 | 24
1.12e-05 | 0 | -23 | -23 | -46 |  | 1.12e-05 | 0 | -23 | -23 | -46
1.14e-05 | -23 | -46 | -70 | -94 |  | 1.14e-05 | -23 | -46 | -70 | -94
1.16e-05 | -46 | -70 | -117 | -141 |  | 1.16e-05 | -46 | -94 | -117 | -141
1.18e-05 | -70 | -117 | -165 | -188 |  | 1.18e-05 | -70 | -117 | -165 | -188
1.2e-05 | -93 | -141 | -188 | -236 |  | 1.2e-05 | -93 | -141 | -188 | -236
1.22e-05 | -117 | -188 | -236 | -307 |  | 1.22e-05 | -117 | -165 | -236 | -283
1.24e-05 | -129 | -212 | -283 | -330 |  | 1.24e-05 | -129 | -212 | -259 | -354
1.26e-05 | -153 | -236 | -307 | -378 |  | 1.26e-05 | -153 | -236 | -307 | -401
1.28e-05 | -176 | -259 | -354 | -425 |  | 1.28e-05 | -164 | -259 | -354 | -449
1.3e-05 | -188 | -283 | -378 | -496 |  | 1.3e-05 | -188 | -307 | -401 | -496
1.32e-05 | -212 | -307 | -425 | -543 |  | 1.32e-05 | -212 | -330 | -425 | -543
1.34e-05 | -223 | -354 | -472 | -567 |  | 1.34e-05 | -235 | -354 | -472 | -591
1.36e-05 | -247 | -378 | -496 | -614 |  | 1.36e-05 | -247 | -378 | -496 | -614
1.38e-05 | -271 | -401 | -543 | -662 |  | 1.38e-05 | -271 | -401 | -543 | -662
1.4e-05 | -282 | -425 | -567 | -709 |  | 1.4e-05 | -282 | -425 | -591 | -733
1.42e-05 | -306 | -472 | -614 | -756 |  | 1.42e-05 | -306 | -472 | -614 | -756
1.44e-05 | -318 | -496 | -638 | -803 |  | 1.44e-05 | -330 | -496 | -638 | -803
1.46e-05 | -342 | -520 | -685 | -851 |  | 1.46e-05 | -342 | -520 | -685 | -851
1.48e-05 | -365 | -543 | -709 | -898 |  | 1.48e-05 | -365 | -543 | -709 | -898
1.5e-05 | -377 | -567 | -756 | -945 |  | 1.5e-05 | -377 | -567 | -756 | -945
1.52e-05 | -389 | -591 | -780 | -969 |  | 1.52e-05 | -389 | -591 | -780 | -969
1.54e-05 | -412 | -614 | -827 | -1016 |  | 1.54e-05 | -412 | -614 | -827 | -1016
1.56e-05 | -424 | -638 | -851 | -1064 |  | 1.56e-05 | -436 | -638 | -851 | -1064
1.58e-05 | -448 | -662 | -874 | -1087 |  | 1.58e-05 | -448 | -662 | -874 | -1087
1.6e-05 | -460 | -685 | -898 | -1135 |  | 1.6e-05 | -460 | -685 | -922 | -1135
1.62e-05 | -471 | -709 | -945 | -1158 |  | 1.62e-05 | -471 | -709 | -945 | -1158
1.64e-05 | -483 | -733 | -969 | -1182 |  | 1.64e-05 | -495 | -733 | -969 | -1206
1.66e-05 | -495 | -756 | -993 | -1229 |  | 1.66e-05 | -507 | -756 | -993 | -1229
1.68e-05 | -519 | -780 | -1016 | -1253 |  | 1.68e-05 | -519 | -780 | -1016 | -1253
1.7e-05 | -530 | -780 | -1040 | -1300 |  | 1.7e-05 | -530 | -803 | -1040 | -1300
1.72e-05 | -542 | -803 | -1064 | -1324 |  | 1.72e-05 | -542 | -803 | -1064 | -1324
1.74e-05 | -554 | -827 | -1087 | -1348 |  | 1.74e-05 | -554 | -827 | -1087 | -1348
1.76e-05 | -566 | -851 | -1111 | -1371 |  | 1.76e-05 | -566 | -851 | -1111 | -1371
1.78e-05 | -578 | -874 | -1135 | -1395 |  | 1.78e-05 | -578 | -874 | -1135 | -1395
1.8e-05 | -590 | -874 | -1158 | -1442 |  | 1.8e-05 | -590 | -874 | -1158 | -1442
1.82e-05 | -601 | -898 | -1182 | -1442 |  | 1.82e-05 | -601 | -898 | -1182 | -1442
1.84e-05 | -601 | -922 | -1182 | -1466 |  | 1.84e-05 | -601 | -922 | -1182 | -1490
1.86e-05 | -613 | -922 | -1206 | -1490 |  | 1.86e-05 | -613 | -922 | -1206 | -1490
1.88e-05 | -625 | -922 | -1229 | -1513 |  | 1.88e-05 | -625 | -922 | -1229 | -1513
1.9e-05 | -637 | -945 | -1253 | -1537 |  | 1.9e-05 | -625 | -945 | -1253 | -1537
1.92e-05 | -637 | -969 | -1253 | -1537 |  | 1.92e-05 | -637 | -945 | -1253 | -1537
1.94e-05 | -649 | -969 | -1253 | -1561 |  | 1.94e-05 | -649 | -969 | -1277 | -1561
1.96e-05 | -649 | -969 | -1277 | -1584 |  | 1.96e-05 | -649 | -969 | -1277 | -1584
1.98e-05 | -660 | -969 | -1277 | -1584 |  | 1.98e-05 | -649 | -969 | -1277 | -1584
2e-05 | -660 | -993 | -1300 | -1584 |  | 2e-05 | -660 | -993 | -1300 | -1608
2.02e-05 | -660 | -993 | -1300 | -1608 |  | 2.02e-05 | -660 | -993 | -1300 | -1608
2.04e-05 | -660 | -993 | -1300 | -1608 |  | 2.04e-05 | -660 | -993 | -1300 | -1632
2.06e-05 | -660 | -993 | -1300 | -1632 |  | 2.06e-05 | -672 | -993 | -1324 | -1632
2.08e-05 | -672 | -1016 | -1324 | -1632 |  | 2.08e-05 | -672 | -1016 | -1324 | -1632
2.1e-05 | -672 | -993 | -1324 | -1632 |  | 2.1e-05 | -672 | -993 | -1324 | -1632
2.12e-05 | -672 | -1016 | -1324 | -1632 |  | 2.12e-05 | -672 | -1016 | -1324 | -1632
2.14e-05 | -672 | -1016 | -1324 | -1632 |  | 2.14e-05 | -672 | -1016 | -1324 | -1632
2.16e-05 | -672 | -993 | -1324 | -1632 |  | 2.16e-05 | -672 | -993 | -1324 | -1632
2.18e-05 | -672 | -993 | -1300 | -1632 |  | 2.18e-05 | -672 | -993 | -1300 | -1632
2.2e-05 | -660 | -993 | -1300 | -1608 |  | 2.2e-05 | -660 | -993 | -1300 | -1632
2.22e-05 | -660 | -993 | -1300 | -1608 |  | 2.22e-05 | -660 | -993 | -1300 | -1584
2.24e-05 | -660 | -969 | -1277 | -1584 |  | 2.24e-05 | -660 | -969 | -1300 | -1584
2.26e-05 | -649 | -969 | -1277 | -1561 |  | 2.26e-05 | -649 | -969 | -1277 | -1561
2.28e-05 | -649 | -969 | -1253 | -1537 |  | 2.28e-05 | -637 | -969 | -1253 | -1537
2.3e-05 | -637 | -945 | -1229 | -1513 |  | 2.3e-05 | -637 | -945 | -1253 | -1513
2.32e-05 | -625 | -945 | -1206 | -1490 |  | 2.32e-05 | -625 | -922 | -1206 | -1513
2.34e-05 | -613 | -922 | -1182 | -1466 |  | 2.34e-05 | -613 | -922 | -1182 | -1466
2.36e-05 | -601 | -898 | -1182 | -1442 |  | 2.36e-05 | -601 | -898 | -1158 | -1442
2.38e-05 | -590 | -874 | -1135 | -1419 |  | 2.38e-05 | -590 | -874 | -1135 | -1419
2.4e-05 | -578 | -851 | -1111 | -1395 |  | 2.4e-05 | -578 | -851 | -1135 | -1395
2.42e-05 | -566 | -827 | -1087 | -1348 |  | 2.42e-05 | -566 | -827 | -1111 | -1348
2.44e-05 | -554 | -827 | -1064 | -1324 |  | 2.44e-05 | -554 | -803 | -1064 | -1324
2.46e-05 | -530 | -803 | -1040 | -1277 |  | 2.46e-05 | -530 | -780 | -1040 | -1277
2.48e-05 | -519 | -780 | -1016 | -1253 |  | 2.48e-05 | -519 | -756 | -1016 | -1253
2.5e-05 | -507 | -756 | -969 | -1206 |  | 2.5e-05 | -507 | -756 | -969 | -1206
2.52e-05 | -483 | -709 | -945 | -1158 |  | 2.52e-05 | -483 | -733 | -945 | -1158
2.54e-05 | -471 | -685 | -898 | -1111 |  | 2.54e-05 | -471 | -685 | -922 | -1111
2.56e-05 | -448 | -662 | -874 | -1064 |  | 2.56e-05 | -448 | -662 | -874 | -1064
2.58e-05 | -424 | -638 | -827 | -1016 |  | 2.58e-05 | -436 | -638 | -827 | -1016
2.6e-05 | -412 | -591 | -780 | -969 |  | 2.6e-05 | -412 | -591 | -780 | -969
2.62e-05 | -389 | -567 | -756 | -922 |  | 2.62e-05 | -389 | -567 | -756 | -922
2.64e-05 | -365 | -543 | -709 | -874 |  | 2.64e-05 | -365 | -543 | -709 | -874
2.66e-05 | -342 | -520 | -662 | -827 |  | 2.66e-05 | -353 | -520 | -662 | -827
2.68e-05 | -318 | -496 | -614 | -756 |  | 2.68e-05 | -330 | -472 | -614 | -780
2.7e-05 | -294 | -425 | -591 | -709 |  | 2.7e-05 | -306 | -449 | -591 | -709
2.72e-05 | -282 | -401 | -543 | -662 |  | 2.72e-05 | -282 | -425 | -543 | -662
2.74e-05 | -259 | -378 | -496 | -591 |  | 2.74e-05 | -259 | -378 | -496 | -614
2.76e-05 | -235 | -354 | -449 | -543 |  | 2.76e-05 | -235 | -354 | -449 | -543
2.78e-05 | -212 | -307 | -401 | -496 |  | 2.78e-05 | -212 | -307 | -401 | -496
2.8e-05 | -188 | -283 | -354 | -449 |  | 2.8e-05 | -188 | -283 | -354 | -449
2.82e-05 | -164 | -236 | -330 | -378 |  | 2.82e-05 | -164 | -259 | -307 | -378
2.84e-05 | -153 | -212 | -283 | -330 |  | 2.84e-05 | -141 | -212 | -283 | -330
2.86e-05 | -129 | -188 | -236 | -283 |  | 2.86e-05 | -129 | -188 | -236 | -283
2.88e-05 | -105 | -141 | -188 | -212 |  | 2.88e-05 | -93 | -141 | -188 | -212
2.9e-05 | -70 | -117 | -141 | -165 |  | 2.9e-05 | -70 | -117 | -141 | -165
2.92e-05 | -58 | -70 | -94 | -117 |  | 2.92e-05 | -58 | -70 | -94 | -117
2.94e-05 | -23 | -46 | -46 | -46 |  | 2.94e-05 | -23 | -46 | -46 | -46
2.96e-05 | -11 | 0 | 0 | 0 |  | 2.96e-05 | -11 | 0 | 0 | 0
2.98e-05 | 12 | 24 | 47 | 71 |  | 2.98e-05 | 24 | 24 | 47 | 47
3e-05 | 35 | 71 | 95 | 118 |  | 3e-05 | 35 | 47 | 95 | 118
3.02e-05 | 59 | 95 | 118 | 166 |  | 3.02e-05 | 59 | 95 | 142 | 166
3.04e-05 | 83 | 118 | 166 | 213 |  | 3.04e-05 | 83 | 118 | 166 | 213
3.06e-05 | 106 | 166 | 213 | 260 |  | 3.06e-05 | 106 | 166 | 213 | 260
3.08e-05 | 118 | 189 | 237 | 308 |  | 3.08e-05 | 130 | 189 | 237 | 308
3.1e-05 | 142 | 213 | 284 | 355 |  | 3.1e-05 | 142 | 213 | 284 | 355
3.12e-05 | 165 | 237 | 331 | 402 |  | 3.12e-05 | 165 | 237 | 331 | 426
3.14e-05 | 177 | 260 | 355 | 450 |  | 3.14e-05 | 177 | 260 | 379 | 450
3.16e-05 | 201 | 308 | 402 | 497 |  | 3.16e-05 | 201 | 308 | 402 | 521
3.18e-05 | 213 | 331 | 426 | 544 |  | 3.18e-05 | 224 | 331 | 426 | 544
3.2e-05 | 236 | 355 | 473 | 592 |  | 3.2e-05 | 236 | 355 | 473 | 592
3.22e-05 | 248 | 379 | 497 | 615 |  | 3.22e-05 | 260 | 379 | 521 | 639
3.24e-05 | 272 | 402 | 544 | 663 |  | 3.24e-05 | 272 | 402 | 544 | 663
3.26e-05 | 283 | 426 | 568 | 710 |  | 3.26e-05 | 283 | 426 | 568 | 710
3.28e-05 | 307 | 450 | 592 | 734 |  | 3.28e-05 | 307 | 450 | 592 | 734
3.3e-05 | 319 | 473 | 615 | 781 |  | 3.3e-05 | 319 | 473 | 615 | 781
3.32e-05 | 331 | 497 | 639 | 804 |  | 3.32e-05 | 331 | 497 | 639 | 804
3.34e-05 | 343 | 521 | 663 | 852 |  | 3.34e-05 | 343 | 521 | 686 | 852
3.36e-05 | 354 | 521 | 710 | 875 |  | 3.36e-05 | 354 | 521 | 710 | 875
3.38e-05 | 366 | 544 | 710 | 899 |  | 3.38e-05 | 366 | 544 | 734 | 899
3.4e-05 | 378 | 568 | 757 | 923 |  | 3.4e-05 | 378 | 568 | 734 | 923
3.42e-05 | 390 | 568 | 757 | 946 |  | 3.42e-05 | 390 | 568 | 757 | 946
3.44e-05 | 402 | 592 | 781 | 970 |  | 3.44e-05 | 402 | 592 | 781 | 970
3.46e-05 | 402 | 592 | 804 | 994 |  | 3.46e-05 | 402 | 592 | 804 | 994
3.48e-05 | 413 | 615 | 804 | 1017 |  | 3.48e-05 | 413 | 615 | 804 | 1017
3.5e-05 | 413 | 615 | 828 | 1017 |  | 3.5e-05 | 425 | 615 | 828 | 1017
3.52e-05 | 425 | 639 | 828 | 1041 |  | 3.52e-05 | 425 | 639 | 828 | 1041
3.54e-05 | 437 | 639 | 852 | 1041 |  | 3.54e-05 | 437 | 639 | 852 | 1041
3.56e-05 | 437 | 639 | 852 | 1065 |  | 3.56e-05 | 437 | 639 | 852 | 1065
3.58e-05 | 437 | 663 | 852 | 1065 |  | 3.58e-05 | 437 | 663 | 875 | 1065
3.6e-05 | 449 | 663 | 875 | 1065 |  | 3.6e-05 | 449 | 663 | 875 | 1065
3.62e-05 | 449 | 663 | 875 | 1065 |  | 3.62e-05 | 449 | 663 | 875 | 1065
3.64e-05 | 449 | 663 | 875 | 1065 |  | 3.64e-05 | 449 | 663 | 875 | 1088
3.66e-05 | 449 | 663 | 875 | 1088 |  | 3.66e-05 | 449 | 663 | 875 | 1088
3.68e-05 | 449 | 663 | 875 | 1088 |  | 3.68e-05 | 449 | 663 | 875 | 1065
3.7e-05 | 449 | 663 | 875 | 1065 |  | 3.7e-05 | 449 | 663 | 875 | 1088
3.72e-05 | 449 | 663 | 875 | 1065 |  | 3.72e-05 | 449 | 663 | 875 | 1065
3.74e-05 | 449 | 639 | 875 | 1065 |  | 3.74e-05 | 449 | 663 | 875 | 1065
3.76e-05 | 437 | 639 | 852 | 1065 |  | 3.76e-05 | 437 | 639 | 852 | 1065
3.78e-05 | 437 | 639 | 852 | 1041 |  | 3.78e-05 | 437 | 639 | 852 | 1041
3.8e-05 | 437 | 639 | 828 | 1041 |  | 3.8e-05 | 437 | 639 | 828 | 1041
3.82e-05 | 425 | 615 | 828 | 1017 |  | 3.82e-05 | 425 | 615 | 828 | 1017
3.84e-05 | 413 | 615 | 804 | 1017 |  | 3.84e-05 | 425 | 615 | 804 | 994
3.86e-05 | 413 | 615 | 804 | 994 |  | 3.86e-05 | 413 | 615 | 804 | 970
3.88e-05 | 402 | 592 | 781 | 970 |  | 3.88e-05 | 402 | 592 | 781 | 970
3.9e-05 | 402 | 592 | 757 | 946 |  | 3.9e-05 | 390 | 592 | 757 | 946
3.92e-05 | 390 | 568 | 757 | 923 |  | 3.92e-05 | 390 | 568 | 757 | 923
3.94e-05 | 378 | 568 | 734 | 899 |  | 3.94e-05 | 378 | 568 | 734 | 899
3.96e-05 | 366 | 544 | 710 | 875 |  | 3.96e-05 | 366 | 544 | 710 | 875
3.98e-05 | 354 | 544 | 686 | 852 |  | 3.98e-05 | 354 | 521 | 686 | 852
4e-05 | 343 | 521 | 686 | 828 |  | 4e-05 | 343 | 521 | 663 | 828
4.02e-05 | 331 | 497 | 639 | 804 |  | 4.02e-05 | 343 | 497 | 639 | 804
4.04e-05 | 319 | 473 | 615 | 781 |  | 4.04e-05 | 319 | 473 | 615 | 781
4.06e-05 | 307 | 450 | 615 | 734 |  | 4.06e-05 | 307 | 450 | 592 | 734
4.08e-05 | 295 | 426 | 568 | 710 |  | 4.08e-05 | 295 | 450 | 568 | 710
4.1e-05 | 283 | 426 | 544 | 686 |  | 4.1e-05 | 283 | 402 | 544 | 686
4.12e-05 | 272 | 402 | 521 | 639 |  | 4.12e-05 | 272 | 402 | 521 | 639
4.14e-05 | 260 | 379 | 497 | 615 |  | 4.14e-05 | 260 | 379 | 497 | 615
4.16e-05 | 248 | 355 | 473 | 568 |  | 4.16e-05 | 248 | 355 | 473 | 568
4.18e-05 | 224 | 331 | 426 | 544 |  | 4.18e-05 | 224 | 331 | 450 | 544
4.2e-05 | 224 | 308 | 402 | 497 |  | 4.2e-05 | 213 | 308 | 402 | 497
4.22e-05 | 201 | 284 | 379 | 473 |  | 4.22e-05 | 201 | 284 | 379 | 473
4.24e-05 | 189 | 260 | 355 | 426 |  | 4.24e-05 | 189 | 260 | 355 | 426
4.26e-05 | 177 | 237 | 308 | 379 |  | 4.26e-05 | 165 | 237 | 331 | 402
4.28e-05 | 154 | 237 | 284 | 355 |  | 4.28e-05 | 154 | 213 | 284 | 355
4.3e-05 | 142 | 213 | 260 | 308 |  | 4.3e-05 | 142 | 213 | 260 | 308
4.32e-05 | 118 | 189 | 237 | 284 |  | 4.32e-05 | 118 | 189 | 237 | 284
4.34e-05 | 118 | 166 | 213 | 237 |  | 4.34e-05 | 106 | 166 | 213 | 237
4.36e-05 | 94 | 142 | 166 | 213 |  | 4.36e-05 | 94 | 142 | 166 | 213
4.38e-05 | 83 | 118 | 142 | 166 |  | 4.38e-05 | 83 | 118 | 142 | 166
4.4e-05 | 71 | 95 | 118 | 142 |  | 4.4e-05 | 59 | 71 | 95 | 142
4.42e-05 | 47 | 71 | 95 | 95 |  | 4.42e-05 | 47 | 71 | 71 | 95
4.44e-05 | 35 | 47 | 47 | 47 |  | 4.44e-05 | 35 | 47 | 47 | 47
4.46e-05 | 24 | 24 | 24 | 24 |  | 4.46e-05 | 12 | 24 | 24 | 24
4.48e-05 | 0 | 0 | 0 | -23 |  | 4.48e-05 | 0 | 0 | -23 | -23
4.5e-05 | -11 | -23 | -46 | -46 |  | 4.5e-05 | -11 | -23 | -46 | -46
4.52e-05 | -34 | -46 | -70 | -94 |  | 4.52e-05 | -34 | -46 | -70 | -94
4.54e-05 | -46 | -70 | -117 | -117 |  | 4.54e-05 | -46 | -70 | -94 | -117
4.56e-05 | -58 | -94 | -117 | -165 |  | 4.56e-05 | -58 | -94 | -141 | -165
4.58e-05 | -70 | -117 | -165 | -188 |  | 4.58e-05 | -70 | -117 | -165 | -188
4.6e-05 | -82 | -141 | -188 | -236 |  | 4.6e-05 | -93 | -141 | -188 | -236
4.62e-05 | -105 | -165 | -212 | -259 |  | 4.62e-05 | -105 | -165 | -212 | -259
4.64e-05 | -117 | -188 | -236 | -307 |  | 4.64e-05 | -117 | -188 | -236 | -307
4.66e-05 | -129 | -188 | -259 | -330 |  | 4.66e-05 | -129 | -212 | -259 | -330
4.68e-05 | -141 | -212 | -283 | -354 |  | 4.68e-05 | -141 | -236 | -283 | -354
4.7e-05 | -153 | -236 | -307 | -378 |  | 4.7e-05 | -153 | -236 | -307 | -378
4.72e-05 | -164 | -259 | -330 | -401 |  | 4.72e-05 | -164 | -259 | -330 | -425
4.74e-05 | -176 | -283 | -354 | -449 |  | 4.74e-05 | -176 | -283 | -354 | -449
4.76e-05 | -188 | -283 | -378 | -472 |  | 4.76e-05 | -188 | -307 | -378 | -472
4.78e-05 | -200 | -307 | -401 | -496 |  | 4.78e-05 | -200 | -307 | -401 | -496
4.8e-05 | -200 | -307 | -401 | -520 |  | 4.8e-05 | -212 | -330 | -425 | -520
4.82e-05 | -223 | -330 | -449 | -543 |  | 4.82e-05 | -212 | -330 | -449 | -543
4.84e-05 | -223 | -354 | -449 | -567 |  | 4.84e-05 | -223 | -354 | -449 | -567
4.86e-05 | -235 | -354 | -472 | -591 |  | 4.86e-05 | -235 | -354 | -472 | -591
4.88e-05 | -247 | -378 | -496 | -614 |  | 4.88e-05 | -247 | -378 | -496 | -614
4.9e-05 | -247 | -378 | -496 | -614 |  | 4.9e-05 | -247 | -378 | -520 | -638
4.92e-05 | -259 | -378 | -520 | -638 |  | 4.92e-05 | -259 | -401 | -520 | -638
4.94e-05 | -271 | -401 | -520 | -662 |  | 4.94e-05 | -271 | -401 | -543 | -662
4.96e-05 | -271 | -401 | -543 | -662 |  | 4.96e-05 | -271 | -425 | -543 | -685
4.98e-05 | -282 | -425 | -543 | -685 |  | 4.98e-05 | -282 | -425 | -567 | -685
5e-05 | -282 | -425 | -567 | -709 |  | 5e-05 | -282 | -425 | -567 | -709
5.02e-05 | -294 | -425 | -567 | -709 |  | 5.02e-05 | -282 | -425 | -567 | -709
5.04e-05 | -294 | -449 | -591 | -709 |  | 5.04e-05 | -294 | -449 | -591 | -733
5.06e-05 | -294 | -449 | -591 | -733 |  | 5.06e-05 | -306 | -449 | -591 | -733
5.08e-05 | -294 | -449 | -591 | -733 |  | 5.08e-05 | -306 | -449 | -591 | -733
5.1e-05 | -306 | -449 | -591 | -733 |  | 5.1e-05 | -306 | -449 | -591 | -733
5.12e-05 | -306 | -449 | -591 | -756 |  | 5.12e-05 | -306 | -449 | -614 | -756
5.14e-05 | -306 | -472 | -591 | -756 |  | 5.14e-05 | -306 | -449 | -614 | -756
5.16e-05 | -306 | -449 | -614 | -756 |  | 5.16e-05 | -306 | -449 | -614 | -756
5.18e-05 | -306 | -472 | -614 | -756 |  | 5.18e-05 | -306 | -472 | -614 | -756
5.2e-05 | -306 | -472 | -614 | -756 |  | 5.2e-05 | -306 | -449 | -614 | -756
5.22e-05 | -306 | -449 | -614 | -756 |  | 5.22e-05 | -306 | -472 | -614 | -756
5.24e-05 | -306 | -449 | -591 | -756 |  | 5.24e-05 | -306 | -449 | -614 | -733
5.26e-05 | -306 | -449 | -591 | -733 |  | 5.26e-05 | -306 | -449 | -591 | -733
5.28e-05 | -306 | -449 | -591 | -733 |  | 5.28e-05 | -306 | -449 | -591 | -733
5.3e-05 | -306 | -449 | -591 | -733 |  | 5.3e-05 | -294 | -449 | -591 | -733
5.32e-05 | -294 | -449 | -591 | -709 |  | 5.32e-05 | -294 | -449 | -591 | -733
5.34e-05 | -294 | -425 | -591 | -709 |  | 5.34e-05 | -294 | -449 | -567 | -709
5.36e-05 | -282 | -425 | -567 | -709 |  | 5.36e-05 | -282 | -449 | -567 | -709
5.38e-05 | -282 | -425 | -567 | -685 |  | 5.38e-05 | -282 | -425 | -567 | -685
5.4e-05 | -282 | -425 | -543 | -685 |  | 5.4e-05 | -282 | -425 | -543 | -685
5.42e-05 | -271 | -401 | -543 | -662 |  | 5.42e-05 | -271 | -401 | -543 | -662
5.44e-05 | -271 | -401 | -543 | -638 |  | 5.44e-05 | -271 | -401 | -520 | -638
5.46e-05 | -259 | -378 | -520 | -638 |  | 5.46e-05 | -259 | -401 | -520 | -638
5.48e-05 | -259 | -378 | -496 | -614 |  | 5.48e-05 | -259 | -378 | -496 | -614
5.5e-05 | -247 | -354 | -496 | -591 |  | 5.5e-05 | -247 | -378 | -496 | -591
5.52e-05 | -247 | -354 | -472 | -591 |  | 5.52e-05 | -247 | -354 | -472 | -591
5.54e-05 | -235 | -354 | -449 | -567 |  | 5.54e-05 | -235 | -354 | -449 | -567
5.56e-05 | -223 | -330 | -449 | -543 |  | 5.56e-05 | -223 | -330 | -449 | -543
5.58e-05 | -212 | -330 | -425 | -520 |  | 5.58e-05 | -212 | -330 | -425 | -520
5.6e-05 | -212 | -307 | -401 | -520 |  | 5.6e-05 | -212 | -307 | -401 | -496
5.62e-05 | -200 | -307 | -378 | -496 |  | 5.62e-05 | -200 | -307 | -378 | -496
5.64e-05 | -188 | -283 | -378 | -449 |  | 5.64e-05 | -188 | -283 | -378 | -449
5.66e-05 | -176 | -283 | -354 | -449 |  | 5.66e-05 | -176 | -283 | -354 | -425
5.68e-05 | -176 | -259 | -330 | -425 |  | 5.68e-05 | -176 | -259 | -330 | -425
5.7e-05 | -164 | -236 | -307 | -378 |  | 5.7e-05 | -164 | -236 | -330 | -378
5.72e-05 | -153 | -236 | -307 | -354 |  | 5.72e-05 | -153 | -236 | -307 | -354
5.74e-05 | -141 | -212 | -283 | -354 |  | 5.74e-05 | -141 | -212 | -283 | -330
5.76e-05 | -141 | -212 | -259 | -330 |  | 5.76e-05 | -129 | -212 | -259 | -307
5.78e-05 | -129 | -188 | -236 | -283 |  | 5.78e-05 | -129 | -188 | -236 | -307
5.8e-05 | -117 | -165 | -212 | -259 |  | 5.8e-05 | -117 | -165 | -212 | -259
5.82e-05 | -105 | -141 | -212 | -236 |  | 5.82e-05 | -105 | -165 | -212 | -236
5.84e-05 | -93 | -141 | -188 | -212 |  | 5.84e-05 | -93 | -141 | -188 | -212
5.86e-05 | -82 | -117 | -165 | -188 |  | 5.86e-05 | -82 | -117 | -141 | -188
5.88e-05 | -70 | -117 | -141 | -165 |  | 5.88e-05 | -70 | -94 | -117 | -165
5.9e-05 | -58 | -94 | -117 | -141 |  | 5.9e-05 | -58 | -94 | -117 | -141
5.92e-05 | -58 | -70 | -94 | -117 |  | 5.92e-05 | -46 | -70 | -94 | -117
5.94e-05 | -34 | -46 | -70 | -94 |  | 5.94e-05 | -34 | -46 | -70 | -94
5.96e-05 | -23 | -46 | -46 | -70 |  | 5.96e-05 | -23 | -46 | -46 | -46
5.98e-05 | -23 | -23 | -23 | -46 |  | 5.98e-05 | -11 | -23 | -23 | -23
6e-05 | -11 | 0 | -23 | 0 |  | 6e-05 | -11 | -23 | 0 | 0
6.02e-05 | 0 | 0 | 0 | 24 |  | 6.02e-05 | 0 | 0 | 24 | 24
6.04e-05 | 12 | 24 | 24 | 47 |  | 6.04e-05 | 12 | 24 | 24 | 47
6.06e-05 | 24 | 47 | 47 | 71 |  | 6.06e-05 | 24 | 47 | 47 | 71
6.08e-05 | 35 | 47 | 71 | 95 |  | 6.08e-05 | 35 | 47 | 71 | 95
6.1e-05 | 47 | 71 | 95 | 118 |  | 6.1e-05 | 47 | 71 | 95 | 118
6.12e-05 | 59 | 71 | 118 | 142 |  | 6.12e-05 | 59 | 71 | 118 | 142
6.14e-05 | 59 | 95 | 118 | 166 |  | 6.14e-05 | 59 | 95 | 142 | 166
6.16e-05 | 71 | 95 | 142 | 189 |  | 6.16e-05 | 71 | 118 | 142 | 189
6.18e-05 | 83 | 118 | 166 | 213 |  | 6.18e-05 | 83 | 118 | 166 | 189
6.2e-05 | 94 | 142 | 189 | 213 |  | 6.2e-05 | 94 | 142 | 189 | 237
6.22e-05 | 94 | 142 | 189 | 237 |  | 6.22e-05 | 94 | 142 | 189 | 237
6.24e-05 | 106 | 166 | 213 | 260 |  | 6.24e-05 | 106 | 166 | 213 | 260
6.26e-05 | 118 | 166 | 237 | 284 |  | 6.26e-05 | 118 | 166 | 237 | 284
6.28e-05 | 118 | 189 | 237 | 308 |  | 6.28e-05 | 118 | 189 | 237 | 284
6.3e-05 | 130 | 189 | 260 | 308 |  | 6.3e-05 | 130 | 189 | 260 | 331
6.32e-05 | 130 | 213 | 260 | 331 |  | 6.32e-05 | 130 | 213 | 260 | 331
6.34e-05 | 142 | 213 | 284 | 355 |  | 6.34e-05 | 142 | 213 | 284 | 355
6.36e-05 | 142 | 213 | 284 | 355 |  | 6.36e-05 | 154 | 237 | 284 | 355
6.38e-05 | 154 | 237 | 308 | 379 |  | 6.38e-05 | 154 | 237 | 308 | 379
6.4e-05 | 154 | 237 | 308 | 402 |  | 6.4e-05 | 165 | 237 | 308 | 402
6.42e-05 | 165 | 237 | 331 | 402 |  | 6.42e-05 | 165 | 237 | 331 | 402
6.44e-05 | 165 | 260 | 331 | 402 |  | 6.44e-05 | 177 | 260 | 331 | 426
6.46e-05 | 177 | 260 | 331 | 426 |  | 6.46e-05 | 177 | 260 | 355 | 426
6.48e-05 | 177 | 260 | 355 | 426 |  | 6.48e-05 | 177 | 260 | 355 | 426
6.5e-05 | 189 | 260 | 355 | 450 |  | 6.5e-05 | 189 | 260 | 355 | 450
6.52e-05 | 189 | 284 | 379 | 450 |  | 6.52e-05 | 189 | 284 | 379 | 450
6.54e-05 | 189 | 284 | 379 | 473 |  | 6.54e-05 | 189 | 284 | 379 | 473
6.56e-05 | 189 | 284 | 379 | 473 |  | 6.56e-05 | 201 | 284 | 379 | 473
6.58e-05 | 201 | 284 | 379 | 473 |  | 6.58e-05 | 201 | 284 | 402 | 473
6.6e-05 | 201 | 284 | 402 | 497 |  | 6.6e-05 | 201 | 308 | 379 | 497
6.62e-05 | 201 | 284 | 402 | 497 |  | 6.62e-05 | 201 | 284 | 402 | 497
6.64e-05 | 201 | 308 | 402 | 497 |  | 6.64e-05 | 201 | 308 | 402 | 497
6.66e-05 | 201 | 308 | 402 | 497 |  | 6.66e-05 | 201 | 308 | 402 | 497
6.68e-05 | 201 | 308 | 402 | 497 |  | 6.68e-05 | 201 | 308 | 402 | 497
6.7e-05 | 213 | 308 | 402 | 497 |  | 6.7e-05 | 201 | 308 | 402 | 497
6.72e-05 | 201 | 308 | 402 | 497 |  | 6.72e-05 | 201 | 308 | 402 | 497
6.74e-05 | 201 | 308 | 402 | 497 |  | 6.74e-05 | 213 | 308 | 402 | 497
6.76e-05 | 213 | 308 | 402 | 497 |  | 6.76e-05 | 201 | 308 | 402 | 497
6.78e-05 | 201 | 308 | 402 | 497 |  | 6.78e-05 | 201 | 308 | 402 | 497
6.8e-05 | 201 | 308 | 402 | 497 |  | 6.8e-05 | 201 | 284 | 402 | 497
6.82e-05 | 201 | 284 | 379 | 473 |  | 6.82e-05 | 201 | 308 | 402 | 497
6.84e-05 | 201 | 284 | 379 | 473 |  | 6.84e-05 | 201 | 284 | 379 | 473
6.86e-05 | 201 | 284 | 379 | 473 |  | 6.86e-05 | 201 | 284 | 379 | 473
6.88e-05 | 201 | 284 | 379 | 473 |  | 6.88e-05 | 189 | 284 | 379 | 473
6.9e-05 | 189 | 284 | 379 | 450 |  | 6.9e-05 | 189 | 284 | 355 | 473
6.92e-05 | 189 | 284 | 355 | 450 |  | 6.92e-05 | 189 | 260 | 355 | 450
6.94e-05 | 189 | 260 | 355 | 450 |  | 6.94e-05 | 189 | 260 | 355 | 450
6.96e-05 | 177 | 260 | 355 | 426 |  | 6.96e-05 | 177 | 260 | 355 | 426
6.98e-05 | 177 | 260 | 331 | 426 |  | 6.98e-05 | 177 | 260 | 331 | 426
7e-05 | 177 | 260 | 331 | 402 |  | 7e-05 | 177 | 260 | 331 | 402
7.02e-05 | 165 | 237 | 331 | 402 |  | 7.02e-05 | 165 | 260 | 331 | 402
7.04e-05 | 165 | 237 | 308 | 379 |  | 7.04e-05 | 165 | 237 | 308 | 379
7.06e-05 | 154 | 237 | 308 | 379 |  | 7.06e-05 | 154 | 237 | 308 | 379
7.08e-05 | 154 | 237 | 284 | 355 |  | 7.08e-05 | 154 | 213 | 284 | 355
7.1e-05 | 154 | 213 | 284 | 355 |  | 7.1e-05 | 154 | 213 | 284 | 355
7.12e-05 | 142 | 213 | 260 | 331 |  | 7.12e-05 | 142 | 213 | 260 | 331
7.14e-05 | 130 | 189 | 260 | 331 |  | 7.14e-05 | 142 | 189 | 260 | 331
7.16e-05 | 130 | 189 | 260 | 308 |  | 7.16e-05 | 130 | 189 | 237 | 308
7.18e-05 | 118 | 189 | 237 | 284 |  | 7.18e-05 | 118 | 189 | 237 | 284
7.2e-05 | 118 | 166 | 237 | 284 |  | 7.2e-05 | 118 | 166 | 237 | 284
7.22e-05 | 118 | 166 | 213 | 260 |  | 7.22e-05 | 106 | 166 | 213 | 260
7.24e-05 | 106 | 166 | 213 | 260 |  | 7.24e-05 | 106 | 166 | 189 | 260
7.26e-05 | 106 | 142 | 189 | 237 |  | 7.26e-05 | 94 | 142 | 189 | 237
7.28e-05 | 94 | 142 | 166 | 213 |  | 7.28e-05 | 94 | 142 | 189 | 213
7.3e-05 | 83 | 118 | 166 | 189 |  | 7.3e-05 | 83 | 118 | 166 | 189
7.32e-05 | 83 | 118 | 142 | 189 |  | 7.32e-05 | 83 | 118 | 142 | 189
7.34e-05 | 71 | 95 | 142 | 166 |  | 7.34e-05 | 71 | 118 | 142 | 166
7.36e-05 | 71 | 95 | 118 | 142 |  | 7.36e-05 | 59 | 95 | 118 | 142
7.38e-05 | 59 | 71 | 118 | 142 |  | 7.38e-05 | 59 | 71 | 118 | 118
7.4e-05 | 47 | 71 | 95 | 118 |  | 7.4e-05 | 47 | 71 | 95 | 118
7.42e-05 | 47 | 47 | 71 | 95 |  | 7.42e-05 | 47 | 71 | 71 | 95
7.44e-05 | 35 | 47 | 71 | 71 |  | 7.44e-05 | 35 | 47 | 71 | 71
7.46e-05 | 35 | 47 | 47 | 71 |  | 7.46e-05 | 24 | 47 | 47 | 47
7.48e-05 | 24 | 24 | 47 | 47 |  | 7.48e-05 | 24 | 24 | 24 | 47
7.5e-05 | 12 | 24 | 24 | 24 |  | 7.5e-05 | 12 | 24 | 24 | 24
7.52e-05 | 12 | 0 | 0 | 0 |  | 7.52e-05 | 0 | 0 | 0 | 0
7.54e-05 | 0 | 0 | 0 | 0 |  | 7.54e-05 | 0 | 0 | 0 | -23
7.56e-05 | 0 | -23 | -23 | -23 |  | 7.56e-05 | -11 | -23 | -23 | -23
7.58e-05 | -11 | -23 | -23 | -46 |  | 7.58e-05 | -11 | -23 | -46 | -46
7.6e-05 | -23 | -46 | -46 | -70 |  | 7.6e-05 | -23 | -46 | -46 | -70
7.62e-05 | -23 | -46 | -46 | -94 |  | 7.62e-05 | -23 | -46 | -70 | -94
7.64e-05 | -34 | -46 | -70 | -94 |  | 7.64e-05 | -34 | -46 | -70 | -94
7.66e-05 | -34 | -70 | -94 | -117 |  | 7.66e-05 | -46 | -70 | -94 | -117
7.68e-05 | -46 | -70 | -94 | -117 |  | 7.68e-05 | -46 | -70 | -94 | -141
7.7e-05 | -58 | -94 | -117 | -141 |  | 7.7e-05 | -58 | -94 | -117 | -141
7.72e-05 | -58 | -94 | -117 | -165 |  | 7.72e-05 | -58 | -94 | -117 | -165
7.74e-05 | -70 | -94 | -141 | -165 |  | 7.74e-05 | -70 | -94 | -141 | -165
7.76e-05 | -70 | -117 | -141 | -165 |  | 7.76e-05 | -70 | -117 | -141 | -188
7.78e-05 | -82 | -117 | -165 | -188 |  | 7.78e-05 | -82 | -117 | -165 | -188
7.8e-05 | -82 | -117 | -165 | -212 |  | 7.8e-05 | -82 | -141 | -165 | -212
7.82e-05 | -93 | -141 | -188 | -236 |  | 7.82e-05 | -93 | -141 | -188 | -212
7.84e-05 | -93 | -141 | -188 | -236 |  | 7.84e-05 | -93 | -141 | -188 | -236
7.86e-05 | -93 | -141 | -188 | -236 |  | 7.86e-05 | -105 | -141 | -212 | -236
7.88e-05 | -105 | -165 | -212 | -259 |  | 7.88e-05 | -105 | -165 | -212 | -259
7.9e-05 | -105 | -165 | -212 | -259 |  | 7.9e-05 | -105 | -165 | -212 | -283
7.92e-05 | -117 | -165 | -212 | -283 |  | 7.92e-05 | -117 | -165 | -236 | -283
7.94e-05 | -117 | -188 | -236 | -283 |  | 7.94e-05 | -117 | -165 | -236 | -307
7.96e-05 | -117 | -165 | -236 | -283 |  | 7.96e-05 | -117 | -188 | -236 | -307
7.98e-05 | -117 | -188 | -236 | -307 |  | 7.98e-05 | -117 | -188 | -236 | -307
8e-05 | -117 | -188 | -259 | -307 |  | 8e-05 | -129 | -188 | -259 | -307
8.02e-05 | -129 | -188 | -259 | -307 |  | 8.02e-05 | -129 | -188 | -259 | -307
8.04e-05 | -129 | -188 | -259 | -307 |  | 8.04e-05 | -129 | -188 | -259 | -330
8.06e-05 | -129 | -212 | -259 | -307 |  | 8.06e-05 | -141 | -212 | -259 | -330
8.08e-05 | -141 | -212 | -259 | -330 |  | 8.08e-05 | -141 | -212 | -259 | -330
8.1e-05 | -141 | -212 | -259 | -330 |  | 8.1e-05 | -141 | -212 | -283 | -330
8.12e-05 | -141 | -212 | -283 | -330 |  | 8.12e-05 | -141 | -212 | -283 | -330
8.14e-05 | -141 | -212 | -283 | -330 |  | 8.14e-05 | -141 | -212 | -283 | -354
8.16e-05 | -141 | -212 | -283 | -330 |  | 8.16e-05 | -141 | -212 | -283 | -330
8.18e-05 | -141 | -212 | -283 | -330 |  | 8.18e-05 | -141 | -212 | -283 | -354
8.2e-05 | -141 | -212 | -283 | -354 |  | 8.2e-05 | -153 | -212 | -283 | -354
8.22e-05 | -141 | -212 | -283 | -354 |  | 8.22e-05 | -141 | -212 | -283 | -354
8.24e-05 | -141 | -212 | -283 | -354 |  | 8.24e-05 | -141 | -212 | -283 | -354
8.26e-05 | -141 | -212 | -283 | -354 |  | 8.26e-05 | -141 | -212 | -283 | -354
8.28e-05 | -141 | -212 | -283 | -330 |  | 8.28e-05 | -141 | -212 | -283 | -354
8.3e-05 | -141 | -212 | -283 | -330 |  | 8.3e-05 | -141 | -212 | -283 | -354
8.32e-05 | -141 | -212 | -283 | -354 |  | 8.32e-05 | -141 | -212 | -283 | -354
8.34e-05 | -141 | -212 | -283 | -330 |  | 8.34e-05 | -141 | -212 | -283 | -354
8.36e-05 | -141 | -212 | -283 | -330 |  | 8.36e-05 | -141 | -212 | -283 | -330
8.38e-05 | -141 | -212 | -283 | -330 |  | 8.38e-05 | -141 | -212 | -283 | -330
8.4e-05 | -141 | -212 | -283 | -330 |  | 8.4e-05 | -141 | -212 | -283 | -330
8.42e-05 | -141 | -212 | -259 | -330 |  | 8.42e-05 | -129 | -212 | -259 | -330
8.44e-05 | -129 | -212 | -259 | -330 |  | 8.44e-05 | -129 | -188 | -259 | -330
8.46e-05 | -129 | -212 | -259 | -307 |  | 8.46e-05 | -129 | -212 | -259 | -307
8.48e-05 | -129 | -188 | -259 | -307 |  | 8.48e-05 | -129 | -188 | -259 | -307
8.5e-05 | -129 | -188 | -259 | -307 |  | 8.5e-05 | -129 | -188 | -259 | -307
8.52e-05 | -129 | -188 | -236 | -283 |  | 8.52e-05 | -117 | -188 | -236 | -307
8.54e-05 | -117 | -188 | -236 | -283 |  | 8.54e-05 | -117 | -188 | -236 | -283
8.56e-05 | -117 | -188 | -236 | -283 |  | 8.56e-05 | -117 | -188 | -236 | -283
8.58e-05 | -117 | -165 | -212 | -283 |  | 8.58e-05 | -117 | -165 | -236 | -283
8.6e-05 | -105 | -165 | -212 | -259 |  | 8.6e-05 | -105 | -165 | -212 | -259
8.62e-05 | -105 | -165 | -212 | -259 |  | 8.62e-05 | -105 | -165 | -212 | -259
8.64e-05 | -105 | -165 | -212 | -259 |  | 8.64e-05 | -105 | -165 | -212 | -259
8.66e-05 | -93 | -141 | -188 | -236 |  | 8.66e-05 | -93 | -165 | -188 | -236
8.68e-05 | -93 | -141 | -188 | -236 |  | 8.68e-05 | -93 | -141 | -188 | -236
8.7e-05 | -93 | -141 | -188 | -212 |  | 8.7e-05 | -93 | -141 | -165 | -212
8.72e-05 | -82 | -141 | -165 | -212 |  | 8.72e-05 | -82 | -141 | -165 | -212
8.74e-05 | -82 | -117 | -165 | -188 |  | 8.74e-05 | -82 | -117 | -165 | -188
8.76e-05 | -70 | -117 | -141 | -188 |  | 8.76e-05 | -82 | -117 | -141 | -188
8.78e-05 | -70 | -117 | -141 | -188 |  | 8.78e-05 | -70 | -117 | -141 | -165
8.8e-05 | -70 | -94 | -141 | -165 |  | 8.8e-05 | -70 | -94 | -141 | -165
8.82e-05 | -58 | -94 | -117 | -165 |  | 8.82e-05 | -58 | -94 | -117 | -141
8.84e-05 | -58 | -94 | -117 | -141 |  | 8.84e-05 | -58 | -94 | -117 | -141
8.86e-05 | -46 | -94 | -117 | -141 |  | 8.86e-05 | -46 | -94 | -94 | -117
8.88e-05 | -46 | -70 | -94 | -117 |  | 8.88e-05 | -46 | -70 | -94 | -117
8.9e-05 | -34 | -70 | -94 | -117 |  | 8.9e-05 | -46 | -70 | -94 | -94
8.92e-05 | -46 | -70 | -94 | -94 |  | 8.92e-05 | -34 | -46 | -70 | -94
8.94e-05 | -34 | -46 | -70 | -70 |  | 8.94e-05 | -34 | -46 | -70 | -70
8.96e-05 | -23 | -46 | -46 | -70 |  | 8.96e-05 | -23 | -46 | -46 | -70
8.98e-05 | -23 | -23 | -46 | -46 |  | 8.98e-05 | -23 | -23 | -46 | -70
9e-05 | -23 | -23 | -46 | -46 |  | 9e-05 | -23 | -23 | -23 | -46
9.02e-05 | -11 | -23 | -23 | -46 |  | 9.02e-05 | -11 | -23 | -23 | -23
9.04e-05 | -11 | -23 | -23 | -23 |  | 9.04e-05 | -11 | -23 | -23 | -23
9.06e-05 | 0 | 0 | 0 | 0 |  | 9.06e-05 | 0 | 0 | 0 | 0
9.08e-05 | 0 | 0 | 0 | 0 |  | 9.08e-05 | 0 | 0 | 0 | 0
9.1e-05 | 0 | 0 | 0 | 24 |  | 9.1e-05 | 0 | 0 | 0 | 24
9.12e-05 | 12 | 24 | 24 | 24 |  | 9.12e-05 | 12 | 24 | 24 | 24
9.14e-05 | 12 | 24 | 24 | 47 |  | 9.14e-05 | 12 | 24 | 24 | 47
9.16e-05 | 24 | 24 | 47 | 47 |  | 9.16e-05 | 24 | 24 | 47 | 47
9.18e-05 | 24 | 24 | 47 | 71 |  | 9.18e-05 | 24 | 47 | 47 | 71
9.2e-05 | 24 | 47 | 47 | 71 |  | 9.2e-05 | 24 | 47 | 47 | 71
9.22e-05 | 35 | 47 | 71 | 71 |  | 9.22e-05 | 35 | 47 | 71 | 95
9.24e-05 | 35 | 47 | 71 | 95 |  | 9.24e-05 | 35 | 47 | 71 | 95
9.26e-05 | 47 | 47 | 71 | 95 |  | 9.26e-05 | 47 | 47 | 95 | 118
9.28e-05 | 47 | 71 | 95 | 118 |  | 9.28e-05 | 47 | 71 | 95 | 118
9.3e-05 | 47 | 71 | 95 | 118 |  | 9.3e-05 | 47 | 71 | 95 | 118
9.32e-05 | 59 | 71 | 95 | 118 |  | 9.32e-05 | 59 | 71 | 118 | 142
9.34e-05 | 59 | 95 | 118 | 142 |  | 9.34e-05 | 59 | 71 | 118 | 142
9.36e-05 | 59 | 95 | 118 | 142 |  | 9.36e-05 | 59 | 71 | 118 | 142
9.38e-05 | 59 | 95 | 118 | 142 |  | 9.38e-05 | 71 | 95 | 118 | 142
9.4e-05 | 71 | 95 | 142 | 166 |  | 9.4e-05 | 71 | 95 | 118 | 166
9.42e-05 | 71 | 95 | 142 | 166 |  | 9.42e-05 | 71 | 95 | 142 | 166
9.44e-05 | 71 | 95 | 142 | 166 |  | 9.44e-05 | 71 | 118 | 142 | 166
9.46e-05 | 83 | 118 | 142 | 189 |  | 9.46e-05 | 83 | 118 | 142 | 189
9.48e-05 | 83 | 118 | 142 | 189 |  | 9.48e-05 | 83 | 118 | 166 | 189
9.5e-05 | 83 | 118 | 166 | 189 |  | 9.5e-05 | 83 | 118 | 166 | 189
9.52e-05 | 83 | 118 | 166 | 213 |  | 9.52e-05 | 83 | 118 | 166 | 213
9.54e-05 | 83 | 118 | 166 | 213 |  | 9.54e-05 | 94 | 118 | 166 | 213
9.56e-05 | 83 | 142 | 166 | 213 |  | 9.56e-05 | 83 | 118 | 166 | 213
9.58e-05 | 83 | 142 | 166 | 213 |  | 9.58e-05 | 94 | 142 | 166 | 213
9.6e-05 | 94 | 118 | 166 | 213 |  | 9.6e-05 | 94 | 142 | 166 | 213
9.62e-05 | 94 | 142 | 189 | 213 |  | 9.62e-05 | 83 | 142 | 189 | 213
9.64e-05 | 94 | 142 | 189 | 237 |  | 9.64e-05 | 94 | 142 | 189 | 213
9.66e-05 | 94 | 142 | 189 | 237 |  | 9.66e-05 | 94 | 142 | 189 | 237
9.68e-05 | 94 | 142 | 189 | 237 |  | 9.68e-05 | 94 | 142 | 189 | 237
9.7e-05 | 94 | 142 | 189 | 237 |  | 9.7e-05 | 94 | 142 | 189 | 237
9.72e-05 | 94 | 142 | 189 | 237 |  | 9.72e-05 | 94 | 142 | 189 | 237
9.74e-05 | 94 | 142 | 189 | 237 |  | 9.74e-05 | 94 | 142 | 189 | 237
9.76e-05 | 94 | 142 | 189 | 237 |  | 9.76e-05 | 94 | 142 | 189 | 237
9.78e-05 | 94 | 142 | 189 | 237 |  | 9.78e-05 | 94 | 142 | 189 | 237
9.8e-05 | 106 | 142 | 189 | 237 |  | 9.8e-05 | 94 | 142 | 189 | 237
9.82e-05 | 94 | 142 | 189 | 237 |  | 9.82e-05 | 94 | 142 | 189 | 237
9.84e-05 | 94 | 142 | 189 | 237 |  | 9.84e-05 | 94 | 142 | 189 | 237
9.86e-05 | 94 | 142 | 189 | 237 |  | 9.86e-05 | 94 | 142 | 189 | 237
9.88e-05 | 94 | 142 | 189 | 237 |  | 9.88e-05 | 94 | 142 | 189 | 213
9.9e-05 | 94 | 142 | 189 | 213 |  | 9.9e-05 | 94 | 142 | 189 | 237
9.92e-05 | 94 | 142 | 189 | 213 |  | 9.92e-05 | 94 | 142 | 189 | 213
9.94e-05 | 94 | 142 | 189 | 213 |  | 9.94e-05 | 94 | 142 | 189 | 213
9.96e-05 | 94 | 142 | 166 | 213 |  | 9.96e-05 | 94 | 142 | 166 | 213
9.98e-05 | 94 | 142 | 166 | 213 |  | 9.98e-05 | 94 | 142 | 166 | 213
0.0001 | 83 | 118 | 166 | 213 |  | 0.0001 | 83 | 118 | 166 | 213
0.0001002 | 83 | 118 | 166 | 213 |  | 0.0001002 | 83 | 118 | 166 | 213
0.0001004 | 83 | 118 | 166 | 189 |  | 0.0001004 | 83 | 118 | 166 | 189
0.0001006 | 83 | 118 | 166 | 189 |  | 0.0001006 | 83 | 118 | 166 | 189
0.0001008 | 83 | 118 | 142 | 189 |  | 0.0001008 | 83 | 118 | 142 | 189
0.000101 | 83 | 118 | 142 | 189 |  | 0.000101 | 83 | 118 | 142 | 189
0.0001012 | 83 | 118 | 142 | 189 |  | 0.0001012 | 83 | 118 | 142 | 189
0.0001014 | 71 | 118 | 142 | 166 |  | 0.0001014 | 71 | 95 | 142 | 166
0.0001016 | 71 | 95 | 142 | 166 |  | 0.0001016 | 71 | 95 | 142 | 166
0.0001018 | 71 | 95 | 118 | 166 |  | 0.0001018 | 59 | 95 | 142 | 166
0.000102 | 71 | 95 | 118 | 142 |  | 0.000102 | 71 | 95 | 118 | 166
0.0001022 | 59 | 95 | 118 | 142 |  | 0.0001022 | 59 | 95 | 118 | 142
| 2 | 2 | 2 | 2 |  |  | 2 | 2 | 2 | 2
| 1000 | 1500 | 2000 | 2500 |  |  | 1000 | 1500 | 2000 | 2500
| 2 | 2 | 2 | 2 |  |  | 2 | 2 | 2 | 2

## Step 2 225 Impulse
Time (s) | 1.0 kV | 1.5 kV | 2.0 kV | 2.5 kV | 225 | Time (s) | 1.0 kV | 1.5 kV | 2.0 kV | 2.5 kV |  | Notes
0 | 0 | 0 | 0 | 0 | A17 Coil Pack QXFA 225 Outer to Inner | 0 | 0 | 0 | 0 | 0 | A17 Coil Pack QXFA 225 Outer to Inner
2e-07 | 0 | 0 | 0 | 0 | Retest for 12b | 2e-07 | 0 | 0 | 0 | 0
4e-07 | 0 | 0 | 0 | 0 |  | 4e-07 | 0 | 0 | 0 | 0
6e-07 | 189 | 379 | 521 | 663 |  | 6e-07 | 189 | 379 | 521 | 663
8e-07 | 555 | 946 | 1372 | 1846 |  | 8e-07 | 555 | 946 | 1372 | 1846
1e-06 | 933 | 1467 | 1964 | 2508 |  | 1e-06 | 933 | 1467 | 1964 | 2508
1.2e-06 | 980 | 1443 | 1869 | 2271 |  | 1.2e-06 | 980 | 1443 | 1869 | 2271
1.4e-06 | 874 | 1278 | 1656 | 2011 |  | 1.4e-06 | 874 | 1278 | 1656 | 2011
1.6e-06 | 827 | 1230 | 1609 | 2011 |  | 1.6e-06 | 827 | 1230 | 1609 | 2011
1.8e-06 | 839 | 1254 | 1656 | 2059 |  | 1.8e-06 | 839 | 1254 | 1656 | 2035
2e-06 | 839 | 1254 | 1609 | 1988 |  | 2e-06 | 839 | 1254 | 1609 | 1988
2.2e-06 | 815 | 1183 | 1562 | 1917 |  | 2.2e-06 | 815 | 1207 | 1562 | 1917
2.4e-06 | 791 | 1183 | 1538 | 1893 |  | 2.4e-06 | 791 | 1183 | 1538 | 1893
2.6e-06 | 791 | 1159 | 1538 | 1893 |  | 2.6e-06 | 780 | 1159 | 1538 | 1893
2.8e-06 | 780 | 1159 | 1491 | 1869 |  | 2.8e-06 | 780 | 1159 | 1514 | 1869
3e-06 | 768 | 1136 | 1467 | 1822 |  | 3e-06 | 768 | 1136 | 1467 | 1822
3.2e-06 | 744 | 1088 | 1443 | 1775 |  | 3.2e-06 | 744 | 1112 | 1443 | 1775
3.4e-06 | 720 | 1065 | 1396 | 1727 |  | 3.4e-06 | 732 | 1065 | 1396 | 1727
3.6e-06 | 709 | 1041 | 1372 | 1704 |  | 3.6e-06 | 709 | 1065 | 1372 | 1704
3.8e-06 | 685 | 1017 | 1349 | 1656 |  | 3.8e-06 | 697 | 1017 | 1349 | 1656
4e-06 | 673 | 994 | 1301 | 1633 |  | 4e-06 | 673 | 994 | 1301 | 1633
4.2e-06 | 661 | 970 | 1278 | 1585 |  | 4.2e-06 | 661 | 970 | 1278 | 1585
4.4e-06 | 650 | 970 | 1254 | 1562 |  | 4.4e-06 | 650 | 970 | 1254 | 1538
4.6e-06 | 638 | 946 | 1230 | 1514 |  | 4.6e-06 | 638 | 946 | 1230 | 1514
4.8e-06 | 626 | 923 | 1207 | 1491 |  | 4.8e-06 | 626 | 946 | 1207 | 1491
5e-06 | 614 | 923 | 1183 | 1467 |  | 5e-06 | 614 | 923 | 1183 | 1467
5.2e-06 | 602 | 899 | 1159 | 1443 |  | 5.2e-06 | 602 | 899 | 1159 | 1443
5.4e-06 | 591 | 875 | 1136 | 1420 |  | 5.4e-06 | 591 | 875 | 1136 | 1420
5.6e-06 | 579 | 852 | 1112 | 1372 |  | 5.6e-06 | 579 | 852 | 1112 | 1372
5.8e-06 | 567 | 828 | 1088 | 1349 |  | 5.8e-06 | 567 | 828 | 1088 | 1349
6e-06 | 555 | 828 | 1065 | 1301 |  | 6e-06 | 543 | 804 | 1065 | 1301
6.2e-06 | 531 | 781 | 1017 | 1278 |  | 6.2e-06 | 531 | 781 | 1017 | 1278
6.4e-06 | 520 | 757 | 994 | 1230 |  | 6.4e-06 | 520 | 757 | 994 | 1230
6.6e-06 | 496 | 734 | 970 | 1183 |  | 6.6e-06 | 496 | 734 | 970 | 1183
6.8e-06 | 484 | 710 | 923 | 1136 |  | 6.8e-06 | 484 | 710 | 923 | 1136
7e-06 | 461 | 686 | 875 | 1088 |  | 7e-06 | 461 | 686 | 875 | 1088
7.2e-06 | 449 | 639 | 852 | 1041 |  | 7.2e-06 | 449 | 639 | 852 | 1041
7.4e-06 | 425 | 615 | 804 | 994 |  | 7.4e-06 | 425 | 615 | 804 | 994
7.6e-06 | 402 | 592 | 757 | 946 |  | 7.6e-06 | 402 | 592 | 757 | 946
7.8e-06 | 378 | 544 | 734 | 899 |  | 7.8e-06 | 378 | 568 | 734 | 899
8e-06 | 354 | 521 | 686 | 828 |  | 8e-06 | 354 | 521 | 686 | 828
8.2e-06 | 343 | 497 | 639 | 781 |  | 8.2e-06 | 331 | 497 | 639 | 781
8.4e-06 | 307 | 450 | 592 | 710 |  | 8.4e-06 | 307 | 450 | 592 | 734
8.6e-06 | 283 | 426 | 544 | 663 |  | 8.6e-06 | 283 | 426 | 544 | 663
8.8e-06 | 260 | 379 | 497 | 615 |  | 8.8e-06 | 260 | 379 | 497 | 615
9e-06 | 236 | 355 | 450 | 544 |  | 9e-06 | 248 | 331 | 450 | 544
9.2e-06 | 213 | 308 | 402 | 497 |  | 9.2e-06 | 213 | 308 | 402 | 497
9.4e-06 | 189 | 260 | 355 | 426 |  | 9.4e-06 | 189 | 260 | 355 | 426
9.6e-06 | 165 | 237 | 308 | 379 |  | 9.6e-06 | 165 | 237 | 308 | 379
9.8e-06 | 154 | 213 | 260 | 308 |  | 9.8e-06 | 142 | 189 | 260 | 308
1e-05 | 118 | 189 | 213 | 260 |  | 1e-05 | 118 | 166 | 213 | 260
1.02e-05 | 94 | 142 | 166 | 213 |  | 1.02e-05 | 106 | 142 | 166 | 213
1.04e-05 | 71 | 95 | 118 | 166 |  | 1.04e-05 | 83 | 95 | 118 | 142
1.06e-05 | 47 | 71 | 71 | 95 |  | 1.06e-05 | 47 | 71 | 71 | 95
1.08e-05 | 24 | 24 | 24 | 47 |  | 1.08e-05 | 24 | 24 | 47 | 47
1.1e-05 | 0 | 0 | 0 | -23 |  | 1.1e-05 | 0 | 0 | 0 | -23
1.12e-05 | -23 | -46 | -46 | -70 |  | 1.12e-05 | -23 | -23 | -46 | -70
1.14e-05 | -46 | -70 | -94 | -117 |  | 1.14e-05 | -34 | -70 | -94 | -117
1.16e-05 | -58 | -94 | -141 | -188 |  | 1.16e-05 | -58 | -94 | -141 | -165
1.18e-05 | -82 | -141 | -188 | -236 |  | 1.18e-05 | -82 | -141 | -188 | -212
1.2e-05 | -105 | -165 | -212 | -283 |  | 1.2e-05 | -105 | -165 | -212 | -283
1.22e-05 | -129 | -212 | -259 | -330 |  | 1.22e-05 | -129 | -188 | -259 | -330
1.24e-05 | -141 | -236 | -307 | -378 |  | 1.24e-05 | -153 | -236 | -307 | -378
1.26e-05 | -164 | -259 | -354 | -425 |  | 1.26e-05 | -164 | -259 | -354 | -425
1.28e-05 | -188 | -307 | -378 | -472 |  | 1.28e-05 | -188 | -283 | -378 | -472
1.3e-05 | -212 | -330 | -425 | -520 |  | 1.3e-05 | -212 | -330 | -425 | -520
1.32e-05 | -223 | -354 | -449 | -567 |  | 1.32e-05 | -223 | -354 | -449 | -567
1.34e-05 | -247 | -378 | -496 | -614 |  | 1.34e-05 | -247 | -378 | -496 | -614
1.36e-05 | -271 | -401 | -543 | -685 |  | 1.36e-05 | -271 | -401 | -543 | -662
1.38e-05 | -282 | -449 | -567 | -709 |  | 1.38e-05 | -282 | -449 | -567 | -709
1.4e-05 | -306 | -472 | -614 | -756 |  | 1.4e-05 | -306 | -472 | -614 | -756
1.42e-05 | -318 | -496 | -638 | -803 |  | 1.42e-05 | -318 | -496 | -638 | -803
1.44e-05 | -342 | -520 | -685 | -851 |  | 1.44e-05 | -342 | -520 | -685 | -851
1.46e-05 | -353 | -543 | -709 | -898 |  | 1.46e-05 | -365 | -543 | -709 | -898
1.48e-05 | -377 | -567 | -756 | -945 |  | 1.48e-05 | -377 | -567 | -756 | -922
1.5e-05 | -389 | -591 | -780 | -969 |  | 1.5e-05 | -401 | -591 | -780 | -969
1.52e-05 | -412 | -614 | -827 | -1016 |  | 1.52e-05 | -412 | -614 | -827 | -1016
1.54e-05 | -436 | -638 | -851 | -1064 |  | 1.54e-05 | -424 | -638 | -851 | -1064
1.56e-05 | -448 | -685 | -874 | -1111 |  | 1.56e-05 | -448 | -662 | -874 | -1087
1.58e-05 | -460 | -709 | -922 | -1135 |  | 1.58e-05 | -460 | -685 | -922 | -1135
1.6e-05 | -471 | -709 | -945 | -1158 |  | 1.6e-05 | -483 | -733 | -945 | -1158
1.62e-05 | -495 | -733 | -969 | -1206 |  | 1.62e-05 | -495 | -733 | -969 | -1206
1.64e-05 | -507 | -756 | -1016 | -1253 |  | 1.64e-05 | -507 | -756 | -993 | -1229
1.66e-05 | -519 | -780 | -1016 | -1277 |  | 1.66e-05 | -519 | -780 | -1016 | -1277
1.68e-05 | -530 | -803 | -1064 | -1300 |  | 1.68e-05 | -530 | -803 | -1064 | -1300
1.7e-05 | -542 | -827 | -1087 | -1324 |  | 1.7e-05 | -542 | -827 | -1087 | -1324
1.72e-05 | -554 | -827 | -1111 | -1348 |  | 1.72e-05 | -554 | -827 | -1111 | -1348
1.74e-05 | -566 | -851 | -1135 | -1395 |  | 1.74e-05 | -566 | -851 | -1135 | -1395
1.76e-05 | -578 | -874 | -1135 | -1419 |  | 1.76e-05 | -578 | -874 | -1158 | -1419
1.78e-05 | -590 | -898 | -1158 | -1442 |  | 1.78e-05 | -590 | -874 | -1158 | -1442
1.8e-05 | -601 | -898 | -1182 | -1466 |  | 1.8e-05 | -601 | -898 | -1182 | -1466
1.82e-05 | -601 | -922 | -1206 | -1490 |  | 1.82e-05 | -613 | -922 | -1206 | -1490
1.84e-05 | -613 | -922 | -1206 | -1513 |  | 1.84e-05 | -625 | -945 | -1229 | -1513
1.86e-05 | -625 | -945 | -1229 | -1513 |  | 1.86e-05 | -625 | -945 | -1229 | -1513
1.88e-05 | -637 | -945 | -1253 | -1537 |  | 1.88e-05 | -637 | -945 | -1253 | -1537
1.9e-05 | -637 | -969 | -1253 | -1561 |  | 1.9e-05 | -649 | -969 | -1253 | -1561
1.92e-05 | -649 | -969 | -1277 | -1584 |  | 1.92e-05 | -649 | -969 | -1277 | -1561
1.94e-05 | -649 | -969 | -1277 | -1584 |  | 1.94e-05 | -660 | -993 | -1277 | -1584
1.96e-05 | -660 | -993 | -1300 | -1584 |  | 1.96e-05 | -660 | -993 | -1300 | -1584
1.98e-05 | -660 | -993 | -1300 | -1608 |  | 1.98e-05 | -660 | -993 | -1300 | -1608
2e-05 | -660 | -993 | -1300 | -1608 |  | 2e-05 | -660 | -993 | -1324 | -1632
2.02e-05 | -672 | -1016 | -1324 | -1632 |  | 2.02e-05 | -672 | -993 | -1324 | -1632
2.04e-05 | -672 | -1016 | -1324 | -1632 |  | 2.04e-05 | -672 | -1016 | -1324 | -1632
2.06e-05 | -672 | -1016 | -1324 | -1632 |  | 2.06e-05 | -672 | -1016 | -1324 | -1632
2.08e-05 | -672 | -1016 | -1324 | -1632 |  | 2.08e-05 | -672 | -1016 | -1324 | -1632
2.1e-05 | -672 | -1016 | -1324 | -1632 |  | 2.1e-05 | -672 | -1016 | -1324 | -1632
2.12e-05 | -672 | -1016 | -1324 | -1632 |  | 2.12e-05 | -672 | -1016 | -1324 | -1632
2.14e-05 | -672 | -1016 | -1324 | -1632 |  | 2.14e-05 | -672 | -1016 | -1324 | -1632
2.16e-05 | -672 | -993 | -1324 | -1632 |  | 2.16e-05 | -672 | -1016 | -1324 | -1632
2.18e-05 | -672 | -993 | -1300 | -1608 |  | 2.18e-05 | -672 | -993 | -1324 | -1608
2.2e-05 | -660 | -993 | -1300 | -1608 |  | 2.2e-05 | -660 | -993 | -1300 | -1608
2.22e-05 | -660 | -969 | -1277 | -1584 |  | 2.22e-05 | -660 | -993 | -1300 | -1584
2.24e-05 | -649 | -969 | -1277 | -1561 |  | 2.24e-05 | -649 | -969 | -1277 | -1561
2.26e-05 | -637 | -945 | -1253 | -1537 |  | 2.26e-05 | -637 | -969 | -1253 | -1537
2.28e-05 | -637 | -945 | -1229 | -1513 |  | 2.28e-05 | -637 | -945 | -1229 | -1513
2.3e-05 | -625 | -922 | -1206 | -1490 |  | 2.3e-05 | -625 | -922 | -1206 | -1490
2.32e-05 | -601 | -898 | -1182 | -1466 |  | 2.32e-05 | -613 | -922 | -1182 | -1466
2.34e-05 | -590 | -898 | -1158 | -1419 |  | 2.34e-05 | -590 | -898 | -1158 | -1442
2.36e-05 | -578 | -874 | -1135 | -1395 |  | 2.36e-05 | -578 | -874 | -1135 | -1395
2.38e-05 | -566 | -851 | -1111 | -1348 |  | 2.38e-05 | -566 | -851 | -1111 | -1371
2.4e-05 | -554 | -827 | -1064 | -1324 |  | 2.4e-05 | -554 | -827 | -1087 | -1324
2.42e-05 | -542 | -803 | -1040 | -1277 |  | 2.42e-05 | -542 | -803 | -1040 | -1300
2.44e-05 | -519 | -780 | -1016 | -1253 |  | 2.44e-05 | -519 | -780 | -1016 | -1253
2.46e-05 | -507 | -756 | -969 | -1206 |  | 2.46e-05 | -507 | -756 | -993 | -1206
2.48e-05 | -483 | -733 | -945 | -1158 |  | 2.48e-05 | -483 | -733 | -945 | -1158
2.5e-05 | -471 | -685 | -898 | -1111 |  | 2.5e-05 | -471 | -709 | -898 | -1111
2.52e-05 | -448 | -662 | -874 | -1064 |  | 2.52e-05 | -448 | -662 | -874 | -1064
2.54e-05 | -424 | -638 | -827 | -1016 |  | 2.54e-05 | -436 | -638 | -827 | -1016
2.56e-05 | -401 | -614 | -780 | -969 |  | 2.56e-05 | -412 | -614 | -803 | -969
2.58e-05 | -377 | -567 | -756 | -922 |  | 2.58e-05 | -389 | -591 | -756 | -922
2.6e-05 | -365 | -543 | -709 | -874 |  | 2.6e-05 | -365 | -543 | -709 | -874
2.62e-05 | -342 | -520 | -662 | -803 |  | 2.62e-05 | -342 | -520 | -662 | -803
2.64e-05 | -318 | -472 | -614 | -756 |  | 2.64e-05 | -318 | -472 | -614 | -756
2.66e-05 | -294 | -425 | -567 | -685 |  | 2.66e-05 | -294 | -449 | -567 | -709
2.68e-05 | -271 | -401 | -520 | -638 |  | 2.68e-05 | -271 | -401 | -543 | -638
2.7e-05 | -259 | -354 | -472 | -567 |  | 2.7e-05 | -259 | -378 | -496 | -591
2.72e-05 | -223 | -330 | -425 | -520 |  | 2.72e-05 | -223 | -330 | -449 | -543
2.74e-05 | -200 | -307 | -378 | -472 |  | 2.74e-05 | -200 | -307 | -401 | -472
2.76e-05 | -176 | -259 | -330 | -401 |  | 2.76e-05 | -188 | -283 | -354 | -425
2.78e-05 | -153 | -236 | -307 | -354 |  | 2.78e-05 | -164 | -236 | -307 | -354
2.8e-05 | -129 | -188 | -236 | -307 |  | 2.8e-05 | -141 | -188 | -259 | -307
2.82e-05 | -105 | -165 | -188 | -236 |  | 2.82e-05 | -117 | -165 | -212 | -259
2.84e-05 | -82 | -117 | -165 | -188 |  | 2.84e-05 | -93 | -117 | -165 | -188
2.86e-05 | -58 | -94 | -117 | -117 |  | 2.86e-05 | -70 | -94 | -117 | -141
2.88e-05 | -34 | -46 | -70 | -70 |  | 2.88e-05 | -34 | -46 | -70 | -70
2.9e-05 | -11 | -23 | 0 | 0 |  | 2.9e-05 | -23 | -23 | -23 | -23
2.92e-05 | 12 | 24 | 47 | 47 |  | 2.92e-05 | 12 | 24 | 24 | 47
2.94e-05 | 35 | 47 | 71 | 118 |  | 2.94e-05 | 35 | 47 | 71 | 95
2.96e-05 | 59 | 95 | 142 | 166 |  | 2.96e-05 | 59 | 71 | 118 | 166
2.98e-05 | 83 | 118 | 166 | 213 |  | 2.98e-05 | 83 | 118 | 166 | 213
3e-05 | 106 | 166 | 213 | 260 |  | 3e-05 | 94 | 142 | 213 | 260
3.02e-05 | 118 | 189 | 260 | 308 |  | 3.02e-05 | 118 | 189 | 237 | 308
3.04e-05 | 142 | 213 | 284 | 355 |  | 3.04e-05 | 142 | 213 | 284 | 355
3.06e-05 | 165 | 260 | 331 | 426 |  | 3.06e-05 | 154 | 237 | 331 | 402
3.08e-05 | 177 | 284 | 379 | 473 |  | 3.08e-05 | 177 | 260 | 355 | 450
3.1e-05 | 201 | 308 | 402 | 521 |  | 3.1e-05 | 201 | 308 | 402 | 497
3.12e-05 | 224 | 331 | 450 | 544 |  | 3.12e-05 | 224 | 331 | 426 | 544
3.14e-05 | 236 | 355 | 473 | 592 |  | 3.14e-05 | 236 | 355 | 473 | 592
3.16e-05 | 260 | 379 | 521 | 639 |  | 3.16e-05 | 260 | 379 | 521 | 639
3.18e-05 | 272 | 402 | 544 | 686 |  | 3.18e-05 | 272 | 402 | 544 | 686
3.2e-05 | 295 | 426 | 568 | 734 |  | 3.2e-05 | 283 | 426 | 568 | 710
3.22e-05 | 307 | 450 | 615 | 757 |  | 3.22e-05 | 307 | 450 | 592 | 734
3.24e-05 | 319 | 473 | 639 | 781 |  | 3.24e-05 | 319 | 473 | 639 | 781
3.26e-05 | 331 | 497 | 663 | 828 |  | 3.26e-05 | 331 | 497 | 663 | 828
3.28e-05 | 343 | 521 | 686 | 852 |  | 3.28e-05 | 343 | 521 | 686 | 852
3.3e-05 | 354 | 544 | 710 | 899 |  | 3.3e-05 | 354 | 544 | 710 | 875
3.32e-05 | 366 | 544 | 734 | 923 |  | 3.32e-05 | 366 | 544 | 734 | 899
3.34e-05 | 378 | 568 | 757 | 946 |  | 3.34e-05 | 378 | 568 | 757 | 923
3.36e-05 | 390 | 592 | 781 | 970 |  | 3.36e-05 | 390 | 592 | 757 | 946
3.38e-05 | 402 | 592 | 781 | 970 |  | 3.38e-05 | 402 | 592 | 781 | 970
3.4e-05 | 413 | 615 | 804 | 994 |  | 3.4e-05 | 402 | 615 | 804 | 994
3.42e-05 | 413 | 615 | 828 | 1017 |  | 3.42e-05 | 413 | 615 | 828 | 1017
3.44e-05 | 425 | 639 | 828 | 1017 |  | 3.44e-05 | 425 | 639 | 828 | 1017
3.46e-05 | 425 | 639 | 828 | 1041 |  | 3.46e-05 | 425 | 639 | 828 | 1041
3.48e-05 | 437 | 639 | 852 | 1041 |  | 3.48e-05 | 437 | 639 | 852 | 1041
3.5e-05 | 437 | 639 | 875 | 1065 |  | 3.5e-05 | 437 | 639 | 852 | 1065
3.52e-05 | 437 | 663 | 875 | 1065 |  | 3.52e-05 | 437 | 639 | 852 | 1065
3.54e-05 | 437 | 663 | 875 | 1065 |  | 3.54e-05 | 449 | 639 | 875 | 1065
3.56e-05 | 437 | 663 | 875 | 1065 |  | 3.56e-05 | 437 | 663 | 875 | 1065
3.58e-05 | 449 | 663 | 875 | 1065 |  | 3.58e-05 | 449 | 639 | 875 | 1065
3.6e-05 | 437 | 663 | 852 | 1065 |  | 3.6e-05 | 449 | 663 | 875 | 1065
3.62e-05 | 437 | 639 | 852 | 1065 |  | 3.62e-05 | 449 | 663 | 852 | 1065
3.64e-05 | 437 | 639 | 852 | 1065 |  | 3.64e-05 | 437 | 663 | 852 | 1065
3.66e-05 | 437 | 639 | 852 | 1041 |  | 3.66e-05 | 437 | 639 | 852 | 1065
3.68e-05 | 437 | 639 | 852 | 1041 |  | 3.68e-05 | 437 | 639 | 852 | 1041
3.7e-05 | 425 | 639 | 828 | 1017 |  | 3.7e-05 | 437 | 639 | 828 | 1041
3.72e-05 | 425 | 615 | 828 | 1017 |  | 3.72e-05 | 425 | 639 | 828 | 1017
3.74e-05 | 413 | 615 | 804 | 994 |  | 3.74e-05 | 413 | 615 | 804 | 1017
3.76e-05 | 413 | 615 | 804 | 994 |  | 3.76e-05 | 413 | 615 | 804 | 994
3.78e-05 | 402 | 592 | 781 | 970 |  | 3.78e-05 | 402 | 592 | 781 | 970
3.8e-05 | 390 | 592 | 757 | 946 |  | 3.8e-05 | 390 | 592 | 781 | 946
3.82e-05 | 390 | 568 | 734 | 923 |  | 3.82e-05 | 390 | 568 | 757 | 923
3.84e-05 | 378 | 568 | 734 | 899 |  | 3.84e-05 | 378 | 568 | 734 | 899
3.86e-05 | 366 | 544 | 710 | 875 |  | 3.86e-05 | 366 | 544 | 710 | 899
3.88e-05 | 354 | 521 | 686 | 852 |  | 3.88e-05 | 354 | 521 | 686 | 852
3.9e-05 | 343 | 521 | 663 | 828 |  | 3.9e-05 | 343 | 521 | 686 | 828
3.92e-05 | 331 | 497 | 639 | 781 |  | 3.92e-05 | 343 | 497 | 639 | 804
3.94e-05 | 319 | 473 | 615 | 757 |  | 3.94e-05 | 331 | 473 | 615 | 781
3.96e-05 | 307 | 450 | 592 | 734 |  | 3.96e-05 | 307 | 450 | 592 | 734
3.98e-05 | 295 | 426 | 568 | 686 |  | 3.98e-05 | 295 | 426 | 568 | 710
4e-05 | 283 | 402 | 544 | 663 |  | 4e-05 | 283 | 402 | 544 | 686
4.02e-05 | 260 | 379 | 521 | 639 |  | 4.02e-05 | 272 | 402 | 521 | 639
4.04e-05 | 248 | 379 | 473 | 592 |  | 4.04e-05 | 260 | 379 | 497 | 615
4.06e-05 | 236 | 355 | 450 | 544 |  | 4.06e-05 | 248 | 355 | 450 | 568
4.08e-05 | 224 | 331 | 426 | 521 |  | 4.08e-05 | 224 | 331 | 450 | 544
4.1e-05 | 201 | 308 | 379 | 497 |  | 4.1e-05 | 213 | 308 | 402 | 497
4.12e-05 | 189 | 284 | 355 | 450 |  | 4.12e-05 | 201 | 284 | 379 | 450
4.14e-05 | 177 | 260 | 331 | 402 |  | 4.14e-05 | 177 | 260 | 331 | 426
4.16e-05 | 154 | 237 | 308 | 355 |  | 4.16e-05 | 165 | 237 | 308 | 379
4.18e-05 | 142 | 213 | 260 | 331 |  | 4.18e-05 | 154 | 213 | 284 | 331
4.2e-05 | 130 | 189 | 237 | 284 |  | 4.2e-05 | 130 | 189 | 260 | 308
4.22e-05 | 118 | 166 | 213 | 260 |  | 4.22e-05 | 118 | 166 | 213 | 284
4.24e-05 | 94 | 142 | 189 | 213 |  | 4.24e-05 | 106 | 142 | 189 | 237
4.26e-05 | 83 | 118 | 142 | 166 |  | 4.26e-05 | 83 | 118 | 166 | 189
4.28e-05 | 59 | 95 | 118 | 142 |  | 4.28e-05 | 71 | 95 | 118 | 166
4.3e-05 | 47 | 71 | 71 | 95 |  | 4.3e-05 | 59 | 71 | 95 | 118
4.32e-05 | 35 | 47 | 47 | 47 |  | 4.32e-05 | 35 | 47 | 71 | 71
4.34e-05 | 12 | 24 | 24 | 24 |  | 4.34e-05 | 24 | 24 | 24 | 47
4.36e-05 | 0 | 0 | 0 | -23 |  | 4.36e-05 | 12 | 0 | 0 | 0
4.38e-05 | -11 | -23 | -46 | -46 |  | 4.38e-05 | -11 | -23 | -23 | -23
4.4e-05 | -23 | -46 | -70 | -94 |  | 4.4e-05 | -23 | -46 | -70 | -70
4.42e-05 | -46 | -70 | -94 | -141 |  | 4.42e-05 | -34 | -70 | -94 | -117
4.44e-05 | -58 | -94 | -141 | -165 |  | 4.44e-05 | -58 | -94 | -117 | -141
4.46e-05 | -70 | -117 | -165 | -212 |  | 4.46e-05 | -70 | -117 | -141 | -188
4.48e-05 | -93 | -141 | -188 | -236 |  | 4.48e-05 | -82 | -141 | -188 | -212
4.5e-05 | -105 | -165 | -212 | -259 |  | 4.5e-05 | -105 | -165 | -212 | -259
4.52e-05 | -117 | -188 | -236 | -307 |  | 4.52e-05 | -117 | -165 | -236 | -283
4.54e-05 | -129 | -188 | -283 | -330 |  | 4.54e-05 | -129 | -188 | -259 | -330
4.56e-05 | -141 | -236 | -307 | -354 |  | 4.56e-05 | -141 | -212 | -283 | -354
4.58e-05 | -153 | -236 | -330 | -401 |  | 4.58e-05 | -153 | -236 | -307 | -378
4.6e-05 | -164 | -259 | -330 | -425 |  | 4.6e-05 | -164 | -259 | -330 | -401
4.62e-05 | -176 | -283 | -354 | -449 |  | 4.62e-05 | -176 | -283 | -354 | -425
4.64e-05 | -188 | -283 | -378 | -472 |  | 4.64e-05 | -188 | -283 | -378 | -472
4.66e-05 | -200 | -307 | -401 | -520 |  | 4.66e-05 | -188 | -307 | -401 | -496
4.68e-05 | -212 | -330 | -425 | -543 |  | 4.68e-05 | -212 | -307 | -425 | -520
4.7e-05 | -223 | -330 | -449 | -567 |  | 4.7e-05 | -223 | -330 | -449 | -543
4.72e-05 | -235 | -354 | -449 | -567 |  | 4.72e-05 | -223 | -354 | -449 | -567
4.74e-05 | -247 | -354 | -472 | -591 |  | 4.74e-05 | -235 | -354 | -472 | -591
4.76e-05 | -247 | -378 | -496 | -614 |  | 4.76e-05 | -247 | -378 | -496 | -614
4.78e-05 | -259 | -378 | -520 | -638 |  | 4.78e-05 | -259 | -401 | -496 | -614
4.8e-05 | -259 | -401 | -520 | -638 |  | 4.8e-05 | -259 | -378 | -520 | -638
4.82e-05 | -271 | -401 | -543 | -662 |  | 4.82e-05 | -271 | -401 | -543 | -662
4.84e-05 | -282 | -401 | -543 | -685 |  | 4.84e-05 | -271 | -401 | -543 | -662
4.86e-05 | -282 | -425 | -543 | -685 |  | 4.86e-05 | -282 | -425 | -543 | -685
4.88e-05 | -282 | -425 | -567 | -709 |  | 4.88e-05 | -282 | -425 | -567 | -709
4.9e-05 | -294 | -425 | -567 | -709 |  | 4.9e-05 | -294 | -425 | -567 | -709
4.92e-05 | -294 | -449 | -591 | -733 |  | 4.92e-05 | -294 | -449 | -567 | -709
4.94e-05 | -294 | -449 | -591 | -733 |  | 4.94e-05 | -294 | -449 | -591 | -733
4.96e-05 | -294 | -449 | -591 | -733 |  | 4.96e-05 | -294 | -449 | -591 | -733
4.98e-05 | -306 | -449 | -591 | -733 |  | 4.98e-05 | -294 | -449 | -614 | -733
5e-05 | -294 | -449 | -591 | -733 |  | 5e-05 | -306 | -449 | -591 | -733
5.02e-05 | -306 | -449 | -591 | -733 |  | 5.02e-05 | -306 | -449 | -591 | -733
5.04e-05 | -306 | -449 | -591 | -733 |  | 5.04e-05 | -306 | -449 | -591 | -756
5.06e-05 | -306 | -449 | -591 | -733 |  | 5.06e-05 | -306 | -449 | -591 | -733
5.08e-05 | -306 | -449 | -591 | -733 |  | 5.08e-05 | -306 | -449 | -591 | -733
5.1e-05 | -294 | -449 | -591 | -733 |  | 5.1e-05 | -306 | -449 | -591 | -733
5.12e-05 | -294 | -449 | -591 | -733 |  | 5.12e-05 | -294 | -449 | -591 | -733
5.14e-05 | -294 | -449 | -567 | -709 |  | 5.14e-05 | -294 | -449 | -591 | -733
5.16e-05 | -294 | -425 | -567 | -709 |  | 5.16e-05 | -294 | -449 | -567 | -709
5.18e-05 | -282 | -425 | -567 | -709 |  | 5.18e-05 | -294 | -449 | -567 | -709
5.2e-05 | -282 | -425 | -567 | -685 |  | 5.2e-05 | -294 | -425 | -567 | -685
5.22e-05 | -282 | -425 | -543 | -685 |  | 5.22e-05 | -282 | -425 | -567 | -685
5.24e-05 | -282 | -401 | -543 | -662 |  | 5.24e-05 | -271 | -425 | -543 | -685
5.26e-05 | -271 | -401 | -520 | -662 |  | 5.26e-05 | -271 | -401 | -543 | -662
5.28e-05 | -271 | -401 | -520 | -638 |  | 5.28e-05 | -271 | -401 | -520 | -638
5.3e-05 | -259 | -378 | -496 | -614 |  | 5.3e-05 | -259 | -401 | -520 | -638
5.32e-05 | -247 | -378 | -496 | -614 |  | 5.32e-05 | -259 | -378 | -496 | -614
5.34e-05 | -247 | -378 | -472 | -591 |  | 5.34e-05 | -247 | -378 | -496 | -591
5.36e-05 | -235 | -354 | -472 | -567 |  | 5.36e-05 | -247 | -354 | -472 | -591
5.38e-05 | -223 | -330 | -449 | -543 |  | 5.38e-05 | -235 | -354 | -449 | -567
5.4e-05 | -223 | -330 | -425 | -543 |  | 5.4e-05 | -223 | -330 | -449 | -543
5.42e-05 | -212 | -307 | -401 | -520 |  | 5.42e-05 | -212 | -330 | -425 | -520
5.44e-05 | -200 | -307 | -401 | -496 |  | 5.44e-05 | -212 | -307 | -401 | -496
5.46e-05 | -188 | -283 | -378 | -472 |  | 5.46e-05 | -200 | -307 | -401 | -472
5.48e-05 | -188 | -283 | -354 | -449 |  | 5.48e-05 | -188 | -283 | -378 | -449
5.5e-05 | -176 | -259 | -354 | -401 |  | 5.5e-05 | -188 | -283 | -354 | -425
5.52e-05 | -164 | -259 | -330 | -401 |  | 5.52e-05 | -176 | -259 | -330 | -401
5.54e-05 | -153 | -236 | -307 | -378 |  | 5.54e-05 | -164 | -259 | -307 | -378
5.56e-05 | -153 | -212 | -283 | -354 |  | 5.56e-05 | -153 | -236 | -307 | -354
5.58e-05 | -141 | -212 | -259 | -330 |  | 5.58e-05 | -141 | -212 | -283 | -330
5.6e-05 | -129 | -188 | -236 | -307 |  | 5.6e-05 | -129 | -188 | -259 | -307
5.62e-05 | -117 | -165 | -212 | -283 |  | 5.62e-05 | -117 | -188 | -236 | -283
5.64e-05 | -105 | -165 | -212 | -236 |  | 5.64e-05 | -117 | -165 | -212 | -259
5.66e-05 | -93 | -141 | -188 | -212 |  | 5.66e-05 | -105 | -141 | -188 | -236
5.68e-05 | -82 | -117 | -165 | -188 |  | 5.68e-05 | -93 | -141 | -165 | -212
5.7e-05 | -70 | -117 | -141 | -165 |  | 5.7e-05 | -82 | -117 | -141 | -188
5.72e-05 | -58 | -94 | -117 | -141 |  | 5.72e-05 | -70 | -94 | -141 | -165
5.74e-05 | -46 | -70 | -94 | -117 |  | 5.74e-05 | -58 | -94 | -117 | -117
5.76e-05 | -34 | -46 | -70 | -94 |  | 5.76e-05 | -46 | -70 | -70 | -94
5.78e-05 | -34 | -46 | -46 | -70 |  | 5.78e-05 | -34 | -46 | -70 | -70
5.8e-05 | -11 | -23 | -23 | -46 |  | 5.8e-05 | -23 | -23 | -46 | -46
5.82e-05 | -11 | 0 | 0 | 0 |  | 5.82e-05 | -11 | -23 | -23 | -23
5.84e-05 | 0 | 0 | 24 | 24 |  | 5.84e-05 | 0 | 0 | 0 | 0
5.86e-05 | 12 | 24 | 24 | 47 |  | 5.86e-05 | 12 | 24 | 24 | 24
5.88e-05 | 24 | 47 | 47 | 71 |  | 5.88e-05 | 24 | 24 | 24 | 47
5.9e-05 | 35 | 47 | 71 | 95 |  | 5.9e-05 | 24 | 47 | 47 | 71
5.92e-05 | 47 | 71 | 95 | 118 |  | 5.92e-05 | 35 | 71 | 95 | 95
5.94e-05 | 59 | 71 | 118 | 142 |  | 5.94e-05 | 47 | 71 | 95 | 118
5.96e-05 | 71 | 95 | 142 | 166 |  | 5.96e-05 | 59 | 95 | 118 | 142
5.98e-05 | 71 | 118 | 142 | 189 |  | 5.98e-05 | 71 | 95 | 142 | 166
6e-05 | 83 | 118 | 166 | 213 |  | 6e-05 | 83 | 118 | 166 | 189
6.02e-05 | 94 | 142 | 189 | 237 |  | 6.02e-05 | 83 | 142 | 166 | 213
6.04e-05 | 94 | 142 | 189 | 237 |  | 6.04e-05 | 94 | 142 | 189 | 237
6.06e-05 | 106 | 166 | 213 | 260 |  | 6.06e-05 | 106 | 142 | 213 | 260
6.08e-05 | 118 | 166 | 237 | 284 |  | 6.08e-05 | 106 | 166 | 213 | 260
6.1e-05 | 118 | 189 | 237 | 308 |  | 6.1e-05 | 118 | 189 | 237 | 284
6.12e-05 | 130 | 189 | 260 | 308 |  | 6.12e-05 | 130 | 189 | 260 | 308
6.14e-05 | 130 | 213 | 260 | 331 |  | 6.14e-05 | 130 | 189 | 260 | 331
6.16e-05 | 142 | 213 | 284 | 355 |  | 6.16e-05 | 142 | 213 | 260 | 331
6.18e-05 | 142 | 237 | 284 | 355 |  | 6.18e-05 | 142 | 213 | 284 | 355
6.2e-05 | 154 | 237 | 308 | 379 |  | 6.2e-05 | 154 | 237 | 308 | 355
6.22e-05 | 165 | 237 | 308 | 402 |  | 6.22e-05 | 154 | 237 | 308 | 379
6.24e-05 | 165 | 237 | 331 | 402 |  | 6.24e-05 | 165 | 237 | 308 | 402
6.26e-05 | 165 | 237 | 331 | 426 |  | 6.26e-05 | 165 | 260 | 331 | 402
6.28e-05 | 177 | 260 | 331 | 426 |  | 6.28e-05 | 165 | 260 | 331 | 426
6.3e-05 | 177 | 260 | 355 | 426 |  | 6.3e-05 | 177 | 260 | 331 | 426
6.32e-05 | 177 | 260 | 355 | 450 |  | 6.32e-05 | 177 | 260 | 355 | 426
6.34e-05 | 177 | 260 | 355 | 450 |  | 6.34e-05 | 189 | 260 | 355 | 450
6.36e-05 | 189 | 284 | 379 | 450 |  | 6.36e-05 | 189 | 284 | 355 | 450
6.38e-05 | 189 | 284 | 379 | 473 |  | 6.38e-05 | 189 | 284 | 379 | 450
6.4e-05 | 189 | 284 | 379 | 473 |  | 6.4e-05 | 189 | 284 | 379 | 473
6.42e-05 | 189 | 284 | 379 | 473 |  | 6.42e-05 | 201 | 284 | 379 | 473
6.44e-05 | 201 | 284 | 379 | 473 |  | 6.44e-05 | 201 | 284 | 379 | 473
6.46e-05 | 201 | 284 | 379 | 473 |  | 6.46e-05 | 201 | 284 | 379 | 473
6.48e-05 | 201 | 284 | 379 | 473 |  | 6.48e-05 | 201 | 284 | 379 | 473
6.5e-05 | 201 | 284 | 379 | 473 |  | 6.5e-05 | 201 | 284 | 379 | 473
6.52e-05 | 201 | 284 | 379 | 473 |  | 6.52e-05 | 201 | 284 | 379 | 473
6.54e-05 | 201 | 284 | 379 | 473 |  | 6.54e-05 | 201 | 284 | 379 | 473
6.56e-05 | 201 | 284 | 379 | 473 |  | 6.56e-05 | 201 | 284 | 379 | 473
6.58e-05 | 189 | 284 | 379 | 473 |  | 6.58e-05 | 201 | 284 | 379 | 473
6.6e-05 | 189 | 284 | 379 | 450 |  | 6.6e-05 | 189 | 284 | 379 | 473
6.62e-05 | 189 | 284 | 379 | 450 |  | 6.62e-05 | 189 | 284 | 379 | 450
6.64e-05 | 189 | 284 | 355 | 450 |  | 6.64e-05 | 189 | 284 | 379 | 450
6.66e-05 | 189 | 284 | 355 | 450 |  | 6.66e-05 | 189 | 284 | 379 | 450
6.68e-05 | 189 | 260 | 355 | 426 |  | 6.68e-05 | 189 | 284 | 355 | 450
6.7e-05 | 177 | 260 | 355 | 426 |  | 6.7e-05 | 177 | 260 | 355 | 450
6.72e-05 | 177 | 260 | 331 | 426 |  | 6.72e-05 | 177 | 260 | 355 | 426
6.74e-05 | 177 | 260 | 331 | 426 |  | 6.74e-05 | 177 | 260 | 331 | 426
6.76e-05 | 165 | 237 | 331 | 402 |  | 6.76e-05 | 177 | 260 | 331 | 402
6.78e-05 | 165 | 237 | 308 | 379 |  | 6.78e-05 | 165 | 237 | 331 | 402
6.8e-05 | 165 | 237 | 308 | 379 |  | 6.8e-05 | 165 | 237 | 308 | 379
6.82e-05 | 154 | 237 | 308 | 355 |  | 6.82e-05 | 154 | 237 | 308 | 379
6.84e-05 | 154 | 213 | 284 | 355 |  | 6.84e-05 | 154 | 237 | 284 | 355
6.86e-05 | 142 | 213 | 284 | 331 |  | 6.86e-05 | 154 | 213 | 284 | 355
6.88e-05 | 130 | 189 | 260 | 331 |  | 6.88e-05 | 142 | 213 | 260 | 331
6.9e-05 | 130 | 189 | 260 | 308 |  | 6.9e-05 | 130 | 213 | 260 | 331
6.92e-05 | 118 | 189 | 237 | 308 |  | 6.92e-05 | 130 | 189 | 260 | 308
6.94e-05 | 118 | 166 | 237 | 284 |  | 6.94e-05 | 130 | 189 | 237 | 284
6.96e-05 | 118 | 166 | 213 | 260 |  | 6.96e-05 | 118 | 166 | 237 | 284
6.98e-05 | 106 | 166 | 213 | 260 |  | 6.98e-05 | 118 | 166 | 213 | 260
7e-05 | 106 | 142 | 189 | 237 |  | 7e-05 | 106 | 166 | 213 | 260
7.02e-05 | 94 | 142 | 189 | 213 |  | 7.02e-05 | 106 | 142 | 189 | 237
7.04e-05 | 83 | 118 | 166 | 213 |  | 7.04e-05 | 94 | 142 | 189 | 213
7.06e-05 | 83 | 118 | 166 | 189 |  | 7.06e-05 | 94 | 118 | 166 | 189
7.08e-05 | 71 | 118 | 142 | 166 |  | 7.08e-05 | 83 | 118 | 142 | 189
7.1e-05 | 71 | 95 | 118 | 142 |  | 7.1e-05 | 71 | 95 | 142 | 166
7.12e-05 | 59 | 95 | 118 | 142 |  | 7.12e-05 | 71 | 95 | 118 | 142
7.14e-05 | 47 | 71 | 95 | 118 |  | 7.14e-05 | 59 | 95 | 118 | 142
7.16e-05 | 47 | 71 | 71 | 95 |  | 7.16e-05 | 59 | 71 | 95 | 118
7.18e-05 | 35 | 47 | 71 | 71 |  | 7.18e-05 | 47 | 71 | 95 | 95
7.2e-05 | 24 | 47 | 47 | 71 |  | 7.2e-05 | 35 | 47 | 71 | 71
7.22e-05 | 24 | 24 | 47 | 47 |  | 7.22e-05 | 24 | 47 | 47 | 71
7.24e-05 | 12 | 24 | 24 | 24 |  | 7.24e-05 | 24 | 24 | 47 | 47
7.26e-05 | 12 | 0 | 0 | 24 |  | 7.26e-05 | 12 | 24 | 24 | 24
7.28e-05 | 0 | 0 | 0 | 0 |  | 7.28e-05 | 12 | 0 | 0 | 0
7.3e-05 | 0 | -23 | -23 | -23 |  | 7.3e-05 | 0 | 0 | 0 | 0
7.32e-05 | -11 | -23 | -23 | -46 |  | 7.32e-05 | 0 | -23 | -23 | -23
7.34e-05 | -23 | -23 | -46 | -46 |  | 7.34e-05 | -11 | -23 | -23 | -46
7.36e-05 | -23 | -46 | -46 | -70 |  | 7.36e-05 | -23 | -23 | -46 | -46
7.38e-05 | -34 | -46 | -70 | -94 |  | 7.38e-05 | -23 | -46 | -46 | -70
7.4e-05 | -34 | -70 | -70 | -117 |  | 7.4e-05 | -34 | -46 | -70 | -94
7.42e-05 | -46 | -70 | -94 | -117 |  | 7.42e-05 | -34 | -70 | -94 | -117
7.44e-05 | -46 | -70 | -117 | -141 |  | 7.44e-05 | -46 | -70 | -94 | -117
7.46e-05 | -58 | -94 | -117 | -141 |  | 7.46e-05 | -46 | -94 | -117 | -141
7.48e-05 | -58 | -94 | -141 | -165 |  | 7.48e-05 | -58 | -94 | -117 | -165
7.5e-05 | -70 | -117 | -141 | -165 |  | 7.5e-05 | -70 | -94 | -141 | -165
7.52e-05 | -70 | -117 | -141 | -188 |  | 7.52e-05 | -70 | -117 | -141 | -188
7.54e-05 | -82 | -117 | -165 | -212 |  | 7.54e-05 | -70 | -117 | -141 | -188
7.56e-05 | -82 | -141 | -165 | -212 |  | 7.56e-05 | -82 | -117 | -165 | -212
7.58e-05 | -93 | -141 | -188 | -236 |  | 7.58e-05 | -82 | -141 | -165 | -212
7.6e-05 | -93 | -141 | -188 | -236 |  | 7.6e-05 | -93 | -141 | -188 | -236
7.62e-05 | -93 | -165 | -188 | -236 |  | 7.62e-05 | -93 | -141 | -188 | -236
7.64e-05 | -105 | -165 | -212 | -259 |  | 7.64e-05 | -105 | -141 | -188 | -236
7.66e-05 | -105 | -165 | -212 | -259 |  | 7.66e-05 | -105 | -165 | -212 | -259
7.68e-05 | -105 | -165 | -236 | -259 |  | 7.68e-05 | -105 | -165 | -212 | -259
7.7e-05 | -105 | -188 | -236 | -283 |  | 7.7e-05 | -105 | -165 | -212 | -283
7.72e-05 | -117 | -188 | -236 | -283 |  | 7.72e-05 | -117 | -188 | -236 | -283
7.74e-05 | -117 | -188 | -236 | -307 |  | 7.74e-05 | -117 | -188 | -236 | -283
7.76e-05 | -117 | -188 | -236 | -307 |  | 7.76e-05 | -117 | -188 | -236 | -307
7.78e-05 | -129 | -188 | -236 | -307 |  | 7.78e-05 | -117 | -188 | -236 | -307
7.8e-05 | -129 | -188 | -259 | -307 |  | 7.8e-05 | -129 | -188 | -236 | -307
7.82e-05 | -129 | -188 | -259 | -307 |  | 7.82e-05 | -129 | -188 | -259 | -307
7.84e-05 | -129 | -212 | -259 | -330 |  | 7.84e-05 | -129 | -188 | -259 | -307
7.86e-05 | -129 | -212 | -259 | -330 |  | 7.86e-05 | -129 | -212 | -259 | -330
7.88e-05 | -141 | -212 | -259 | -330 |  | 7.88e-05 | -129 | -212 | -259 | -330
7.9e-05 | -129 | -212 | -259 | -330 |  | 7.9e-05 | -141 | -212 | -259 | -330
7.92e-05 | -129 | -212 | -259 | -330 |  | 7.92e-05 | -129 | -212 | -259 | -330
7.94e-05 | -129 | -212 | -259 | -330 |  | 7.94e-05 | -141 | -212 | -259 | -330
7.96e-05 | -129 | -212 | -259 | -330 |  | 7.96e-05 | -141 | -212 | -283 | -330
7.98e-05 | -141 | -212 | -283 | -330 |  | 7.98e-05 | -141 | -212 | -259 | -330
8e-05 | -129 | -212 | -259 | -330 |  | 8e-05 | -129 | -212 | -259 | -330
8.02e-05 | -129 | -212 | -259 | -330 |  | 8.02e-05 | -141 | -212 | -259 | -330
8.04e-05 | -141 | -188 | -259 | -330 |  | 8.04e-05 | -141 | -212 | -259 | -330
8.06e-05 | -129 | -212 | -259 | -330 |  | 8.06e-05 | -129 | -212 | -259 | -330
8.08e-05 | -129 | -212 | -259 | -307 |  | 8.08e-05 | -129 | -212 | -259 | -330
8.1e-05 | -129 | -212 | -259 | -307 |  | 8.1e-05 | -129 | -212 | -259 | -307
8.12e-05 | -129 | -188 | -259 | -307 |  | 8.12e-05 | -129 | -188 | -259 | -307
8.14e-05 | -129 | -188 | -259 | -307 |  | 8.14e-05 | -129 | -188 | -259 | -307
8.16e-05 | -117 | -188 | -236 | -307 |  | 8.16e-05 | -129 | -188 | -259 | -307
8.18e-05 | -117 | -188 | -236 | -307 |  | 8.18e-05 | -117 | -188 | -259 | -307
8.2e-05 | -117 | -188 | -236 | -283 |  | 8.2e-05 | -117 | -188 | -236 | -307
8.22e-05 | -117 | -188 | -236 | -283 |  | 8.22e-05 | -117 | -188 | -236 | -283
8.24e-05 | -117 | -165 | -236 | -283 |  | 8.24e-05 | -117 | -188 | -236 | -283
8.26e-05 | -105 | -165 | -212 | -259 |  | 8.26e-05 | -117 | -165 | -236 | -283
8.28e-05 | -105 | -165 | -212 | -259 |  | 8.28e-05 | -117 | -165 | -212 | -259
8.3e-05 | -105 | -165 | -212 | -259 |  | 8.3e-05 | -105 | -165 | -212 | -259
8.32e-05 | -105 | -141 | -188 | -236 |  | 8.32e-05 | -105 | -165 | -212 | -259
8.34e-05 | -93 | -141 | -188 | -236 |  | 8.34e-05 | -105 | -165 | -188 | -236
8.36e-05 | -93 | -141 | -188 | -236 |  | 8.36e-05 | -93 | -165 | -188 | -236
8.38e-05 | -93 | -141 | -165 | -212 |  | 8.38e-05 | -93 | -141 | -188 | -212
8.4e-05 | -82 | -141 | -165 | -188 |  | 8.4e-05 | -93 | -141 | -165 | -212
8.42e-05 | -82 | -117 | -165 | -188 |  | 8.42e-05 | -82 | -141 | -165 | -212
8.44e-05 | -82 | -117 | -141 | -188 |  | 8.44e-05 | -82 | -117 | -165 | -188
8.46e-05 | -70 | -117 | -141 | -165 |  | 8.46e-05 | -82 | -117 | -141 | -188
8.48e-05 | -70 | -94 | -141 | -165 |  | 8.48e-05 | -70 | -117 | -141 | -165
8.5e-05 | -58 | -94 | -117 | -141 |  | 8.5e-05 | -70 | -94 | -117 | -165
8.52e-05 | -58 | -94 | -117 | -141 |  | 8.52e-05 | -58 | -94 | -117 | -141
8.54e-05 | -58 | -94 | -117 | -117 |  | 8.54e-05 | -58 | -94 | -117 | -141
8.56e-05 | -46 | -70 | -94 | -117 |  | 8.56e-05 | -58 | -70 | -94 | -117
8.58e-05 | -46 | -70 | -94 | -117 |  | 8.58e-05 | -46 | -70 | -94 | -117
8.6e-05 | -34 | -46 | -70 | -94 |  | 8.6e-05 | -46 | -70 | -94 | -117
8.62e-05 | -34 | -46 | -70 | -70 |  | 8.62e-05 | -34 | -46 | -70 | -94
8.64e-05 | -23 | -46 | -46 | -70 |  | 8.64e-05 | -34 | -46 | -70 | -70
8.66e-05 | -23 | -46 | -46 | -70 |  | 8.66e-05 | -34 | -46 | -46 | -70
8.68e-05 | -23 | -23 | -46 | -46 |  | 8.68e-05 | -23 | -46 | -46 | -46
8.7e-05 | -11 | -23 | -23 | -23 |  | 8.7e-05 | -23 | -23 | -46 | -46
8.72e-05 | -11 | -23 | -23 | -23 |  | 8.72e-05 | -11 | -23 | -23 | -23
8.74e-05 | -11 | -23 | -23 | 0 |  | 8.74e-05 | -11 | -23 | -23 | -23
8.76e-05 | 0 | 0 | 0 | 0 |  | 8.76e-05 | 0 | 0 | 0 | 0
8.78e-05 | 0 | 0 | 0 | 0 |  | 8.78e-05 | 0 | 0 | 0 | 0
8.8e-05 | 12 | 0 | 24 | 24 |  | 8.8e-05 | 0 | 0 | 0 | 24
8.82e-05 | 12 | 24 | 24 | 24 |  | 8.82e-05 | 12 | 0 | 24 | 24
8.84e-05 | 24 | 24 | 24 | 47 |  | 8.84e-05 | 12 | 24 | 24 | 24
8.86e-05 | 24 | 24 | 47 | 47 |  | 8.86e-05 | 24 | 24 | 47 | 47
8.88e-05 | 24 | 47 | 47 | 71 |  | 8.88e-05 | 24 | 24 | 47 | 47
8.9e-05 | 35 | 47 | 71 | 71 |  | 8.9e-05 | 24 | 47 | 47 | 71
8.92e-05 | 35 | 47 | 71 | 71 |  | 8.92e-05 | 35 | 47 | 71 | 71
8.94e-05 | 35 | 47 | 71 | 95 |  | 8.94e-05 | 35 | 47 | 71 | 95
8.96e-05 | 47 | 47 | 95 | 95 |  | 8.96e-05 | 35 | 47 | 71 | 95
8.98e-05 | 47 | 71 | 95 | 118 |  | 8.98e-05 | 35 | 71 | 71 | 95
9e-05 | 47 | 71 | 95 | 118 |  | 9e-05 | 35 | 71 | 95 | 118
9.02e-05 | 59 | 71 | 95 | 118 |  | 9.02e-05 | 47 | 71 | 95 | 118
9.04e-05 | 59 | 71 | 118 | 142 |  | 9.04e-05 | 47 | 71 | 95 | 142
9.06e-05 | 59 | 95 | 118 | 142 |  | 9.06e-05 | 59 | 95 | 118 | 142
9.08e-05 | 59 | 95 | 118 | 166 |  | 9.08e-05 | 59 | 95 | 118 | 142
9.1e-05 | 71 | 95 | 118 | 166 |  | 9.1e-05 | 59 | 95 | 118 | 142
9.12e-05 | 71 | 95 | 142 | 166 |  | 9.12e-05 | 59 | 95 | 118 | 166
9.14e-05 | 71 | 95 | 142 | 166 |  | 9.14e-05 | 71 | 95 | 142 | 166
9.16e-05 | 71 | 95 | 142 | 166 |  | 9.16e-05 | 71 | 95 | 142 | 166
9.18e-05 | 83 | 118 | 142 | 189 |  | 9.18e-05 | 71 | 118 | 142 | 189
9.2e-05 | 83 | 118 | 166 | 189 |  | 9.2e-05 | 71 | 118 | 142 | 189
9.22e-05 | 71 | 118 | 166 | 189 |  | 9.22e-05 | 83 | 118 | 142 | 189
9.24e-05 | 83 | 118 | 166 | 189 |  | 9.24e-05 | 83 | 118 | 142 | 189
9.26e-05 | 83 | 118 | 166 | 189 |  | 9.26e-05 | 83 | 118 | 166 | 189
9.28e-05 | 83 | 118 | 166 | 189 |  | 9.28e-05 | 83 | 118 | 166 | 189
9.3e-05 | 83 | 118 | 166 | 189 |  | 9.3e-05 | 83 | 118 | 166 | 189
9.32e-05 | 83 | 118 | 166 | 213 |  | 9.32e-05 | 83 | 118 | 166 | 213
9.34e-05 | 83 | 118 | 166 | 189 |  | 9.34e-05 | 83 | 118 | 166 | 213
9.36e-05 | 83 | 118 | 166 | 213 |  | 9.36e-05 | 83 | 118 | 166 | 213
9.38e-05 | 83 | 142 | 166 | 213 |  | 9.38e-05 | 83 | 118 | 166 | 213
9.4e-05 | 94 | 118 | 166 | 213 |  | 9.4e-05 | 83 | 118 | 166 | 213
9.42e-05 | 83 | 142 | 166 | 213 |  | 9.42e-05 | 94 | 118 | 166 | 213
9.44e-05 | 83 | 142 | 166 | 213 |  | 9.44e-05 | 83 | 142 | 166 | 213
9.46e-05 | 83 | 118 | 166 | 213 |  | 9.46e-05 | 83 | 118 | 166 | 213
9.48e-05 | 83 | 142 | 166 | 213 |  | 9.48e-05 | 83 | 142 | 166 | 213
9.5e-05 | 83 | 118 | 166 | 213 |  | 9.5e-05 | 83 | 118 | 166 | 213
9.52e-05 | 83 | 118 | 166 | 213 |  | 9.52e-05 | 83 | 142 | 166 | 213
9.54e-05 | 83 | 118 | 166 | 213 |  | 9.54e-05 | 83 | 118 | 166 | 213
9.56e-05 | 83 | 118 | 166 | 189 |  | 9.56e-05 | 83 | 142 | 166 | 213
9.58e-05 | 83 | 118 | 166 | 189 |  | 9.58e-05 | 83 | 118 | 166 | 189
9.6e-05 | 83 | 118 | 166 | 189 |  | 9.6e-05 | 83 | 118 | 166 | 189
9.62e-05 | 83 | 118 | 166 | 189 |  | 9.62e-05 | 83 | 118 | 166 | 189
9.64e-05 | 83 | 118 | 142 | 189 |  | 9.64e-05 | 83 | 118 | 166 | 189
9.66e-05 | 83 | 118 | 142 | 189 |  | 9.66e-05 | 83 | 118 | 142 | 189
9.68e-05 | 71 | 118 | 142 | 166 |  | 9.68e-05 | 83 | 118 | 142 | 189
9.7e-05 | 71 | 95 | 142 | 166 |  | 9.7e-05 | 83 | 118 | 142 | 189
9.72e-05 | 71 | 95 | 142 | 166 |  | 9.72e-05 | 71 | 118 | 142 | 166
9.74e-05 | 71 | 95 | 142 | 166 |  | 9.74e-05 | 71 | 95 | 142 | 166
9.76e-05 | 71 | 95 | 142 | 166 |  | 9.76e-05 | 71 | 95 | 142 | 166
9.78e-05 | 71 | 95 | 118 | 142 |  | 9.78e-05 | 71 | 95 | 118 | 166
9.8e-05 | 59 | 95 | 118 | 142 |  | 9.8e-05 | 71 | 95 | 118 | 142
9.82e-05 | 59 | 71 | 118 | 142 |  | 9.82e-05 | 59 | 95 | 118 | 142
9.84e-05 | 59 | 71 | 118 | 142 |  | 9.84e-05 | 59 | 95 | 118 | 142
9.86e-05 | 59 | 71 | 95 | 118 |  | 9.86e-05 | 59 | 71 | 118 | 142
9.88e-05 | 47 | 71 | 95 | 118 |  | 9.88e-05 | 59 | 71 | 95 | 118
9.9e-05 | 47 | 71 | 95 | 118 |  | 9.9e-05 | 47 | 71 | 95 | 118
9.92e-05 | 47 | 71 | 95 | 95 |  | 9.92e-05 | 47 | 71 | 95 | 118
9.94e-05 | 35 | 71 | 71 | 95 |  | 9.94e-05 | 47 | 71 | 95 | 95
9.96e-05 | 35 | 47 | 71 | 95 |  | 9.96e-05 | 35 | 47 | 71 | 95
9.98e-05 | 35 | 47 | 71 | 71 |  | 9.98e-05 | 35 | 47 | 71 | 95
0.0001 | 35 | 47 | 71 | 71 |  | 0.0001 | 35 | 47 | 71 | 95
0.0001002 | 24 | 47 | 47 | 71 |  | 0.0001002 | 35 | 47 | 47 | 71
0.0001004 | 24 | 24 | 47 | 71 |  | 0.0001004 | 35 | 47 | 47 | 71
0.0001006 | 24 | 24 | 47 | 47 |  | 0.0001006 | 24 | 47 | 47 | 71
0.0001008 | 24 | 24 | 47 | 47 |  | 0.0001008 | 24 | 24 | 47 | 47
0.000101 | 12 | 24 | 24 | 24 |  | 0.000101 | 24 | 24 | 47 | 47
0.0001012 | 12 | 24 | 24 | 24 |  | 0.0001012 | 12 | 24 | 24 | 24
0.0001014 | 12 | 24 | 24 | 24 |  | 0.0001014 | 12 | 24 | 24 | 24
0.0001016 | 12 | 0 | 0 | 0 |  | 0.0001016 | 12 | 24 | 24 | 24
0.0001018 | 12 | 0 | 0 | 0 |  | 0.0001018 | 12 | 0 | 24 | 24
0.000102 | 0 | 0 | 0 | 0 |  | 0.000102 | 0 | 0 | 0 | 0
0.0001022 | 0 | 0 | 0 | 0 |  | 0.0001022 | 0 | 0 | 0 | 0
| 2 | 2 | 2 | 2 |  |  | 2 | 2 | 2 | 2
| 1000 | 1500 | 2000 | 2500 |  |  | 1000 | 1500 | 2000 | 2500
| 2 | 2 | 2 | 2 |  |  | 2 | 2 | 2 | 2

## Step 2 150 Impulse
Time (s) | 1000 V | 1500 V | 2000 V | 2500 V | 150 | Time (s) | 1000 V | 1500 V | 2000 V | 2500 V |  | Notes
0 | 0 | 0 | 0 | 0 | A17 Coil Pack QXFA 150 Outer to Inner | 0 | 0 | 0 | 0 | 0 | A17 Coil Pack QXFA 150 Outer to Inner
2e-07 | 0 | 0 | 0 | 0 |  | 2e-07 | 0 | 0 | 0 | 0
4e-07 | 0 | 0 | 0 | 0 |  | 4e-07 | 0 | 0 | 0 | 0
6e-07 | 154 | 308 | 450 | 592 |  | 6e-07 | 154 | 308 | 426 | 592
8e-07 | 626 | 1065 | 1514 | 2035 |  | 8e-07 | 626 | 1065 | 1538 | 2035
1e-06 | 980 | 1514 | 2011 | 2484 |  | 1e-06 | 980 | 1491 | 2011 | 2484
1.2e-06 | 945 | 1372 | 1751 | 2130 |  | 1.2e-06 | 945 | 1372 | 1775 | 2130
1.4e-06 | 850 | 1254 | 1633 | 2011 |  | 1.4e-06 | 850 | 1254 | 1633 | 2011
1.6e-06 | 850 | 1278 | 1656 | 2059 |  | 1.6e-06 | 850 | 1254 | 1656 | 2059
1.8e-06 | 850 | 1254 | 1633 | 2011 |  | 1.8e-06 | 850 | 1254 | 1633 | 2035
2e-06 | 827 | 1207 | 1585 | 1940 |  | 2e-06 | 827 | 1207 | 1585 | 1964
2.2e-06 | 815 | 1183 | 1562 | 1917 |  | 2.2e-06 | 815 | 1183 | 1562 | 1940
2.4e-06 | 803 | 1183 | 1562 | 1893 |  | 2.4e-06 | 803 | 1183 | 1562 | 1917
2.6e-06 | 791 | 1159 | 1538 | 1869 |  | 2.6e-06 | 791 | 1159 | 1538 | 1893
2.8e-06 | 780 | 1159 | 1491 | 1846 |  | 2.8e-06 | 780 | 1159 | 1514 | 1846
3e-06 | 768 | 1112 | 1467 | 1798 |  | 3e-06 | 768 | 1136 | 1467 | 1822
3.2e-06 | 744 | 1088 | 1443 | 1775 |  | 3.2e-06 | 744 | 1088 | 1443 | 1775
3.4e-06 | 720 | 1065 | 1396 | 1704 |  | 3.4e-06 | 720 | 1065 | 1396 | 1727
3.6e-06 | 709 | 1041 | 1349 | 1680 |  | 3.6e-06 | 709 | 1041 | 1372 | 1680
3.8e-06 | 685 | 1017 | 1325 | 1633 |  | 3.8e-06 | 685 | 1017 | 1325 | 1633
4e-06 | 673 | 994 | 1301 | 1609 |  | 4e-06 | 673 | 994 | 1301 | 1609
4.2e-06 | 661 | 970 | 1278 | 1562 |  | 4.2e-06 | 661 | 970 | 1278 | 1562
4.4e-06 | 650 | 946 | 1254 | 1538 |  | 4.4e-06 | 650 | 970 | 1254 | 1538
4.6e-06 | 638 | 946 | 1230 | 1514 |  | 4.6e-06 | 638 | 946 | 1230 | 1514
4.8e-06 | 626 | 923 | 1207 | 1491 |  | 4.8e-06 | 626 | 923 | 1207 | 1491
5e-06 | 614 | 899 | 1183 | 1467 |  | 5e-06 | 614 | 899 | 1183 | 1467
5.2e-06 | 602 | 899 | 1159 | 1443 |  | 5.2e-06 | 602 | 899 | 1159 | 1443
5.4e-06 | 591 | 875 | 1136 | 1396 |  | 5.4e-06 | 591 | 875 | 1136 | 1396
5.6e-06 | 579 | 852 | 1112 | 1372 |  | 5.6e-06 | 579 | 852 | 1112 | 1372
5.8e-06 | 555 | 828 | 1065 | 1325 |  | 5.8e-06 | 555 | 828 | 1088 | 1325
6e-06 | 543 | 804 | 1041 | 1278 |  | 6e-06 | 543 | 804 | 1041 | 1278
6.2e-06 | 531 | 781 | 1017 | 1230 |  | 6.2e-06 | 531 | 781 | 1017 | 1254
6.4e-06 | 508 | 734 | 970 | 1183 |  | 6.4e-06 | 508 | 734 | 970 | 1207
6.6e-06 | 484 | 734 | 946 | 1159 |  | 6.6e-06 | 496 | 710 | 946 | 1159
6.8e-06 | 472 | 686 | 899 | 1112 |  | 6.8e-06 | 472 | 686 | 899 | 1112
7e-06 | 449 | 663 | 875 | 1065 |  | 7e-06 | 449 | 663 | 875 | 1065
7.2e-06 | 437 | 615 | 828 | 994 |  | 7.2e-06 | 437 | 615 | 828 | 1017
7.4e-06 | 413 | 592 | 781 | 946 |  | 7.4e-06 | 413 | 592 | 781 | 970
7.6e-06 | 390 | 568 | 734 | 899 |  | 7.6e-06 | 390 | 568 | 734 | 899
7.8e-06 | 366 | 521 | 686 | 852 |  | 7.8e-06 | 366 | 521 | 686 | 852
8e-06 | 343 | 497 | 639 | 781 |  | 8e-06 | 343 | 497 | 639 | 804
8.2e-06 | 319 | 450 | 592 | 734 |  | 8.2e-06 | 319 | 450 | 592 | 734
8.4e-06 | 295 | 402 | 544 | 663 |  | 8.4e-06 | 295 | 426 | 544 | 663
8.6e-06 | 272 | 379 | 497 | 615 |  | 8.6e-06 | 260 | 379 | 497 | 615
8.8e-06 | 236 | 355 | 450 | 544 |  | 8.8e-06 | 248 | 355 | 450 | 544
9e-06 | 213 | 308 | 402 | 497 |  | 9e-06 | 213 | 308 | 402 | 497
9.2e-06 | 189 | 260 | 355 | 426 |  | 9.2e-06 | 189 | 260 | 355 | 426
9.4e-06 | 165 | 237 | 308 | 379 |  | 9.4e-06 | 165 | 237 | 308 | 379
9.6e-06 | 142 | 213 | 260 | 308 |  | 9.6e-06 | 142 | 213 | 260 | 308
9.8e-06 | 118 | 166 | 213 | 260 |  | 9.8e-06 | 118 | 166 | 213 | 260
1e-05 | 94 | 142 | 166 | 213 |  | 1e-05 | 94 | 142 | 166 | 213
1.02e-05 | 71 | 95 | 118 | 142 |  | 1.02e-05 | 71 | 95 | 118 | 142
1.04e-05 | 47 | 47 | 71 | 71 |  | 1.04e-05 | 47 | 71 | 71 | 95
1.06e-05 | 24 | 24 | 24 | 24 |  | 1.06e-05 | 24 | 24 | 24 | 24
1.08e-05 | 0 | -23 | -23 | -23 |  | 1.08e-05 | 0 | 0 | -23 | -23
1.1e-05 | -23 | -46 | -70 | -70 |  | 1.1e-05 | -23 | -46 | -70 | -70
1.12e-05 | -46 | -70 | -117 | -141 |  | 1.12e-05 | -46 | -70 | -94 | -141
1.14e-05 | -70 | -117 | -141 | -188 |  | 1.14e-05 | -70 | -117 | -165 | -188
1.16e-05 | -93 | -141 | -188 | -236 |  | 1.16e-05 | -93 | -141 | -188 | -236
1.18e-05 | -117 | -188 | -236 | -307 |  | 1.18e-05 | -117 | -188 | -236 | -283
1.2e-05 | -129 | -212 | -283 | -354 |  | 1.2e-05 | -141 | -212 | -283 | -330
1.22e-05 | -153 | -236 | -307 | -378 |  | 1.22e-05 | -153 | -236 | -307 | -401
1.24e-05 | -176 | -283 | -354 | -449 |  | 1.24e-05 | -176 | -283 | -354 | -449
1.26e-05 | -200 | -307 | -401 | -496 |  | 1.26e-05 | -200 | -307 | -401 | -496
1.28e-05 | -212 | -330 | -449 | -543 |  | 1.28e-05 | -212 | -330 | -449 | -543
1.3e-05 | -235 | -354 | -496 | -591 |  | 1.3e-05 | -235 | -354 | -472 | -591
1.32e-05 | -259 | -401 | -520 | -638 |  | 1.32e-05 | -259 | -401 | -520 | -638
1.34e-05 | -271 | -425 | -567 | -685 |  | 1.34e-05 | -282 | -425 | -567 | -685
1.36e-05 | -294 | -449 | -591 | -733 |  | 1.36e-05 | -294 | -449 | -591 | -756
1.38e-05 | -318 | -472 | -638 | -780 |  | 1.38e-05 | -318 | -496 | -638 | -780
1.4e-05 | -330 | -520 | -662 | -827 |  | 1.4e-05 | -342 | -520 | -662 | -827
1.42e-05 | -353 | -543 | -709 | -874 |  | 1.42e-05 | -353 | -543 | -709 | -874
1.44e-05 | -377 | -567 | -733 | -922 |  | 1.44e-05 | -377 | -567 | -733 | -922
1.46e-05 | -389 | -591 | -780 | -969 |  | 1.46e-05 | -389 | -591 | -780 | -969
1.48e-05 | -412 | -614 | -827 | -993 |  | 1.48e-05 | -412 | -614 | -803 | -993
1.5e-05 | -424 | -638 | -851 | -1040 |  | 1.5e-05 | -424 | -638 | -851 | -1040
1.52e-05 | -448 | -662 | -874 | -1087 |  | 1.52e-05 | -448 | -662 | -874 | -1087
1.54e-05 | -460 | -685 | -922 | -1135 |  | 1.54e-05 | -460 | -685 | -922 | -1135
1.56e-05 | -471 | -709 | -945 | -1158 |  | 1.56e-05 | -483 | -709 | -945 | -1158
1.58e-05 | -495 | -733 | -969 | -1182 |  | 1.58e-05 | -495 | -733 | -969 | -1206
1.6e-05 | -507 | -756 | -993 | -1229 |  | 1.6e-05 | -507 | -756 | -993 | -1229
1.62e-05 | -519 | -780 | -1016 | -1277 |  | 1.62e-05 | -519 | -780 | -1016 | -1277
1.64e-05 | -530 | -803 | -1064 | -1300 |  | 1.64e-05 | -530 | -803 | -1064 | -1300
1.66e-05 | -542 | -827 | -1087 | -1324 |  | 1.66e-05 | -554 | -827 | -1087 | -1324
1.68e-05 | -554 | -851 | -1111 | -1348 |  | 1.68e-05 | -566 | -851 | -1111 | -1348
1.7e-05 | -566 | -874 | -1135 | -1395 |  | 1.7e-05 | -578 | -874 | -1135 | -1395
1.72e-05 | -590 | -874 | -1158 | -1419 |  | 1.72e-05 | -578 | -874 | -1158 | -1419
1.74e-05 | -590 | -898 | -1182 | -1442 |  | 1.74e-05 | -590 | -898 | -1182 | -1442
1.76e-05 | -601 | -922 | -1182 | -1466 |  | 1.76e-05 | -601 | -922 | -1182 | -1466
1.78e-05 | -613 | -922 | -1206 | -1490 |  | 1.78e-05 | -613 | -922 | -1206 | -1490
1.8e-05 | -625 | -945 | -1229 | -1513 |  | 1.8e-05 | -625 | -945 | -1229 | -1513
1.82e-05 | -637 | -945 | -1253 | -1537 |  | 1.82e-05 | -637 | -945 | -1253 | -1537
1.84e-05 | -637 | -969 | -1253 | -1537 |  | 1.84e-05 | -637 | -969 | -1277 | -1537
1.86e-05 | -649 | -969 | -1277 | -1561 |  | 1.86e-05 | -649 | -969 | -1277 | -1561
1.88e-05 | -660 | -969 | -1277 | -1584 |  | 1.88e-05 | -649 | -969 | -1277 | -1584
1.9e-05 | -660 | -993 | -1300 | -1584 |  | 1.9e-05 | -660 | -993 | -1300 | -1584
1.92e-05 | -660 | -993 | -1300 | -1608 |  | 1.92e-05 | -660 | -993 | -1300 | -1608
1.94e-05 | -672 | -993 | -1300 | -1608 |  | 1.94e-05 | -672 | -1016 | -1324 | -1632
1.96e-05 | -672 | -1016 | -1324 | -1632 |  | 1.96e-05 | -672 | -1016 | -1324 | -1632
1.98e-05 | -672 | -1016 | -1324 | -1632 |  | 1.98e-05 | -672 | -1016 | -1324 | -1632
2e-05 | -672 | -1016 | -1324 | -1632 |  | 2e-05 | -684 | -1016 | -1324 | -1632
2.02e-05 | -672 | -1016 | -1324 | -1655 |  | 2.02e-05 | -684 | -1016 | -1324 | -1655
2.04e-05 | -684 | -1016 | -1324 | -1655 |  | 2.04e-05 | -684 | -1016 | -1348 | -1655
2.06e-05 | -684 | -1016 | -1324 | -1655 |  | 2.06e-05 | -684 | -1016 | -1348 | -1655
2.08e-05 | -684 | -1016 | -1348 | -1632 |  | 2.08e-05 | -684 | -1016 | -1348 | -1655
2.1e-05 | -684 | -1016 | -1324 | -1632 |  | 2.1e-05 | -684 | -1016 | -1348 | -1655
2.12e-05 | -672 | -1016 | -1324 | -1632 |  | 2.12e-05 | -684 | -1016 | -1324 | -1632
2.14e-05 | -672 | -993 | -1324 | -1632 |  | 2.14e-05 | -672 | -1016 | -1324 | -1608
2.16e-05 | -660 | -993 | -1300 | -1608 |  | 2.16e-05 | -672 | -993 | -1300 | -1608
2.18e-05 | -660 | -993 | -1277 | -1561 |  | 2.18e-05 | -660 | -993 | -1300 | -1584
2.2e-05 | -649 | -969 | -1277 | -1561 |  | 2.2e-05 | -649 | -969 | -1277 | -1561
2.22e-05 | -637 | -945 | -1253 | -1513 |  | 2.22e-05 | -649 | -969 | -1253 | -1537
2.24e-05 | -625 | -945 | -1229 | -1490 |  | 2.24e-05 | -625 | -945 | -1229 | -1490
2.26e-05 | -613 | -922 | -1182 | -1466 |  | 2.26e-05 | -625 | -922 | -1206 | -1466
2.28e-05 | -601 | -898 | -1158 | -1419 |  | 2.28e-05 | -601 | -898 | -1158 | -1442
2.3e-05 | -590 | -874 | -1135 | -1395 |  | 2.3e-05 | -590 | -874 | -1135 | -1395
2.32e-05 | -566 | -851 | -1111 | -1348 |  | 2.32e-05 | -566 | -851 | -1111 | -1371
2.34e-05 | -554 | -827 | -1087 | -1324 |  | 2.34e-05 | -554 | -827 | -1087 | -1324
2.36e-05 | -542 | -803 | -1040 | -1277 |  | 2.36e-05 | -542 | -803 | -1040 | -1300
2.38e-05 | -519 | -780 | -1016 | -1229 |  | 2.38e-05 | -519 | -780 | -1016 | -1253
2.4e-05 | -495 | -756 | -969 | -1182 |  | 2.4e-05 | -507 | -756 | -969 | -1182
2.42e-05 | -483 | -709 | -922 | -1135 |  | 2.42e-05 | -483 | -733 | -945 | -1158
2.44e-05 | -460 | -685 | -898 | -1087 |  | 2.44e-05 | -460 | -685 | -898 | -1087
2.46e-05 | -436 | -662 | -851 | -1040 |  | 2.46e-05 | -448 | -662 | -874 | -1040
2.48e-05 | -412 | -614 | -803 | -993 |  | 2.48e-05 | -424 | -614 | -803 | -993
2.5e-05 | -389 | -591 | -756 | -945 |  | 2.5e-05 | -401 | -591 | -756 | -945
2.52e-05 | -377 | -543 | -709 | -874 |  | 2.52e-05 | -377 | -567 | -733 | -874
2.54e-05 | -353 | -520 | -662 | -803 |  | 2.54e-05 | -353 | -520 | -685 | -827
2.56e-05 | -330 | -472 | -614 | -756 |  | 2.56e-05 | -330 | -496 | -638 | -756
2.58e-05 | -306 | -449 | -567 | -709 |  | 2.58e-05 | -306 | -449 | -567 | -709
2.6e-05 | -282 | -401 | -520 | -638 |  | 2.6e-05 | -271 | -401 | -520 | -638
2.62e-05 | -247 | -378 | -472 | -567 |  | 2.62e-05 | -247 | -378 | -472 | -567
2.64e-05 | -223 | -330 | -425 | -520 |  | 2.64e-05 | -223 | -330 | -425 | -520
2.66e-05 | -200 | -307 | -378 | -449 |  | 2.66e-05 | -200 | -307 | -378 | -449
2.68e-05 | -176 | -259 | -330 | -401 |  | 2.68e-05 | -176 | -259 | -330 | -401
2.7e-05 | -153 | -212 | -283 | -330 |  | 2.7e-05 | -153 | -212 | -283 | -330
2.72e-05 | -117 | -188 | -212 | -259 |  | 2.72e-05 | -117 | -188 | -236 | -283
2.74e-05 | -93 | -141 | -165 | -212 |  | 2.74e-05 | -105 | -141 | -188 | -212
2.76e-05 | -70 | -94 | -117 | -141 |  | 2.76e-05 | -70 | -117 | -141 | -165
2.78e-05 | -46 | -70 | -70 | -94 |  | 2.78e-05 | -46 | -70 | -70 | -94
2.8e-05 | -11 | -23 | -23 | -23 |  | 2.8e-05 | -23 | -23 | -23 | -23
2.82e-05 | 12 | 24 | 24 | 47 |  | 2.82e-05 | 12 | 0 | 24 | 47
2.84e-05 | 35 | 47 | 71 | 118 |  | 2.84e-05 | 24 | 47 | 71 | 95
2.86e-05 | 59 | 95 | 118 | 166 |  | 2.86e-05 | 59 | 95 | 118 | 166
2.88e-05 | 83 | 142 | 166 | 213 |  | 2.88e-05 | 83 | 118 | 166 | 213
2.9e-05 | 106 | 166 | 213 | 260 |  | 2.9e-05 | 106 | 166 | 213 | 260
2.92e-05 | 130 | 189 | 260 | 331 |  | 2.92e-05 | 130 | 189 | 260 | 331
2.94e-05 | 154 | 237 | 308 | 379 |  | 2.94e-05 | 142 | 237 | 308 | 379
2.96e-05 | 177 | 260 | 355 | 426 |  | 2.96e-05 | 165 | 260 | 355 | 426
2.98e-05 | 189 | 284 | 379 | 497 |  | 2.98e-05 | 189 | 284 | 379 | 497
3e-05 | 213 | 331 | 426 | 544 |  | 3e-05 | 213 | 331 | 426 | 544
3.02e-05 | 236 | 355 | 473 | 592 |  | 3.02e-05 | 236 | 355 | 473 | 592
3.04e-05 | 260 | 379 | 521 | 639 |  | 3.04e-05 | 248 | 379 | 497 | 639
3.06e-05 | 272 | 402 | 544 | 663 |  | 3.06e-05 | 272 | 402 | 544 | 663
3.08e-05 | 295 | 426 | 568 | 710 |  | 3.08e-05 | 283 | 426 | 568 | 710
3.1e-05 | 307 | 450 | 615 | 757 |  | 3.1e-05 | 307 | 450 | 615 | 757
3.12e-05 | 319 | 473 | 639 | 804 |  | 3.12e-05 | 319 | 473 | 639 | 804
3.14e-05 | 343 | 497 | 663 | 828 |  | 3.14e-05 | 331 | 497 | 663 | 828
3.16e-05 | 354 | 521 | 686 | 852 |  | 3.16e-05 | 354 | 521 | 686 | 852
3.18e-05 | 366 | 544 | 710 | 899 |  | 3.18e-05 | 366 | 544 | 710 | 899
3.2e-05 | 378 | 568 | 734 | 923 |  | 3.2e-05 | 378 | 544 | 734 | 923
3.22e-05 | 390 | 568 | 757 | 946 |  | 3.22e-05 | 390 | 568 | 781 | 946
3.24e-05 | 402 | 592 | 781 | 970 |  | 3.24e-05 | 402 | 592 | 781 | 970
3.26e-05 | 413 | 615 | 804 | 994 |  | 3.26e-05 | 402 | 592 | 804 | 994
3.28e-05 | 413 | 615 | 804 | 1017 |  | 3.28e-05 | 413 | 615 | 804 | 1017
3.3e-05 | 425 | 639 | 828 | 1017 |  | 3.3e-05 | 425 | 639 | 828 | 1041
3.32e-05 | 437 | 639 | 852 | 1041 |  | 3.32e-05 | 437 | 639 | 852 | 1041
3.34e-05 | 437 | 639 | 852 | 1065 |  | 3.34e-05 | 437 | 639 | 852 | 1065
3.36e-05 | 449 | 663 | 875 | 1065 |  | 3.36e-05 | 449 | 663 | 875 | 1065
3.38e-05 | 449 | 663 | 875 | 1065 |  | 3.38e-05 | 449 | 663 | 875 | 1065
3.4e-05 | 449 | 663 | 875 | 1065 |  | 3.4e-05 | 449 | 663 | 875 | 1088
3.42e-05 | 449 | 663 | 875 | 1065 |  | 3.42e-05 | 449 | 663 | 875 | 1088
3.44e-05 | 449 | 663 | 875 | 1088 |  | 3.44e-05 | 449 | 663 | 875 | 1088
3.46e-05 | 449 | 663 | 875 | 1065 |  | 3.46e-05 | 449 | 663 | 875 | 1088
3.48e-05 | 449 | 663 | 875 | 1065 |  | 3.48e-05 | 449 | 663 | 875 | 1065
3.5e-05 | 449 | 639 | 875 | 1065 |  | 3.5e-05 | 449 | 663 | 875 | 1065
3.52e-05 | 437 | 639 | 852 | 1065 |  | 3.52e-05 | 449 | 639 | 852 | 1065
3.54e-05 | 437 | 639 | 852 | 1041 |  | 3.54e-05 | 437 | 639 | 852 | 1041
3.56e-05 | 425 | 639 | 828 | 1017 |  | 3.56e-05 | 437 | 639 | 828 | 1041
3.58e-05 | 425 | 639 | 828 | 1017 |  | 3.58e-05 | 425 | 615 | 828 | 1017
3.6e-05 | 413 | 615 | 804 | 994 |  | 3.6e-05 | 413 | 615 | 804 | 994
3.62e-05 | 413 | 615 | 804 | 970 |  | 3.62e-05 | 413 | 592 | 804 | 994
3.64e-05 | 402 | 592 | 781 | 946 |  | 3.64e-05 | 402 | 592 | 781 | 970
3.66e-05 | 390 | 568 | 757 | 923 |  | 3.66e-05 | 390 | 568 | 757 | 946
3.68e-05 | 378 | 568 | 734 | 899 |  | 3.68e-05 | 378 | 568 | 734 | 899
3.7e-05 | 366 | 544 | 710 | 875 |  | 3.7e-05 | 366 | 544 | 710 | 899
3.72e-05 | 354 | 521 | 686 | 852 |  | 3.72e-05 | 366 | 521 | 710 | 875
3.74e-05 | 343 | 521 | 663 | 828 |  | 3.74e-05 | 343 | 521 | 686 | 828
3.76e-05 | 331 | 497 | 639 | 781 |  | 3.76e-05 | 331 | 497 | 639 | 804
3.78e-05 | 319 | 473 | 615 | 757 |  | 3.78e-05 | 319 | 473 | 615 | 757
3.8e-05 | 307 | 450 | 592 | 710 |  | 3.8e-05 | 307 | 450 | 592 | 734
3.82e-05 | 295 | 426 | 544 | 686 |  | 3.82e-05 | 295 | 426 | 568 | 686
3.84e-05 | 272 | 402 | 521 | 639 |  | 3.84e-05 | 272 | 402 | 521 | 663
3.86e-05 | 260 | 379 | 497 | 615 |  | 3.86e-05 | 260 | 379 | 497 | 615
3.88e-05 | 248 | 355 | 473 | 568 |  | 3.88e-05 | 248 | 355 | 473 | 568
3.9e-05 | 224 | 331 | 426 | 521 |  | 3.9e-05 | 224 | 331 | 426 | 544
3.92e-05 | 213 | 308 | 402 | 497 |  | 3.92e-05 | 213 | 308 | 402 | 497
3.94e-05 | 201 | 284 | 355 | 450 |  | 3.94e-05 | 189 | 284 | 379 | 450
3.96e-05 | 177 | 260 | 331 | 402 |  | 3.96e-05 | 177 | 260 | 331 | 426
3.98e-05 | 165 | 237 | 308 | 379 |  | 3.98e-05 | 165 | 237 | 308 | 379
4e-05 | 142 | 213 | 260 | 331 |  | 4e-05 | 142 | 213 | 260 | 331
4.02e-05 | 130 | 189 | 237 | 284 |  | 4.02e-05 | 130 | 189 | 237 | 308
4.04e-05 | 106 | 166 | 189 | 237 |  | 4.04e-05 | 118 | 166 | 213 | 260
4.06e-05 | 94 | 142 | 166 | 213 |  | 4.06e-05 | 94 | 142 | 166 | 213
4.08e-05 | 71 | 95 | 142 | 166 |  | 4.08e-05 | 83 | 118 | 142 | 166
4.1e-05 | 59 | 71 | 95 | 118 |  | 4.1e-05 | 59 | 71 | 118 | 142
4.12e-05 | 35 | 47 | 71 | 71 |  | 4.12e-05 | 47 | 47 | 71 | 71
4.14e-05 | 24 | 24 | 24 | 47 |  | 4.14e-05 | 24 | 24 | 24 | 47
4.16e-05 | 0 | 0 | 0 | 0 |  | 4.16e-05 | 12 | 0 | 0 | 0
4.18e-05 | -11 | -23 | -46 | -46 |  | 4.18e-05 | -11 | -23 | -23 | -46
4.2e-05 | -23 | -46 | -70 | -94 |  | 4.2e-05 | -23 | -46 | -70 | -94
4.22e-05 | -46 | -70 | -94 | -141 |  | 4.22e-05 | -46 | -70 | -94 | -141
4.24e-05 | -58 | -117 | -141 | -188 |  | 4.24e-05 | -58 | -94 | -141 | -165
4.26e-05 | -82 | -117 | -165 | -212 |  | 4.26e-05 | -82 | -117 | -165 | -212
4.28e-05 | -93 | -141 | -212 | -259 |  | 4.28e-05 | -93 | -141 | -188 | -236
4.3e-05 | -105 | -165 | -236 | -283 |  | 4.3e-05 | -105 | -165 | -236 | -283
4.32e-05 | -129 | -188 | -259 | -307 |  | 4.32e-05 | -129 | -188 | -259 | -307
4.34e-05 | -141 | -212 | -283 | -354 |  | 4.34e-05 | -141 | -212 | -283 | -354
4.36e-05 | -153 | -236 | -307 | -378 |  | 4.36e-05 | -153 | -236 | -307 | -378
4.38e-05 | -164 | -259 | -330 | -425 |  | 4.38e-05 | -164 | -259 | -330 | -401
4.4e-05 | -176 | -283 | -354 | -449 |  | 4.4e-05 | -176 | -283 | -354 | -449
4.42e-05 | -188 | -307 | -378 | -472 |  | 4.42e-05 | -188 | -307 | -378 | -472
4.44e-05 | -200 | -307 | -401 | -520 |  | 4.44e-05 | -200 | -307 | -401 | -520
4.46e-05 | -212 | -330 | -425 | -543 |  | 4.46e-05 | -212 | -330 | -425 | -520
4.48e-05 | -223 | -330 | -449 | -567 |  | 4.48e-05 | -223 | -354 | -449 | -567
4.5e-05 | -235 | -354 | -472 | -591 |  | 4.5e-05 | -235 | -354 | -472 | -591
4.52e-05 | -247 | -378 | -496 | -614 |  | 4.52e-05 | -247 | -378 | -496 | -591
4.54e-05 | -259 | -378 | -520 | -638 |  | 4.54e-05 | -259 | -378 | -496 | -614
4.56e-05 | -259 | -401 | -520 | -638 |  | 4.56e-05 | -259 | -401 | -520 | -638
4.58e-05 | -271 | -401 | -543 | -662 |  | 4.58e-05 | -271 | -401 | -543 | -662
4.6e-05 | -282 | -425 | -567 | -685 |  | 4.6e-05 | -282 | -425 | -543 | -685
4.62e-05 | -282 | -425 | -567 | -685 |  | 4.62e-05 | -282 | -425 | -567 | -709
4.64e-05 | -294 | -449 | -567 | -709 |  | 4.64e-05 | -294 | -425 | -567 | -709
4.66e-05 | -294 | -449 | -591 | -733 |  | 4.66e-05 | -294 | -449 | -591 | -733
4.68e-05 | -306 | -449 | -591 | -733 |  | 4.68e-05 | -306 | -449 | -591 | -733
4.7e-05 | -306 | -449 | -591 | -733 |  | 4.7e-05 | -306 | -449 | -591 | -733
4.72e-05 | -306 | -449 | -591 | -733 |  | 4.72e-05 | -306 | -472 | -614 | -756
4.74e-05 | -306 | -472 | -614 | -756 |  | 4.74e-05 | -306 | -472 | -614 | -756
4.76e-05 | -306 | -472 | -614 | -756 |  | 4.76e-05 | -318 | -472 | -614 | -756
4.78e-05 | -318 | -472 | -614 | -756 |  | 4.78e-05 | -306 | -472 | -614 | -756
4.8e-05 | -306 | -472 | -614 | -756 |  | 4.8e-05 | -318 | -472 | -614 | -756
4.82e-05 | -306 | -472 | -614 | -756 |  | 4.82e-05 | -306 | -472 | -614 | -756
4.84e-05 | -306 | -472 | -614 | -756 |  | 4.84e-05 | -318 | -472 | -614 | -756
4.86e-05 | -306 | -472 | -591 | -756 |  | 4.86e-05 | -306 | -472 | -614 | -756
4.88e-05 | -306 | -449 | -591 | -733 |  | 4.88e-05 | -306 | -449 | -614 | -733
4.9e-05 | -306 | -449 | -591 | -733 |  | 4.9e-05 | -306 | -449 | -591 | -733
4.92e-05 | -294 | -449 | -591 | -733 |  | 4.92e-05 | -294 | -449 | -591 | -733
4.94e-05 | -294 | -449 | -591 | -709 |  | 4.94e-05 | -294 | -449 | -591 | -709
4.96e-05 | -294 | -425 | -567 | -709 |  | 4.96e-05 | -294 | -425 | -591 | -709
4.98e-05 | -282 | -425 | -567 | -685 |  | 4.98e-05 | -294 | -425 | -567 | -685
5e-05 | -282 | -425 | -543 | -662 |  | 5e-05 | -282 | -425 | -567 | -685
5.02e-05 | -271 | -401 | -543 | -662 |  | 5.02e-05 | -271 | -401 | -543 | -662
5.04e-05 | -271 | -401 | -520 | -638 |  | 5.04e-05 | -271 | -401 | -520 | -638
5.06e-05 | -259 | -378 | -520 | -614 |  | 5.06e-05 | -259 | -378 | -520 | -638
5.08e-05 | -247 | -378 | -496 | -591 |  | 5.08e-05 | -259 | -378 | -496 | -614
5.1e-05 | -247 | -354 | -472 | -591 |  | 5.1e-05 | -247 | -354 | -496 | -591
5.12e-05 | -235 | -354 | -449 | -567 |  | 5.12e-05 | -235 | -354 | -472 | -567
5.14e-05 | -223 | -330 | -449 | -543 |  | 5.14e-05 | -223 | -354 | -449 | -543
5.16e-05 | -212 | -330 | -425 | -520 |  | 5.16e-05 | -212 | -330 | -425 | -520
5.18e-05 | -200 | -307 | -401 | -496 |  | 5.18e-05 | -212 | -307 | -401 | -496
5.2e-05 | -200 | -283 | -378 | -449 |  | 5.2e-05 | -200 | -307 | -378 | -472
5.22e-05 | -188 | -283 | -354 | -425 |  | 5.22e-05 | -188 | -283 | -378 | -449
5.24e-05 | -176 | -259 | -330 | -425 |  | 5.24e-05 | -176 | -259 | -354 | -425
5.26e-05 | -164 | -259 | -307 | -378 |  | 5.26e-05 | -164 | -259 | -330 | -401
5.28e-05 | -153 | -236 | -307 | -354 |  | 5.28e-05 | -153 | -236 | -307 | -354
5.3e-05 | -141 | -212 | -283 | -330 |  | 5.3e-05 | -141 | -212 | -283 | -330
5.32e-05 | -129 | -188 | -259 | -307 |  | 5.32e-05 | -129 | -212 | -259 | -307
5.34e-05 | -117 | -188 | -236 | -283 |  | 5.34e-05 | -129 | -188 | -236 | -283
5.36e-05 | -105 | -165 | -212 | -259 |  | 5.36e-05 | -117 | -165 | -212 | -259
5.38e-05 | -93 | -141 | -188 | -212 |  | 5.38e-05 | -105 | -141 | -188 | -236
5.4e-05 | -82 | -117 | -165 | -188 |  | 5.4e-05 | -82 | -141 | -165 | -188
5.42e-05 | -70 | -117 | -141 | -165 |  | 5.42e-05 | -70 | -117 | -141 | -165
5.44e-05 | -58 | -94 | -117 | -141 |  | 5.44e-05 | -70 | -94 | -117 | -141
5.46e-05 | -46 | -70 | -94 | -117 |  | 5.46e-05 | -46 | -70 | -94 | -117
5.48e-05 | -34 | -46 | -46 | -70 |  | 5.48e-05 | -34 | -46 | -70 | -94
5.5e-05 | -23 | -23 | -46 | -46 |  | 5.5e-05 | -23 | -46 | -46 | -46
5.52e-05 | -11 | -23 | -23 | -23 |  | 5.52e-05 | -11 | -23 | -23 | -23
5.54e-05 | 0 | 0 | 0 | 0 |  | 5.54e-05 | 0 | 0 | 0 | 0
5.56e-05 | 12 | 24 | 24 | 47 |  | 5.56e-05 | 12 | 24 | 24 | 24
5.58e-05 | 24 | 24 | 47 | 71 |  | 5.58e-05 | 24 | 24 | 47 | 47
5.6e-05 | 35 | 47 | 71 | 95 |  | 5.6e-05 | 35 | 47 | 71 | 95
5.62e-05 | 47 | 71 | 95 | 118 |  | 5.62e-05 | 47 | 71 | 95 | 118
5.64e-05 | 59 | 95 | 118 | 142 |  | 5.64e-05 | 59 | 71 | 118 | 142
5.66e-05 | 71 | 95 | 142 | 166 |  | 5.66e-05 | 59 | 95 | 142 | 166
5.68e-05 | 71 | 118 | 166 | 189 |  | 5.68e-05 | 71 | 118 | 166 | 189
5.7e-05 | 94 | 118 | 166 | 213 |  | 5.7e-05 | 94 | 118 | 166 | 213
5.72e-05 | 94 | 142 | 189 | 237 |  | 5.72e-05 | 94 | 142 | 189 | 237
5.74e-05 | 106 | 166 | 213 | 260 |  | 5.74e-05 | 106 | 166 | 213 | 260
5.76e-05 | 118 | 189 | 237 | 284 |  | 5.76e-05 | 118 | 166 | 237 | 284
5.78e-05 | 118 | 189 | 237 | 308 |  | 5.78e-05 | 118 | 189 | 237 | 308
5.8e-05 | 130 | 189 | 260 | 331 |  | 5.8e-05 | 130 | 189 | 260 | 331
5.82e-05 | 142 | 213 | 284 | 355 |  | 5.82e-05 | 130 | 213 | 260 | 331
5.84e-05 | 142 | 213 | 284 | 355 |  | 5.84e-05 | 142 | 213 | 284 | 355
5.86e-05 | 154 | 237 | 308 | 379 |  | 5.86e-05 | 154 | 237 | 308 | 379
5.88e-05 | 165 | 237 | 308 | 379 |  | 5.88e-05 | 154 | 237 | 308 | 379
5.9e-05 | 165 | 260 | 331 | 402 |  | 5.9e-05 | 165 | 237 | 331 | 402
5.92e-05 | 165 | 260 | 331 | 426 |  | 5.92e-05 | 177 | 237 | 331 | 426
5.94e-05 | 177 | 260 | 355 | 426 |  | 5.94e-05 | 177 | 260 | 355 | 426
5.96e-05 | 177 | 260 | 355 | 426 |  | 5.96e-05 | 177 | 260 | 355 | 450
5.98e-05 | 189 | 260 | 355 | 450 |  | 5.98e-05 | 177 | 260 | 355 | 450
6e-05 | 189 | 284 | 379 | 450 |  | 6e-05 | 189 | 284 | 379 | 450
6.02e-05 | 189 | 284 | 379 | 473 |  | 6.02e-05 | 189 | 284 | 379 | 473
6.04e-05 | 201 | 284 | 379 | 473 |  | 6.04e-05 | 201 | 284 | 379 | 473
6.06e-05 | 201 | 284 | 379 | 497 |  | 6.06e-05 | 201 | 284 | 379 | 497
6.08e-05 | 201 | 308 | 402 | 497 |  | 6.08e-05 | 201 | 308 | 402 | 497
6.1e-05 | 201 | 308 | 379 | 497 |  | 6.1e-05 | 201 | 308 | 402 | 497
6.12e-05 | 201 | 308 | 402 | 497 |  | 6.12e-05 | 201 | 308 | 402 | 497
6.14e-05 | 201 | 308 | 402 | 497 |  | 6.14e-05 | 201 | 308 | 402 | 497
6.16e-05 | 213 | 308 | 402 | 497 |  | 6.16e-05 | 201 | 308 | 402 | 497
6.18e-05 | 201 | 308 | 402 | 497 |  | 6.18e-05 | 201 | 308 | 402 | 497
6.2e-05 | 201 | 308 | 402 | 497 |  | 6.2e-05 | 201 | 308 | 402 | 497
6.22e-05 | 201 | 308 | 402 | 497 |  | 6.22e-05 | 201 | 308 | 402 | 497
6.24e-05 | 201 | 308 | 379 | 497 |  | 6.24e-05 | 201 | 308 | 402 | 497
6.26e-05 | 201 | 308 | 379 | 473 |  | 6.26e-05 | 201 | 308 | 379 | 497
6.28e-05 | 201 | 308 | 379 | 473 |  | 6.28e-05 | 201 | 284 | 379 | 473
6.3e-05 | 201 | 284 | 379 | 473 |  | 6.3e-05 | 201 | 284 | 379 | 473
6.32e-05 | 189 | 284 | 379 | 450 |  | 6.32e-05 | 201 | 284 | 379 | 473
6.34e-05 | 189 | 284 | 355 | 450 |  | 6.34e-05 | 189 | 284 | 355 | 450
6.36e-05 | 189 | 260 | 355 | 450 |  | 6.36e-05 | 189 | 284 | 355 | 450
6.38e-05 | 177 | 260 | 355 | 426 |  | 6.38e-05 | 177 | 260 | 355 | 450
6.4e-05 | 177 | 260 | 331 | 426 |  | 6.4e-05 | 177 | 260 | 355 | 426
6.42e-05 | 177 | 237 | 331 | 402 |  | 6.42e-05 | 177 | 260 | 331 | 426
6.44e-05 | 165 | 260 | 308 | 379 |  | 6.44e-05 | 165 | 260 | 331 | 402
6.46e-05 | 165 | 237 | 308 | 379 |  | 6.46e-05 | 165 | 237 | 308 | 379
6.48e-05 | 154 | 237 | 284 | 355 |  | 6.48e-05 | 154 | 237 | 308 | 379
6.5e-05 | 154 | 213 | 284 | 355 |  | 6.5e-05 | 154 | 237 | 284 | 355
6.52e-05 | 142 | 213 | 260 | 331 |  | 6.52e-05 | 142 | 213 | 260 | 331
6.54e-05 | 130 | 213 | 260 | 308 |  | 6.54e-05 | 142 | 213 | 260 | 331
6.56e-05 | 130 | 189 | 237 | 308 |  | 6.56e-05 | 130 | 189 | 260 | 308
6.58e-05 | 118 | 189 | 237 | 284 |  | 6.58e-05 | 118 | 189 | 237 | 284
6.6e-05 | 118 | 166 | 237 | 260 |  | 6.6e-05 | 118 | 166 | 237 | 284
6.62e-05 | 106 | 166 | 213 | 260 |  | 6.62e-05 | 106 | 166 | 213 | 260
6.64e-05 | 106 | 142 | 189 | 237 |  | 6.64e-05 | 106 | 166 | 189 | 237
6.66e-05 | 94 | 142 | 166 | 213 |  | 6.66e-05 | 94 | 142 | 189 | 213
6.68e-05 | 83 | 118 | 166 | 189 |  | 6.68e-05 | 83 | 118 | 166 | 213
6.7e-05 | 83 | 118 | 142 | 189 |  | 6.7e-05 | 83 | 118 | 142 | 189
6.72e-05 | 71 | 95 | 142 | 166 |  | 6.72e-05 | 71 | 118 | 142 | 166
6.74e-05 | 71 | 95 | 118 | 142 |  | 6.74e-05 | 59 | 95 | 118 | 142
6.76e-05 | 59 | 71 | 95 | 118 |  | 6.76e-05 | 59 | 71 | 118 | 142
6.78e-05 | 47 | 71 | 95 | 95 |  | 6.78e-05 | 47 | 71 | 95 | 118
6.8e-05 | 35 | 47 | 71 | 71 |  | 6.8e-05 | 35 | 47 | 71 | 95
6.82e-05 | 35 | 47 | 47 | 71 |  | 6.82e-05 | 35 | 47 | 47 | 71
6.84e-05 | 24 | 24 | 24 | 47 |  | 6.84e-05 | 24 | 24 | 47 | 47
6.86e-05 | 12 | 24 | 24 | 24 |  | 6.86e-05 | 12 | 24 | 24 | 47
6.88e-05 | 0 | 0 | 0 | 0 |  | 6.88e-05 | 12 | 0 | 0 | 0
6.9e-05 | 0 | 0 | -23 | -23 |  | 6.9e-05 | 0 | 0 | 0 | 0
6.92e-05 | -11 | -23 | -23 | -23 |  | 6.92e-05 | 0 | -23 | -23 | -23
6.94e-05 | -11 | -23 | -46 | -46 |  | 6.94e-05 | -11 | -23 | -46 | -46
6.96e-05 | -23 | -46 | -46 | -70 |  | 6.96e-05 | -23 | -46 | -46 | -70
6.98e-05 | -34 | -46 | -70 | -94 |  | 6.98e-05 | -34 | -46 | -70 | -94
7e-05 | -34 | -70 | -94 | -117 |  | 7e-05 | -34 | -46 | -70 | -94
7.02e-05 | -46 | -70 | -94 | -117 |  | 7.02e-05 | -46 | -70 | -94 | -117
7.04e-05 | -58 | -94 | -117 | -141 |  | 7.04e-05 | -46 | -70 | -117 | -141
7.06e-05 | -58 | -94 | -141 | -165 |  | 7.06e-05 | -58 | -94 | -117 | -165
7.08e-05 | -70 | -117 | -141 | -165 |  | 7.08e-05 | -70 | -94 | -141 | -165
7.1e-05 | -70 | -117 | -165 | -188 |  | 7.1e-05 | -70 | -117 | -141 | -188
7.12e-05 | -82 | -117 | -165 | -188 |  | 7.12e-05 | -82 | -117 | -165 | -188
7.14e-05 | -82 | -141 | -165 | -212 |  | 7.14e-05 | -82 | -141 | -165 | -212
7.16e-05 | -93 | -141 | -188 | -236 |  | 7.16e-05 | -93 | -141 | -188 | -236
7.18e-05 | -93 | -141 | -212 | -236 |  | 7.18e-05 | -93 | -141 | -188 | -236
7.2e-05 | -105 | -165 | -212 | -259 |  | 7.2e-05 | -105 | -165 | -212 | -236
7.22e-05 | -105 | -165 | -212 | -259 |  | 7.22e-05 | -105 | -165 | -212 | -259
7.24e-05 | -105 | -165 | -236 | -283 |  | 7.24e-05 | -117 | -165 | -236 | -283
7.26e-05 | -117 | -165 | -236 | -283 |  | 7.26e-05 | -117 | -188 | -236 | -283
7.28e-05 | -117 | -188 | -236 | -307 |  | 7.28e-05 | -117 | -188 | -236 | -283
7.3e-05 | -117 | -188 | -236 | -307 |  | 7.3e-05 | -117 | -188 | -259 | -307
7.32e-05 | -129 | -188 | -259 | -307 |  | 7.32e-05 | -129 | -188 | -259 | -307
7.34e-05 | -129 | -212 | -259 | -307 |  | 7.34e-05 | -129 | -188 | -259 | -307
7.36e-05 | -129 | -212 | -259 | -330 |  | 7.36e-05 | -129 | -212 | -259 | -330
7.38e-05 | -141 | -212 | -259 | -330 |  | 7.38e-05 | -129 | -212 | -259 | -330
7.4e-05 | -141 | -212 | -283 | -330 |  | 7.4e-05 | -141 | -212 | -283 | -330
7.42e-05 | -141 | -212 | -283 | -330 |  | 7.42e-05 | -141 | -212 | -283 | -330
7.44e-05 | -141 | -212 | -283 | -330 |  | 7.44e-05 | -141 | -212 | -283 | -330
7.46e-05 | -141 | -212 | -283 | -330 |  | 7.46e-05 | -141 | -212 | -283 | -354
7.48e-05 | -141 | -212 | -283 | -354 |  | 7.48e-05 | -141 | -212 | -283 | -354
7.5e-05 | -141 | -212 | -283 | -354 |  | 7.5e-05 | -141 | -212 | -283 | -354
7.52e-05 | -153 | -212 | -283 | -330 |  | 7.52e-05 | -141 | -212 | -283 | -354
7.54e-05 | -141 | -212 | -283 | -330 |  | 7.54e-05 | -141 | -212 | -283 | -354
7.56e-05 | -141 | -212 | -283 | -330 |  | 7.56e-05 | -141 | -236 | -283 | -354
7.58e-05 | -141 | -212 | -283 | -354 |  | 7.58e-05 | -141 | -212 | -283 | -354
7.6e-05 | -141 | -212 | -283 | -330 |  | 7.6e-05 | -141 | -236 | -283 | -354
7.62e-05 | -141 | -212 | -283 | -330 |  | 7.62e-05 | -141 | -212 | -283 | -330
7.64e-05 | -141 | -212 | -283 | -330 |  | 7.64e-05 | -141 | -212 | -283 | -330
7.66e-05 | -141 | -212 | -283 | -330 |  | 7.66e-05 | -141 | -212 | -283 | -330
7.68e-05 | -141 | -212 | -259 | -330 |  | 7.68e-05 | -141 | -212 | -283 | -330
7.7e-05 | -129 | -212 | -259 | -307 |  | 7.7e-05 | -129 | -212 | -259 | -330
7.72e-05 | -129 | -188 | -259 | -307 |  | 7.72e-05 | -129 | -212 | -259 | -307
7.74e-05 | -129 | -188 | -259 | -307 |  | 7.74e-05 | -129 | -188 | -259 | -307
7.76e-05 | -129 | -188 | -236 | -307 |  | 7.76e-05 | -129 | -188 | -259 | -307
7.78e-05 | -117 | -188 | -236 | -307 |  | 7.78e-05 | -117 | -188 | -236 | -283
7.8e-05 | -117 | -188 | -236 | -283 |  | 7.8e-05 | -117 | -188 | -236 | -283
7.82e-05 | -117 | -165 | -236 | -283 |  | 7.82e-05 | -117 | -188 | -236 | -283
7.84e-05 | -105 | -165 | -212 | -259 |  | 7.84e-05 | -105 | -165 | -212 | -283
7.86e-05 | -105 | -165 | -212 | -259 |  | 7.86e-05 | -105 | -165 | -212 | -259
7.88e-05 | -105 | -165 | -212 | -236 |  | 7.88e-05 | -105 | -165 | -212 | -259
7.9e-05 | -105 | -141 | -188 | -236 |  | 7.9e-05 | -93 | -141 | -188 | -236
7.92e-05 | -93 | -141 | -188 | -212 |  | 7.92e-05 | -93 | -141 | -188 | -236
7.94e-05 | -93 | -141 | -188 | -212 |  | 7.94e-05 | -93 | -141 | -165 | -212
7.96e-05 | -82 | -117 | -165 | -188 |  | 7.96e-05 | -82 | -117 | -165 | -212
7.98e-05 | -82 | -117 | -165 | -188 |  | 7.98e-05 | -82 | -117 | -165 | -188
8e-05 | -70 | -117 | -141 | -165 |  | 8e-05 | -82 | -117 | -141 | -188
8.02e-05 | -70 | -94 | -141 | -165 |  | 8.02e-05 | -70 | -117 | -141 | -165
8.04e-05 | -58 | -94 | -117 | -165 |  | 8.04e-05 | -70 | -94 | -117 | -141
8.06e-05 | -58 | -94 | -117 | -141 |  | 8.06e-05 | -58 | -94 | -117 | -141
8.08e-05 | -46 | -70 | -117 | -117 |  | 8.08e-05 | -58 | -94 | -117 | -141
8.1e-05 | -46 | -70 | -94 | -117 |  | 8.1e-05 | -46 | -70 | -94 | -117
8.12e-05 | -46 | -70 | -70 | -94 |  | 8.12e-05 | -46 | -70 | -94 | -94
8.14e-05 | -34 | -46 | -70 | -94 |  | 8.14e-05 | -34 | -70 | -70 | -94
8.16e-05 | -34 | -46 | -46 | -70 |  | 8.16e-05 | -34 | -46 | -70 | -70
8.18e-05 | -23 | -46 | -46 | -46 |  | 8.18e-05 | -23 | -46 | -70 | -70
8.2e-05 | -23 | -23 | -46 | -46 |  | 8.2e-05 | -23 | -23 | -46 | -46
8.22e-05 | -11 | -23 | -23 | -23 |  | 8.22e-05 | -11 | -23 | -23 | -46
8.24e-05 | -11 | -23 | -23 | -23 |  | 8.24e-05 | -11 | -23 | -23 | -23
8.26e-05 | 0 | 0 | 0 | 0 |  | 8.26e-05 | -11 | -23 | 0 | -23
8.28e-05 | 0 | 0 | 0 | 0 |  | 8.28e-05 | 0 | 0 | 0 | 0
8.3e-05 | 12 | 0 | 24 | 24 |  | 8.3e-05 | 12 | 0 | 0 | 24
8.32e-05 | 12 | 24 | 24 | 24 |  | 8.32e-05 | 12 | 24 | 24 | 24
8.34e-05 | 24 | 24 | 24 | 47 |  | 8.34e-05 | 12 | 24 | 24 | 47
8.36e-05 | 24 | 24 | 47 | 47 |  | 8.36e-05 | 24 | 24 | 47 | 47
8.38e-05 | 24 | 47 | 47 | 71 |  | 8.38e-05 | 24 | 47 | 47 | 71
8.4e-05 | 35 | 47 | 47 | 71 |  | 8.4e-05 | 35 | 47 | 71 | 71
8.42e-05 | 35 | 47 | 71 | 95 |  | 8.42e-05 | 35 | 47 | 71 | 95
8.44e-05 | 35 | 71 | 71 | 95 |  | 8.44e-05 | 35 | 47 | 71 | 95
8.46e-05 | 47 | 71 | 95 | 118 |  | 8.46e-05 | 47 | 71 | 71 | 118
8.48e-05 | 47 | 71 | 95 | 118 |  | 8.48e-05 | 47 | 71 | 95 | 118
8.5e-05 | 59 | 71 | 95 | 118 |  | 8.5e-05 | 59 | 71 | 95 | 118
8.52e-05 | 59 | 71 | 118 | 142 |  | 8.52e-05 | 59 | 71 | 118 | 142
8.54e-05 | 59 | 95 | 118 | 142 |  | 8.54e-05 | 59 | 95 | 118 | 142
8.56e-05 | 71 | 95 | 118 | 166 |  | 8.56e-05 | 59 | 95 | 118 | 142
8.58e-05 | 71 | 95 | 118 | 166 |  | 8.58e-05 | 71 | 95 | 118 | 166
8.6e-05 | 71 | 95 | 142 | 166 |  | 8.6e-05 | 71 | 118 | 142 | 166
8.62e-05 | 71 | 118 | 142 | 189 |  | 8.62e-05 | 71 | 95 | 142 | 189
8.64e-05 | 71 | 118 | 166 | 189 |  | 8.64e-05 | 83 | 118 | 142 | 189
8.66e-05 | 83 | 118 | 166 | 189 |  | 8.66e-05 | 71 | 118 | 142 | 189
8.68e-05 | 83 | 118 | 166 | 189 |  | 8.68e-05 | 83 | 118 | 142 | 189
8.7e-05 | 83 | 118 | 166 | 189 |  | 8.7e-05 | 83 | 118 | 166 | 213
8.72e-05 | 83 | 118 | 166 | 213 |  | 8.72e-05 | 83 | 118 | 166 | 213
8.74e-05 | 83 | 142 | 166 | 213 |  | 8.74e-05 | 83 | 142 | 166 | 213
8.76e-05 | 94 | 142 | 166 | 213 |  | 8.76e-05 | 94 | 142 | 166 | 213
8.78e-05 | 94 | 142 | 189 | 213 |  | 8.78e-05 | 94 | 142 | 166 | 213
8.8e-05 | 94 | 142 | 189 | 213 |  | 8.8e-05 | 94 | 142 | 189 | 213
8.82e-05 | 94 | 142 | 189 | 213 |  | 8.82e-05 | 94 | 142 | 189 | 213
8.84e-05 | 94 | 142 | 189 | 237 |  | 8.84e-05 | 94 | 142 | 189 | 213
8.86e-05 | 94 | 142 | 189 | 213 |  | 8.86e-05 | 94 | 142 | 189 | 213
8.88e-05 | 94 | 142 | 189 | 213 |  | 8.88e-05 | 94 | 142 | 189 | 213
8.9e-05 | 94 | 142 | 189 | 213 |  | 8.9e-05 | 94 | 142 | 189 | 237
8.92e-05 | 94 | 142 | 189 | 213 |  | 8.92e-05 | 94 | 142 | 189 | 213
8.94e-05 | 94 | 142 | 189 | 237 |  | 8.94e-05 | 94 | 142 | 189 | 213
8.96e-05 | 94 | 142 | 189 | 213 |  | 8.96e-05 | 94 | 142 | 189 | 213
8.98e-05 | 94 | 142 | 166 | 213 |  | 8.98e-05 | 94 | 142 | 166 | 213
9e-05 | 94 | 142 | 166 | 213 |  | 9e-05 | 94 | 142 | 166 | 213
9.02e-05 | 94 | 142 | 166 | 213 |  | 9.02e-05 | 94 | 142 | 166 | 213
9.04e-05 | 94 | 118 | 166 | 213 |  | 9.04e-05 | 83 | 142 | 166 | 213
9.06e-05 | 83 | 118 | 166 | 213 |  | 9.06e-05 | 83 | 118 | 166 | 213
9.08e-05 | 83 | 142 | 166 | 213 |  | 9.08e-05 | 83 | 142 | 166 | 213
9.1e-05 | 83 | 118 | 166 | 189 |  | 9.1e-05 | 83 | 118 | 166 | 213
9.12e-05 | 83 | 118 | 166 | 189 |  | 9.12e-05 | 83 | 118 | 166 | 213
9.14e-05 | 83 | 118 | 166 | 189 |  | 9.14e-05 | 83 | 118 | 166 | 189
9.16e-05 | 83 | 118 | 142 | 189 |  | 9.16e-05 | 83 | 118 | 142 | 189
9.18e-05 | 71 | 118 | 142 | 166 |  | 9.18e-05 | 83 | 118 | 142 | 189
9.2e-05 | 71 | 95 | 142 | 166 |  | 9.2e-05 | 83 | 95 | 142 | 189
9.22e-05 | 71 | 95 | 142 | 166 |  | 9.22e-05 | 71 | 118 | 142 | 166
9.24e-05 | 71 | 95 | 118 | 166 |  | 9.24e-05 | 71 | 95 | 142 | 166
9.26e-05 | 59 | 95 | 118 | 142 |  | 9.26e-05 | 71 | 95 | 118 | 166
9.28e-05 | 59 | 95 | 118 | 142 |  | 9.28e-05 | 59 | 95 | 118 | 142
9.3e-05 | 59 | 71 | 118 | 142 |  | 9.3e-05 | 59 | 95 | 118 | 142
9.32e-05 | 59 | 71 | 118 | 118 |  | 9.32e-05 | 59 | 71 | 95 | 118
9.34e-05 | 59 | 71 | 95 | 118 |  | 9.34e-05 | 47 | 71 | 95 | 118
9.36e-05 | 47 | 71 | 95 | 118 |  | 9.36e-05 | 59 | 71 | 95 | 118
9.38e-05 | 47 | 71 | 71 | 95 |  | 9.38e-05 | 47 | 71 | 95 | 118
9.4e-05 | 47 | 47 | 71 | 95 |  | 9.4e-05 | 47 | 71 | 71 | 95
9.42e-05 | 35 | 47 | 71 | 95 |  | 9.42e-05 | 35 | 47 | 71 | 95
9.44e-05 | 35 | 47 | 71 | 71 |  | 9.44e-05 | 35 | 47 | 71 | 71
9.46e-05 | 35 | 47 | 47 | 71 |  | 9.46e-05 | 35 | 47 | 47 | 71
9.48e-05 | 24 | 47 | 47 | 47 |  | 9.48e-05 | 24 | 47 | 47 | 71
9.5e-05 | 24 | 24 | 47 | 47 |  | 9.5e-05 | 24 | 47 | 47 | 71
9.52e-05 | 24 | 24 | 47 | 47 |  | 9.52e-05 | 24 | 24 | 47 | 47
9.54e-05 | 12 | 24 | 24 | 47 |  | 9.54e-05 | 24 | 24 | 24 | 47
9.56e-05 | 12 | 24 | 24 | 24 |  | 9.56e-05 | 12 | 24 | 24 | 24
9.58e-05 | 12 | 24 | 0 | 24 |  | 9.58e-05 | 12 | 0 | 0 | 24
9.6e-05 | 12 | 0 | 0 | 0 |  | 9.6e-05 | 12 | 0 | 24 | 24
9.62e-05 | 0 | 0 | 0 | 0 |  | 9.62e-05 | 0 | 0 | 0 | 0
9.64e-05 | 0 | 0 | 0 | -23 |  | 9.64e-05 | 0 | 0 | 0 | 0
9.66e-05 | -11 | 0 | -23 | -23 |  | 9.66e-05 | 0 | -23 | -23 | -23
9.68e-05 | -11 | -23 | -23 | -23 |  | 9.68e-05 | -11 | -23 | -23 | -23
9.7e-05 | -11 | -23 | -23 | -46 |  | 9.7e-05 | -11 | -23 | -23 | -23
9.72e-05 | -11 | -23 | -46 | -46 |  | 9.72e-05 | -11 | -23 | -23 | -46
9.74e-05 | -23 | -46 | -46 | -46 |  | 9.74e-05 | -23 | -23 | -23 | -46
9.76e-05 | -23 | -46 | -46 | -46 |  | 9.76e-05 | -23 | -46 | -46 | -70
9.78e-05 | -23 | -46 | -46 | -70 |  | 9.78e-05 | -23 | -46 | -46 | -70
9.8e-05 | -23 | -46 | -70 | -70 |  | 9.8e-05 | -23 | -46 | -46 | -70
9.82e-05 | -23 | -46 | -70 | -70 |  | 9.82e-05 | -23 | -46 | -70 | -70
9.84e-05 | -34 | -46 | -70 | -94 |  | 9.84e-05 | -34 | -46 | -70 | -94
9.86e-05 | -34 | -70 | -70 | -94 |  | 9.86e-05 | -34 | -70 | -70 | -94
9.88e-05 | -34 | -70 | -94 | -94 |  | 9.88e-05 | -34 | -70 | -94 | -94
9.9e-05 | -46 | -70 | -94 | -117 |  | 9.9e-05 | -46 | -70 | -94 | -94
9.92e-05 | -46 | -70 | -94 | -117 |  | 9.92e-05 | -46 | -70 | -94 | -117
9.94e-05 | -46 | -70 | -94 | -117 |  | 9.94e-05 | -46 | -70 | -94 | -117
9.96e-05 | -46 | -70 | -94 | -117 |  | 9.96e-05 | -46 | -70 | -94 | -117
9.98e-05 | -58 | -70 | -117 | -117 |  | 9.98e-05 | -46 | -70 | -117 | -117
0.0001 | -46 | -94 | -117 | -141 |  | 0.0001 | -46 | -94 | -117 | -141
0.0001002 | -46 | -94 | -117 | -141 |  | 0.0001002 | -58 | -94 | -117 | -141
0.0001004 | -58 | -94 | -117 | -141 |  | 0.0001004 | -58 | -94 | -117 | -141
0.0001006 | -58 | -94 | -117 | -141 |  | 0.0001006 | -58 | -94 | -117 | -141
0.0001008 | -58 | -94 | -117 | -141 |  | 0.0001008 | -58 | -94 | -117 | -141
0.000101 | -58 | -94 | -117 | -141 |  | 0.000101 | -58 | -94 | -117 | -141
0.0001012 | -58 | -94 | -117 | -141 |  | 0.0001012 | -58 | -94 | -117 | -141
0.0001014 | -70 | -94 | -117 | -165 |  | 0.0001014 | -70 | -94 | -117 | -165
0.0001016 | -70 | -94 | -117 | -165 |  | 0.0001016 | -70 | -94 | -117 | -165
0.0001018 | -70 | -94 | -117 | -165 |  | 0.0001018 | -58 | -94 | -141 | -165
0.000102 | -70 | -94 | -141 | -165 |  | 0.000102 | -70 | -94 | -141 | -165
0.0001022 | -70 | -94 | -141 | -165 |  | 0.0001022 | -58 | -94 | -141 | -165
| 2 | 2 | 2 | 2 |  |  | 2 | 2 | 2 | 2
| 1000 | 1500 | 2000 | 2500 |  |  | 1000 | 1500 | 2000 | 2500
| 2 | 2 | 2 | 2 |  |  | 2 | 2 | 2 | 2

## Step 2 2000 V Combined
Time (s) | Coil 158 | Coil 159 | Coil 224 | Coil 225 | Coil 150 | Coil --- | Time (s) | Coil 158 | Coil 159 | Coil 224 | Coil 225 | Coil 150 | Coil --- | Notes
0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 0 | 0 | 0 | 0
4e-07 | 0 | 0 | 0 | 0 | 0 |  | 4e-07 | 0 | 0 | 0 | 0 | 0
8e-07 | 0 | 0 | 0 | 0 | 0 |  | 8e-07 | 0 | 0 | 0 | 0 | 0
1.2e-06 | 450 | 426 | 473 | 521 | 450 |  | 1.2e-06 | 426 | 450 | 473 | 521 | 426
1.6e-06 | 1514 | 1514 | 1372 | 1372 | 1514 |  | 1.6e-06 | 1514 | 1514 | 1372 | 1372 | 1538
2e-06 | 2011 | 2011 | 1964 | 1964 | 2011 |  | 2e-06 | 2011 | 2011 | 1964 | 1964 | 2011
2.4e-06 | 1775 | 1751 | 1869 | 1869 | 1751 |  | 2.4e-06 | 1775 | 1775 | 1869 | 1869 | 1775
2.8e-06 | 1633 | 1633 | 1656 | 1656 | 1633 |  | 2.8e-06 | 1633 | 1633 | 1656 | 1656 | 1633
3.2e-06 | 1680 | 1680 | 1633 | 1609 | 1656 |  | 3.2e-06 | 1656 | 1680 | 1609 | 1609 | 1656
3.6e-06 | 1656 | 1656 | 1656 | 1656 | 1633 |  | 3.6e-06 | 1656 | 1656 | 1656 | 1656 | 1633
4e-06 | 1609 | 1609 | 1633 | 1609 | 1585 |  | 4e-06 | 1585 | 1609 | 1609 | 1609 | 1585
4.4e-06 | 1585 | 1585 | 1562 | 1562 | 1562 |  | 4.4e-06 | 1562 | 1585 | 1562 | 1562 | 1562
4.8e-06 | 1562 | 1562 | 1538 | 1538 | 1562 |  | 4.8e-06 | 1562 | 1562 | 1538 | 1538 | 1562
5.2e-06 | 1538 | 1538 | 1514 | 1538 | 1538 |  | 5.2e-06 | 1538 | 1538 | 1514 | 1538 | 1538
5.6e-06 | 1514 | 1514 | 1491 | 1491 | 1491 |  | 5.6e-06 | 1514 | 1514 | 1491 | 1514 | 1514
6e-06 | 1491 | 1491 | 1467 | 1467 | 1467 |  | 6e-06 | 1491 | 1491 | 1467 | 1467 | 1467
6.4e-06 | 1443 | 1443 | 1443 | 1443 | 1443 |  | 6.4e-06 | 1443 | 1443 | 1443 | 1443 | 1443
6.8e-06 | 1420 | 1396 | 1396 | 1396 | 1396 |  | 6.8e-06 | 1396 | 1396 | 1396 | 1396 | 1396
7.2e-06 | 1372 | 1372 | 1372 | 1372 | 1349 |  | 7.2e-06 | 1372 | 1372 | 1372 | 1372 | 1372
7.6e-06 | 1325 | 1349 | 1349 | 1349 | 1325 |  | 7.6e-06 | 1325 | 1349 | 1325 | 1349 | 1325
8e-06 | 1325 | 1301 | 1301 | 1301 | 1301 |  | 8e-06 | 1301 | 1301 | 1301 | 1301 | 1301
8.4e-06 | 1278 | 1278 | 1278 | 1278 | 1278 |  | 8.4e-06 | 1301 | 1278 | 1278 | 1278 | 1278
8.8e-06 | 1278 | 1254 | 1254 | 1254 | 1254 |  | 8.8e-06 | 1254 | 1254 | 1254 | 1254 | 1254
9.2e-06 | 1254 | 1230 | 1230 | 1230 | 1230 |  | 9.2e-06 | 1254 | 1254 | 1230 | 1230 | 1230
9.6e-06 | 1207 | 1207 | 1207 | 1207 | 1207 |  | 9.6e-06 | 1207 | 1207 | 1207 | 1207 | 1207
1e-05 | 1207 | 1183 | 1183 | 1183 | 1183 |  | 1e-05 | 1183 | 1183 | 1183 | 1183 | 1183
1.04e-05 | 1159 | 1159 | 1159 | 1159 | 1159 |  | 1.04e-05 | 1159 | 1159 | 1159 | 1159 | 1159
1.08e-05 | 1136 | 1136 | 1136 | 1136 | 1136 |  | 1.08e-05 | 1136 | 1136 | 1136 | 1136 | 1136
1.12e-05 | 1112 | 1112 | 1112 | 1112 | 1112 |  | 1.12e-05 | 1112 | 1112 | 1112 | 1112 | 1112
1.16e-05 | 1088 | 1065 | 1088 | 1088 | 1065 |  | 1.16e-05 | 1088 | 1088 | 1088 | 1088 | 1088
1.2e-05 | 1041 | 1041 | 1065 | 1065 | 1041 |  | 1.2e-05 | 1041 | 1041 | 1065 | 1065 | 1041
1.24e-05 | 1017 | 1017 | 1041 | 1017 | 1017 |  | 1.24e-05 | 1017 | 1017 | 1041 | 1017 | 1017
1.28e-05 | 970 | 970 | 994 | 994 | 970 |  | 1.28e-05 | 970 | 970 | 994 | 994 | 970
1.32e-05 | 946 | 946 | 970 | 970 | 946 |  | 1.32e-05 | 923 | 946 | 970 | 970 | 946
1.36e-05 | 899 | 899 | 946 | 923 | 899 |  | 1.36e-05 | 899 | 899 | 946 | 923 | 899
1.4e-05 | 852 | 852 | 899 | 875 | 875 |  | 1.4e-05 | 852 | 852 | 899 | 875 | 875
1.44e-05 | 804 | 828 | 875 | 852 | 828 |  | 1.44e-05 | 804 | 828 | 875 | 852 | 828
1.48e-05 | 781 | 781 | 828 | 804 | 781 |  | 1.48e-05 | 757 | 781 | 828 | 804 | 781
1.52e-05 | 734 | 734 | 781 | 757 | 734 |  | 1.52e-05 | 710 | 734 | 781 | 757 | 734
1.56e-05 | 686 | 686 | 734 | 734 | 686 |  | 1.56e-05 | 686 | 686 | 734 | 734 | 686
1.6e-05 | 639 | 639 | 686 | 686 | 639 |  | 1.6e-05 | 639 | 639 | 710 | 686 | 639
1.64e-05 | 592 | 592 | 663 | 639 | 592 |  | 1.64e-05 | 592 | 592 | 663 | 639 | 592
1.68e-05 | 544 | 544 | 615 | 592 | 544 |  | 1.68e-05 | 521 | 544 | 615 | 592 | 544
1.72e-05 | 497 | 497 | 568 | 544 | 497 |  | 1.72e-05 | 497 | 497 | 568 | 544 | 497
1.76e-05 | 426 | 450 | 521 | 497 | 450 |  | 1.76e-05 | 426 | 450 | 521 | 497 | 450
1.8e-05 | 379 | 402 | 473 | 450 | 402 |  | 1.8e-05 | 379 | 402 | 473 | 450 | 402
1.84e-05 | 331 | 331 | 426 | 402 | 355 |  | 1.84e-05 | 331 | 331 | 426 | 402 | 355
1.88e-05 | 284 | 308 | 379 | 355 | 308 |  | 1.88e-05 | 284 | 308 | 379 | 355 | 308
1.92e-05 | 237 | 260 | 331 | 308 | 260 |  | 1.92e-05 | 237 | 237 | 331 | 308 | 260
1.96e-05 | 189 | 213 | 284 | 260 | 213 |  | 1.96e-05 | 189 | 213 | 284 | 260 | 213
2e-05 | 166 | 166 | 237 | 213 | 166 |  | 2e-05 | 142 | 166 | 237 | 213 | 166
2.04e-05 | 118 | 95 | 189 | 166 | 118 |  | 2.04e-05 | 118 | 118 | 189 | 166 | 118
2.08e-05 | 47 | 71 | 142 | 118 | 71 |  | 2.08e-05 | 47 | 71 | 142 | 118 | 71
2.12e-05 | 0 | 0 | 118 | 71 | 24 |  | 2.12e-05 | 24 | 24 | 118 | 71 | 24
2.16e-05 | -46 | -23 | 71 | 24 | -23 |  | 2.16e-05 | -23 | -23 | 71 | 47 | -23
2.2e-05 | -70 | -70 | 24 | 0 | -70 |  | 2.2e-05 | -70 | -70 | 24 | 0 | -70
2.24e-05 | -117 | -117 | -23 | -46 | -117 |  | 2.24e-05 | -117 | -117 | -23 | -46 | -94
2.28e-05 | -165 | -165 | -70 | -94 | -141 |  | 2.28e-05 | -165 | -165 | -70 | -94 | -165
2.32e-05 | -212 | -212 | -117 | -141 | -188 |  | 2.32e-05 | -212 | -212 | -117 | -141 | -188
2.36e-05 | -259 | -259 | -165 | -188 | -236 |  | 2.36e-05 | -259 | -259 | -165 | -188 | -236
2.4e-05 | -283 | -283 | -188 | -212 | -283 |  | 2.4e-05 | -283 | -307 | -188 | -212 | -283
2.44e-05 | -330 | -330 | -236 | -259 | -307 |  | 2.44e-05 | -330 | -330 | -236 | -259 | -307
2.48e-05 | -378 | -378 | -283 | -307 | -354 |  | 2.48e-05 | -378 | -378 | -259 | -307 | -354
2.52e-05 | -425 | -425 | -307 | -354 | -401 |  | 2.52e-05 | -425 | -425 | -307 | -354 | -401
2.56e-05 | -472 | -449 | -354 | -378 | -449 |  | 2.56e-05 | -449 | -449 | -354 | -378 | -449
2.6e-05 | -496 | -496 | -378 | -425 | -496 |  | 2.6e-05 | -496 | -496 | -401 | -425 | -472
2.64e-05 | -543 | -543 | -425 | -449 | -520 |  | 2.64e-05 | -543 | -543 | -425 | -449 | -520
2.68e-05 | -591 | -567 | -472 | -496 | -567 |  | 2.68e-05 | -567 | -567 | -472 | -496 | -567
2.72e-05 | -614 | -614 | -496 | -543 | -591 |  | 2.72e-05 | -614 | -614 | -496 | -543 | -591
2.76e-05 | -662 | -662 | -543 | -567 | -638 |  | 2.76e-05 | -662 | -638 | -543 | -567 | -638
2.8e-05 | -685 | -685 | -567 | -614 | -662 |  | 2.8e-05 | -685 | -685 | -591 | -614 | -662
2.84e-05 | -733 | -733 | -614 | -638 | -709 |  | 2.84e-05 | -733 | -733 | -614 | -638 | -709
2.88e-05 | -756 | -756 | -638 | -685 | -733 |  | 2.88e-05 | -756 | -756 | -638 | -685 | -733
2.92e-05 | -803 | -803 | -685 | -709 | -780 |  | 2.92e-05 | -803 | -803 | -685 | -709 | -780
2.96e-05 | -827 | -827 | -709 | -756 | -827 |  | 2.96e-05 | -827 | -827 | -709 | -756 | -803
3e-05 | -874 | -874 | -756 | -780 | -851 |  | 3e-05 | -874 | -874 | -756 | -780 | -851
3.04e-05 | -898 | -898 | -780 | -827 | -874 |  | 3.04e-05 | -898 | -898 | -780 | -827 | -874
3.08e-05 | -945 | -922 | -827 | -851 | -922 |  | 3.08e-05 | -945 | -922 | -827 | -851 | -922
3.12e-05 | -969 | -969 | -851 | -874 | -945 |  | 3.12e-05 | -969 | -969 | -851 | -874 | -945
3.16e-05 | -993 | -993 | -874 | -922 | -969 |  | 3.16e-05 | -993 | -993 | -874 | -922 | -969
3.2e-05 | -1016 | -1016 | -898 | -945 | -993 |  | 3.2e-05 | -1016 | -1016 | -922 | -945 | -993
3.24e-05 | -1040 | -1040 | -945 | -969 | -1016 |  | 3.24e-05 | -1040 | -1040 | -945 | -969 | -1016
3.28e-05 | -1087 | -1064 | -969 | -1016 | -1064 |  | 3.28e-05 | -1064 | -1064 | -969 | -993 | -1064
3.32e-05 | -1087 | -1111 | -993 | -1016 | -1087 |  | 3.32e-05 | -1087 | -1087 | -993 | -1016 | -1087
3.36e-05 | -1135 | -1135 | -1016 | -1064 | -1111 |  | 3.36e-05 | -1111 | -1111 | -1016 | -1064 | -1111
3.4e-05 | -1135 | -1135 | -1040 | -1087 | -1135 |  | 3.4e-05 | -1135 | -1135 | -1040 | -1087 | -1135
3.44e-05 | -1158 | -1158 | -1064 | -1111 | -1158 |  | 3.44e-05 | -1158 | -1158 | -1064 | -1111 | -1158
3.48e-05 | -1182 | -1182 | -1087 | -1135 | -1182 |  | 3.48e-05 | -1182 | -1182 | -1087 | -1135 | -1182
3.52e-05 | -1206 | -1206 | -1111 | -1135 | -1182 |  | 3.52e-05 | -1206 | -1206 | -1111 | -1158 | -1182
3.56e-05 | -1206 | -1206 | -1135 | -1158 | -1206 |  | 3.56e-05 | -1206 | -1229 | -1135 | -1158 | -1206
3.6e-05 | -1253 | -1253 | -1158 | -1182 | -1229 |  | 3.6e-05 | -1229 | -1253 | -1158 | -1182 | -1229
3.64e-05 | -1253 | -1253 | -1182 | -1206 | -1253 |  | 3.64e-05 | -1253 | -1253 | -1182 | -1206 | -1253
3.68e-05 | -1277 | -1277 | -1182 | -1206 | -1253 |  | 3.68e-05 | -1277 | -1277 | -1182 | -1229 | -1277
3.72e-05 | -1277 | -1300 | -1206 | -1229 | -1277 |  | 3.72e-05 | -1277 | -1277 | -1206 | -1229 | -1277
3.76e-05 | -1300 | -1300 | -1229 | -1253 | -1277 |  | 3.76e-05 | -1300 | -1300 | -1229 | -1253 | -1277
3.8e-05 | -1300 | -1300 | -1253 | -1253 | -1300 |  | 3.8e-05 | -1300 | -1300 | -1253 | -1253 | -1300
3.84e-05 | -1324 | -1324 | -1253 | -1277 | -1300 |  | 3.84e-05 | -1324 | -1324 | -1253 | -1277 | -1300
3.88e-05 | -1324 | -1324 | -1253 | -1277 | -1300 |  | 3.88e-05 | -1324 | -1324 | -1277 | -1277 | -1324
3.92e-05 | -1324 | -1324 | -1277 | -1300 | -1324 |  | 3.92e-05 | -1324 | -1324 | -1277 | -1300 | -1324
3.96e-05 | -1324 | -1324 | -1277 | -1300 | -1324 |  | 3.96e-05 | -1324 | -1324 | -1277 | -1300 | -1324
4e-05 | -1324 | -1348 | -1300 | -1300 | -1324 |  | 4e-05 | -1324 | -1348 | -1300 | -1324 | -1324
4.04e-05 | -1348 | -1348 | -1300 | -1324 | -1324 |  | 4.04e-05 | -1348 | -1348 | -1300 | -1324 | -1324
4.08e-05 | -1324 | -1348 | -1300 | -1324 | -1324 |  | 4.08e-05 | -1348 | -1348 | -1300 | -1324 | -1348
4.12e-05 | -1324 | -1348 | -1300 | -1324 | -1324 |  | 4.12e-05 | -1348 | -1348 | -1324 | -1324 | -1348
4.16e-05 | -1324 | -1324 | -1324 | -1324 | -1348 |  | 4.16e-05 | -1324 | -1348 | -1324 | -1324 | -1348
4.2e-05 | -1324 | -1324 | -1324 | -1324 | -1324 |  | 4.2e-05 | -1324 | -1324 | -1324 | -1324 | -1348
4.24e-05 | -1324 | -1324 | -1324 | -1324 | -1324 |  | 4.24e-05 | -1324 | -1324 | -1324 | -1324 | -1324
4.28e-05 | -1300 | -1324 | -1324 | -1324 | -1324 |  | 4.28e-05 | -1300 | -1324 | -1324 | -1324 | -1324
4.32e-05 | -1277 | -1300 | -1324 | -1324 | -1300 |  | 4.32e-05 | -1300 | -1300 | -1324 | -1324 | -1300
4.36e-05 | -1277 | -1277 | -1300 | -1300 | -1277 |  | 4.36e-05 | -1277 | -1277 | -1300 | -1324 | -1300
4.4e-05 | -1253 | -1253 | -1300 | -1300 | -1277 |  | 4.4e-05 | -1253 | -1253 | -1300 | -1300 | -1277
4.44e-05 | -1206 | -1229 | -1300 | -1277 | -1253 |  | 4.44e-05 | -1229 | -1229 | -1300 | -1300 | -1253
4.48e-05 | -1182 | -1206 | -1277 | -1277 | -1229 |  | 4.48e-05 | -1206 | -1206 | -1300 | -1277 | -1229
4.52e-05 | -1158 | -1182 | -1277 | -1253 | -1182 |  | 4.52e-05 | -1158 | -1182 | -1277 | -1253 | -1206
4.56e-05 | -1135 | -1158 | -1253 | -1229 | -1158 |  | 4.56e-05 | -1135 | -1158 | -1253 | -1229 | -1158
4.6e-05 | -1111 | -1111 | -1229 | -1206 | -1135 |  | 4.6e-05 | -1111 | -1135 | -1253 | -1206 | -1135
4.64e-05 | -1064 | -1087 | -1206 | -1182 | -1111 |  | 4.64e-05 | -1087 | -1087 | -1206 | -1182 | -1111
4.68e-05 | -1040 | -1040 | -1182 | -1158 | -1087 |  | 4.68e-05 | -1040 | -1064 | -1182 | -1158 | -1087
4.72e-05 | -993 | -1016 | -1182 | -1135 | -1040 |  | 4.72e-05 | -993 | -1016 | -1158 | -1135 | -1040
4.76e-05 | -969 | -969 | -1135 | -1111 | -1016 |  | 4.76e-05 | -969 | -993 | -1135 | -1111 | -1016
4.8e-05 | -922 | -945 | -1111 | -1064 | -969 |  | 4.8e-05 | -922 | -945 | -1135 | -1087 | -969
4.84e-05 | -874 | -898 | -1087 | -1040 | -922 |  | 4.84e-05 | -898 | -898 | -1111 | -1040 | -945
4.88e-05 | -827 | -851 | -1064 | -1016 | -898 |  | 4.88e-05 | -827 | -874 | -1064 | -1016 | -898
4.92e-05 | -780 | -803 | -1040 | -969 | -851 |  | 4.92e-05 | -803 | -803 | -1040 | -993 | -874
4.96e-05 | -733 | -756 | -1016 | -945 | -803 |  | 4.96e-05 | -756 | -780 | -1016 | -945 | -803
5e-05 | -709 | -709 | -969 | -898 | -756 |  | 5e-05 | -709 | -709 | -969 | -898 | -756
5.04e-05 | -638 | -662 | -945 | -874 | -709 |  | 5.04e-05 | -662 | -685 | -945 | -874 | -733
5.08e-05 | -591 | -614 | -898 | -827 | -662 |  | 5.08e-05 | -614 | -614 | -922 | -827 | -685
5.12e-05 | -543 | -567 | -874 | -780 | -614 |  | 5.12e-05 | -567 | -567 | -874 | -803 | -638
5.16e-05 | -496 | -520 | -827 | -756 | -567 |  | 5.16e-05 | -520 | -520 | -827 | -756 | -567
5.2e-05 | -449 | -472 | -780 | -709 | -520 |  | 5.2e-05 | -449 | -472 | -780 | -709 | -520
5.24e-05 | -401 | -425 | -756 | -662 | -472 |  | 5.24e-05 | -401 | -425 | -756 | -662 | -472
5.28e-05 | -330 | -354 | -709 | -614 | -425 |  | 5.28e-05 | -354 | -378 | -709 | -614 | -425
5.32e-05 | -307 | -307 | -662 | -567 | -378 |  | 5.32e-05 | -307 | -330 | -662 | -567 | -378
5.36e-05 | -236 | -259 | -614 | -520 | -330 |  | 5.36e-05 | -259 | -283 | -614 | -543 | -330
5.4e-05 | -188 | -212 | -591 | -472 | -283 |  | 5.4e-05 | -188 | -212 | -591 | -496 | -283
5.44e-05 | -141 | -165 | -543 | -425 | -212 |  | 5.44e-05 | -141 | -165 | -543 | -449 | -236
5.48e-05 | -94 | -117 | -496 | -378 | -165 |  | 5.48e-05 | -94 | -117 | -496 | -401 | -188
5.52e-05 | -46 | -70 | -449 | -330 | -117 |  | 5.52e-05 | -46 | -70 | -449 | -354 | -141
5.56e-05 | 24 | 0 | -401 | -307 | -70 |  | 5.56e-05 | 0 | -23 | -401 | -307 | -70
5.6e-05 | 71 | 47 | -354 | -236 | -23 |  | 5.6e-05 | 47 | 24 | -354 | -259 | -23
5.64e-05 | 118 | 95 | -330 | -188 | 24 |  | 5.64e-05 | 95 | 95 | -307 | -212 | 24
5.68e-05 | 166 | 142 | -283 | -165 | 71 |  | 5.68e-05 | 142 | 142 | -283 | -165 | 71
5.72e-05 | 213 | 189 | -236 | -117 | 118 |  | 5.72e-05 | 213 | 189 | -236 | -117 | 118
5.76e-05 | 260 | 237 | -188 | -70 | 166 |  | 5.76e-05 | 237 | 237 | -188 | -70 | 166
5.8e-05 | 284 | 284 | -141 | 0 | 213 |  | 5.8e-05 | 284 | 260 | -141 | -23 | 213
5.84e-05 | 331 | 331 | -94 | 47 | 260 |  | 5.84e-05 | 331 | 331 | -94 | 24 | 260
5.88e-05 | 379 | 379 | -46 | 71 | 308 |  | 5.88e-05 | 355 | 355 | -46 | 71 | 308
5.92e-05 | 426 | 402 | 0 | 142 | 355 |  | 5.92e-05 | 426 | 402 | 0 | 118 | 355
5.96e-05 | 473 | 450 | 47 | 166 | 379 |  | 5.96e-05 | 450 | 450 | 47 | 166 | 379
6e-05 | 497 | 497 | 95 | 213 | 426 |  | 6e-05 | 497 | 497 | 95 | 213 | 426
6.04e-05 | 544 | 521 | 118 | 260 | 473 |  | 6.04e-05 | 521 | 521 | 142 | 237 | 473
6.08e-05 | 568 | 568 | 166 | 284 | 521 |  | 6.08e-05 | 568 | 544 | 166 | 284 | 497
6.12e-05 | 592 | 592 | 213 | 331 | 544 |  | 6.12e-05 | 592 | 592 | 213 | 331 | 544
6.16e-05 | 639 | 615 | 237 | 379 | 568 |  | 6.16e-05 | 615 | 615 | 237 | 355 | 568
6.2e-05 | 663 | 663 | 284 | 402 | 615 |  | 6.2e-05 | 663 | 639 | 284 | 402 | 615
6.24e-05 | 686 | 686 | 331 | 450 | 639 |  | 6.24e-05 | 686 | 686 | 331 | 426 | 639
6.28e-05 | 710 | 710 | 355 | 473 | 663 |  | 6.28e-05 | 710 | 710 | 379 | 473 | 663
6.32e-05 | 734 | 734 | 402 | 521 | 686 |  | 6.32e-05 | 734 | 734 | 402 | 521 | 686
6.36e-05 | 757 | 757 | 426 | 544 | 710 |  | 6.36e-05 | 757 | 757 | 426 | 544 | 710
6.4e-05 | 781 | 781 | 473 | 568 | 734 |  | 6.4e-05 | 757 | 781 | 473 | 568 | 734
6.44e-05 | 804 | 804 | 497 | 615 | 757 |  | 6.44e-05 | 781 | 804 | 521 | 592 | 781
6.48e-05 | 804 | 804 | 544 | 639 | 781 |  | 6.48e-05 | 804 | 804 | 544 | 639 | 781
6.52e-05 | 828 | 828 | 568 | 663 | 804 |  | 6.52e-05 | 828 | 828 | 568 | 663 | 804
6.56e-05 | 828 | 828 | 592 | 686 | 804 |  | 6.56e-05 | 828 | 852 | 592 | 686 | 804
6.6e-05 | 852 | 852 | 615 | 710 | 828 |  | 6.6e-05 | 852 | 852 | 615 | 710 | 828
6.64e-05 | 852 | 875 | 639 | 734 | 852 |  | 6.64e-05 | 852 | 875 | 639 | 734 | 852
6.68e-05 | 852 | 875 | 663 | 757 | 852 |  | 6.68e-05 | 852 | 875 | 686 | 757 | 852
6.72e-05 | 875 | 875 | 710 | 781 | 875 |  | 6.72e-05 | 852 | 875 | 710 | 757 | 875
6.76e-05 | 875 | 875 | 710 | 781 | 875 |  | 6.76e-05 | 852 | 875 | 734 | 781 | 875
6.8e-05 | 852 | 875 | 757 | 804 | 875 |  | 6.8e-05 | 852 | 875 | 734 | 804 | 875
6.84e-05 | 852 | 875 | 757 | 828 | 875 |  | 6.84e-05 | 852 | 875 | 757 | 828 | 875
6.88e-05 | 852 | 875 | 781 | 828 | 875 |  | 6.88e-05 | 852 | 875 | 781 | 828 | 875
6.92e-05 | 852 | 875 | 804 | 828 | 875 |  | 6.92e-05 | 852 | 875 | 804 | 828 | 875
6.96e-05 | 828 | 852 | 804 | 852 | 875 |  | 6.96e-05 | 828 | 852 | 804 | 852 | 875
7e-05 | 828 | 828 | 828 | 875 | 875 |  | 7e-05 | 828 | 852 | 828 | 852 | 875
7.04e-05 | 804 | 828 | 828 | 875 | 852 |  | 7.04e-05 | 804 | 828 | 828 | 852 | 852
7.08e-05 | 804 | 828 | 852 | 875 | 852 |  | 7.08e-05 | 804 | 828 | 852 | 875 | 852
7.12e-05 | 781 | 804 | 852 | 875 | 828 |  | 7.12e-05 | 781 | 804 | 852 | 875 | 828
7.16e-05 | 757 | 781 | 852 | 875 | 828 |  | 7.16e-05 | 757 | 804 | 875 | 875 | 828
7.2e-05 | 734 | 757 | 875 | 852 | 804 |  | 7.2e-05 | 757 | 781 | 875 | 875 | 804
7.24e-05 | 734 | 757 | 875 | 852 | 804 |  | 7.24e-05 | 734 | 757 | 875 | 852 | 804
7.28e-05 | 710 | 734 | 875 | 852 | 781 |  | 7.28e-05 | 710 | 734 | 875 | 852 | 781
7.32e-05 | 686 | 710 | 875 | 852 | 757 |  | 7.32e-05 | 686 | 710 | 875 | 852 | 757
7.36e-05 | 663 | 686 | 875 | 852 | 734 |  | 7.36e-05 | 663 | 686 | 875 | 852 | 734
7.4e-05 | 639 | 663 | 875 | 828 | 710 |  | 7.4e-05 | 639 | 663 | 875 | 828 | 710
7.44e-05 | 592 | 639 | 875 | 828 | 686 |  | 7.44e-05 | 615 | 639 | 875 | 828 | 710
7.48e-05 | 568 | 615 | 875 | 804 | 663 |  | 7.48e-05 | 592 | 615 | 875 | 804 | 686
7.52e-05 | 544 | 568 | 852 | 804 | 639 |  | 7.52e-05 | 544 | 592 | 852 | 804 | 639
7.56e-05 | 521 | 544 | 852 | 781 | 615 |  | 7.56e-05 | 521 | 544 | 852 | 781 | 615
7.6e-05 | 473 | 521 | 828 | 757 | 592 |  | 7.6e-05 | 497 | 521 | 828 | 781 | 592
7.64e-05 | 450 | 497 | 828 | 734 | 544 |  | 7.64e-05 | 450 | 497 | 828 | 757 | 568
7.68e-05 | 426 | 450 | 804 | 734 | 521 |  | 7.68e-05 | 426 | 450 | 804 | 734 | 521
7.72e-05 | 379 | 426 | 804 | 710 | 497 |  | 7.72e-05 | 402 | 426 | 804 | 710 | 497
7.76e-05 | 355 | 379 | 781 | 686 | 473 |  | 7.76e-05 | 355 | 402 | 781 | 686 | 473
7.8e-05 | 308 | 355 | 757 | 663 | 426 |  | 7.8e-05 | 331 | 355 | 757 | 686 | 426
7.84e-05 | 284 | 308 | 757 | 639 | 402 |  | 7.84e-05 | 308 | 331 | 757 | 639 | 402
7.88e-05 | 260 | 284 | 734 | 615 | 355 |  | 7.88e-05 | 260 | 284 | 734 | 615 | 379
7.92e-05 | 237 | 260 | 710 | 592 | 331 |  | 7.92e-05 | 237 | 260 | 710 | 592 | 331
7.96e-05 | 189 | 213 | 686 | 568 | 308 |  | 7.96e-05 | 213 | 237 | 686 | 568 | 308
8e-05 | 166 | 189 | 686 | 544 | 260 |  | 8e-05 | 166 | 189 | 663 | 544 | 260
8.04e-05 | 118 | 142 | 639 | 521 | 237 |  | 8.04e-05 | 142 | 166 | 639 | 521 | 237
8.08e-05 | 95 | 118 | 615 | 473 | 189 |  | 8.08e-05 | 95 | 118 | 615 | 497 | 213
8.12e-05 | 47 | 71 | 615 | 450 | 166 |  | 8.12e-05 | 71 | 95 | 592 | 450 | 166
8.16e-05 | 24 | 47 | 568 | 426 | 142 |  | 8.16e-05 | 24 | 47 | 568 | 450 | 142
8.2e-05 | 0 | 0 | 544 | 379 | 95 |  | 8.2e-05 | 0 | 24 | 544 | 402 | 118
8.24e-05 | -46 | -23 | 521 | 355 | 71 |  | 8.24e-05 | -46 | -23 | 521 | 379 | 71
8.28e-05 | -94 | -70 | 497 | 331 | 24 |  | 8.28e-05 | -70 | -46 | 497 | 331 | 24
8.32e-05 | -117 | -94 | 473 | 308 | 0 |  | 8.32e-05 | -94 | -94 | 473 | 308 | 0
8.36e-05 | -141 | -141 | 426 | 260 | -46 |  | 8.36e-05 | -141 | -117 | 450 | 284 | -23
8.4e-05 | -188 | -165 | 402 | 237 | -70 |  | 8.4e-05 | -165 | -165 | 402 | 260 | -70
8.44e-05 | -212 | -212 | 379 | 213 | -94 |  | 8.44e-05 | -188 | -188 | 379 | 213 | -94
8.48e-05 | -236 | -236 | 355 | 189 | -141 |  | 8.48e-05 | -236 | -212 | 355 | 189 | -141
8.52e-05 | -259 | -259 | 308 | 142 | -165 |  | 8.52e-05 | -259 | -236 | 331 | 166 | -165
8.56e-05 | -307 | -283 | 284 | 118 | -212 |  | 8.56e-05 | -283 | -259 | 284 | 118 | -188
8.6e-05 | -307 | -307 | 260 | 71 | -236 |  | 8.6e-05 | -307 | -307 | 260 | 95 | -236
8.64e-05 | -330 | -330 | 237 | 47 | -259 |  | 8.64e-05 | -330 | -330 | 237 | 71 | -259
8.68e-05 | -354 | -354 | 213 | 24 | -283 |  | 8.68e-05 | -354 | -354 | 213 | 24 | -283
8.72e-05 | -378 | -378 | 166 | 0 | -307 |  | 8.72e-05 | -378 | -378 | 166 | 0 | -307
8.76e-05 | -401 | -401 | 142 | -46 | -330 |  | 8.76e-05 | -401 | -401 | 142 | -23 | -330
8.8e-05 | -425 | -425 | 118 | -70 | -354 |  | 8.8e-05 | -425 | -425 | 95 | -70 | -354
8.84e-05 | -449 | -449 | 95 | -94 | -378 |  | 8.84e-05 | -449 | -449 | 71 | -94 | -378
8.88e-05 | -449 | -472 | 47 | -141 | -401 |  | 8.88e-05 | -449 | -472 | 47 | -117 | -401
8.92e-05 | -496 | -496 | 24 | -165 | -425 |  | 8.92e-05 | -472 | -496 | 24 | -141 | -425
8.96e-05 | -496 | -520 | 0 | -188 | -449 |  | 8.96e-05 | -496 | -496 | -23 | -188 | -449
9e-05 | -520 | -520 | -46 | -212 | -472 |  | 9e-05 | -496 | -520 | -46 | -212 | -472
9.04e-05 | -543 | -543 | -70 | -236 | -496 |  | 9.04e-05 | -520 | -543 | -70 | -236 | -496
9.08e-05 | -543 | -543 | -117 | -283 | -520 |  | 9.08e-05 | -543 | -543 | -94 | -259 | -496
9.12e-05 | -543 | -567 | -117 | -307 | -520 |  | 9.12e-05 | -543 | -567 | -141 | -283 | -520
9.16e-05 | -567 | -567 | -165 | -330 | -543 |  | 9.16e-05 | -567 | -567 | -165 | -307 | -543
9.2e-05 | -567 | -591 | -188 | -330 | -567 |  | 9.2e-05 | -567 | -567 | -188 | -330 | -543
9.24e-05 | -567 | -591 | -212 | -354 | -567 |  | 9.24e-05 | -567 | -591 | -212 | -354 | -567
9.28e-05 | -567 | -591 | -236 | -378 | -567 |  | 9.28e-05 | -567 | -591 | -236 | -378 | -567
9.32e-05 | -591 | -591 | -259 | -401 | -591 |  | 9.32e-05 | -567 | -591 | -259 | -401 | -591
9.36e-05 | -567 | -614 | -283 | -425 | -591 |  | 9.36e-05 | -567 | -591 | -283 | -425 | -591
9.4e-05 | -591 | -614 | -307 | -449 | -591 |  | 9.4e-05 | -591 | -614 | -307 | -449 | -591
9.44e-05 | -591 | -614 | -330 | -449 | -591 |  | 9.44e-05 | -591 | -614 | -330 | -449 | -614
9.48e-05 | -591 | -614 | -354 | -472 | -614 |  | 9.48e-05 | -591 | -614 | -354 | -472 | -614
9.52e-05 | -567 | -614 | -378 | -496 | -614 |  | 9.52e-05 | -567 | -614 | -378 | -496 | -614
9.56e-05 | -567 | -614 | -401 | -520 | -614 |  | 9.56e-05 | -591 | -614 | -401 | -496 | -614
9.6e-05 | -567 | -591 | -401 | -520 | -614 |  | 9.6e-05 | -567 | -614 | -425 | -520 | -614
9.64e-05 | -567 | -591 | -449 | -543 | -614 |  | 9.64e-05 | -567 | -591 | -449 | -543 | -614
9.68e-05 | -567 | -591 | -449 | -543 | -614 |  | 9.68e-05 | -567 | -591 | -449 | -543 | -614
9.72e-05 | -543 | -567 | -472 | -543 | -591 |  | 9.72e-05 | -543 | -591 | -472 | -543 | -614
9.76e-05 | -543 | -567 | -496 | -567 | -591 |  | 9.76e-05 | -543 | -567 | -496 | -567 | -614
9.8e-05 | -520 | -567 | -496 | -567 | -591 |  | 9.8e-05 | -543 | -567 | -520 | -567 | -591
9.84e-05 | -520 | -543 | -520 | -591 | -591 |  | 9.84e-05 | -520 | -543 | -520 | -567 | -591
9.88e-05 | -496 | -543 | -520 | -591 | -591 |  | 9.88e-05 | -520 | -543 | -543 | -591 | -591
9.92e-05 | -472 | -520 | -543 | -591 | -567 |  | 9.92e-05 | -496 | -520 | -543 | -591 | -591
9.96e-05 | -472 | -520 | -543 | -591 | -567 |  | 9.96e-05 | -472 | -520 | -567 | -614 | -567
0.0001 | -449 | -496 | -567 | -591 | -543 |  | 0.0001 | -472 | -496 | -567 | -591 | -567
0.0001004 | -425 | -472 | -567 | -591 | -543 |  | 0.0001004 | -449 | -496 | -567 | -591 | -543
0.0001008 | -425 | -449 | -591 | -591 | -520 |  | 0.0001008 | -425 | -472 | -591 | -591 | -520
0.0001012 | -401 | -449 | -591 | -591 | -520 |  | 0.0001012 | -401 | -449 | -591 | -591 | -520
0.0001016 | -378 | -425 | -591 | -591 | -496 |  | 0.0001016 | -401 | -425 | -591 | -591 | -496
0.000102 | -354 | -401 | -591 | -591 | -472 |  | 0.000102 | -378 | -401 | -591 | -591 | -496
0.0001024 | -330 | -378 | -591 | -591 | -449 |  | 0.0001024 | -354 | -378 | -614 | -591 | -472
0.0001028 | -330 | -354 | -591 | -567 | -449 |  | 0.0001028 | -330 | -378 | -614 | -591 | -449
0.0001032 | -307 | -330 | -614 | -567 | -425 |  | 0.0001032 | -307 | -354 | -614 | -567 | -425
0.0001036 | -283 | -330 | -614 | -567 | -401 |  | 0.0001036 | -307 | -330 | -614 | -567 | -401
0.000104 | -259 | -307 | -614 | -567 | -378 |  | 0.000104 | -283 | -307 | -614 | -567 | -378
0.0001044 | -236 | -283 | -614 | -543 | -354 |  | 0.0001044 | -259 | -283 | -614 | -567 | -378
0.0001048 | -212 | -259 | -591 | -543 | -330 |  | 0.0001048 | -236 | -283 | -614 | -543 | -354
0.0001052 | -212 | -236 | -591 | -520 | -307 |  | 0.0001052 | -212 | -236 | -591 | -543 | -330
0.0001056 | -188 | -212 | -591 | -520 | -307 |  | 0.0001056 | -188 | -212 | -591 | -520 | -307
0.000106 | -141 | -188 | -591 | -496 | -283 |  | 0.000106 | -165 | -188 | -591 | -520 | -283
0.0001064 | -141 | -165 | -591 | -496 | -259 |  | 0.0001064 | -141 | -165 | -591 | -496 | -259
0.0001068 | -94 | -141 | -591 | -472 | -236 |  | 0.0001068 | -141 | -141 | -567 | -496 | -236
0.0001072 | -94 | -117 | -567 | -472 | -212 |  | 0.0001072 | -94 | -117 | -567 | -472 | -212
0.0001076 | -70 | -94 | -567 | -449 | -188 |  | 0.0001076 | -70 | -94 | -567 | -449 | -188
0.000108 | -46 | -70 | -543 | -425 | -165 |  | 0.000108 | -46 | -70 | -543 | -449 | -165
0.0001084 | -23 | -46 | -543 | -401 | -141 |  | 0.0001084 | -23 | -46 | -543 | -425 | -141
0.0001088 | 0 | -23 | -543 | -401 | -117 |  | 0.0001088 | 0 | -23 | -520 | -401 | -117
0.0001092 | 24 | 0 | -520 | -378 | -94 |  | 0.0001092 | 24 | 0 | -520 | -401 | -94
0.0001096 | 47 | 24 | -496 | -354 | -46 |  | 0.0001096 | 47 | 24 | -496 | -378 | -70
0.00011 | 71 | 47 | -496 | -354 | -46 |  | 0.00011 | 47 | 47 | -496 | -354 | -46
0.0001104 | 95 | 71 | -472 | -330 | -23 |  | 0.0001104 | 71 | 71 | -472 | -330 | -23
0.0001108 | 95 | 95 | -449 | -307 | 0 |  | 0.0001108 | 95 | 71 | -449 | -307 | 0
0.0001112 | 142 | 118 | -449 | -283 | 24 |  | 0.0001112 | 118 | 95 | -449 | -307 | 24
0.0001116 | 142 | 142 | -425 | -259 | 47 |  | 0.0001116 | 142 | 142 | -425 | -283 | 47
0.000112 | 166 | 166 | -401 | -236 | 71 |  | 0.000112 | 166 | 142 | -401 | -259 | 71
0.0001124 | 189 | 166 | -378 | -212 | 95 |  | 0.0001124 | 166 | 166 | -378 | -236 | 95
0.0001128 | 189 | 189 | -378 | -212 | 118 |  | 0.0001128 | 189 | 189 | -378 | -212 | 118
0.0001132 | 213 | 213 | -354 | -188 | 142 |  | 0.0001132 | 213 | 213 | -354 | -188 | 142
0.0001136 | 237 | 237 | -330 | -165 | 166 |  | 0.0001136 | 237 | 213 | -330 | -165 | 166
0.000114 | 237 | 237 | -307 | -141 | 166 |  | 0.000114 | 237 | 237 | -330 | -141 | 166
0.0001144 | 260 | 260 | -307 | -117 | 189 |  | 0.0001144 | 237 | 260 | -307 | -141 | 189
0.0001148 | 260 | 260 | -283 | -94 | 213 |  | 0.0001148 | 260 | 260 | -283 | -117 | 213
0.0001152 | 284 | 284 | -259 | -70 | 237 |  | 0.0001152 | 260 | 284 | -259 | -70 | 237
0.0001156 | 284 | 308 | -236 | -46 | 237 |  | 0.0001156 | 284 | 308 | -236 | -70 | 237
0.000116 | 308 | 308 | -212 | -23 | 260 |  | 0.000116 | 308 | 308 | -212 | -46 | 260
0.0001164 | 308 | 331 | -212 | 0 | 284 |  | 0.0001164 | 308 | 331 | -212 | -23 | 260
0.0001168 | 331 | 331 | -188 | 24 | 284 |  | 0.0001168 | 308 | 331 | -188 | 0 | 284
0.0001172 | 331 | 355 | -165 | 24 | 308 |  | 0.0001172 | 331 | 355 | -141 | 24 | 308
0.0001176 | 331 | 355 | -141 | 47 | 308 |  | 0.0001176 | 331 | 355 | -117 | 24 | 308
0.000118 | 331 | 355 | -117 | 71 | 331 |  | 0.000118 | 355 | 355 | -117 | 47 | 331
0.0001184 | 355 | 379 | -94 | 95 | 331 |  | 0.0001184 | 355 | 379 | -94 | 95 | 331
0.0001188 | 355 | 379 | -70 | 118 | 355 |  | 0.0001188 | 355 | 379 | -70 | 95 | 355
0.0001192 | 355 | 379 | -46 | 142 | 355 |  | 0.0001192 | 355 | 379 | -46 | 118 | 355
0.0001196 | 355 | 379 | -23 | 142 | 355 |  | 0.0001196 | 355 | 379 | -23 | 142 | 355
0.00012 | 355 | 379 | -23 | 166 | 379 |  | 0.00012 | 355 | 379 | 0 | 166 | 379
0.0001204 | 355 | 379 | 0 | 189 | 379 |  | 0.0001204 | 355 | 402 | 24 | 166 | 379
0.0001208 | 355 | 379 | 24 | 189 | 379 |  | 0.0001208 | 355 | 402 | 24 | 189 | 379
0.0001212 | 355 | 379 | 47 | 213 | 379 |  | 0.0001212 | 379 | 402 | 47 | 213 | 379
0.0001216 | 355 | 379 | 71 | 237 | 402 |  | 0.0001216 | 355 | 402 | 71 | 213 | 402
0.000122 | 355 | 379 | 95 | 237 | 379 |  | 0.000122 | 355 | 379 | 95 | 237 | 402
0.0001224 | 355 | 402 | 118 | 260 | 402 |  | 0.0001224 | 355 | 402 | 118 | 260 | 402
0.0001228 | 355 | 379 | 118 | 260 | 402 |  | 0.0001228 | 355 | 379 | 142 | 260 | 402
0.0001232 | 355 | 379 | 142 | 284 | 402 |  | 0.0001232 | 355 | 379 | 142 | 260 | 402
0.0001236 | 331 | 379 | 166 | 284 | 402 |  | 0.0001236 | 355 | 379 | 166 | 284 | 402
0.000124 | 331 | 379 | 189 | 308 | 402 |  | 0.000124 | 331 | 379 | 189 | 308 | 402
0.0001244 | 331 | 355 | 189 | 308 | 402 |  | 0.0001244 | 331 | 379 | 189 | 308 | 402
0.0001248 | 308 | 355 | 213 | 331 | 379 |  | 0.0001248 | 331 | 355 | 213 | 308 | 402
0.0001252 | 308 | 355 | 237 | 331 | 379 |  | 0.0001252 | 331 | 355 | 237 | 331 | 379
0.0001256 | 308 | 331 | 237 | 331 | 379 |  | 0.0001256 | 308 | 355 | 237 | 331 | 379
0.000126 | 284 | 331 | 260 | 355 | 379 |  | 0.000126 | 308 | 355 | 260 | 331 | 379
0.0001264 | 284 | 331 | 260 | 355 | 379 |  | 0.0001264 | 284 | 331 | 260 | 355 | 379
0.0001268 | 260 | 308 | 284 | 355 | 355 |  | 0.0001268 | 284 | 331 | 284 | 355 | 355
0.0001272 | 260 | 308 | 284 | 379 | 355 |  | 0.0001272 | 284 | 308 | 284 | 355 | 355
0.0001276 | 260 | 284 | 308 | 379 | 355 |  | 0.0001276 | 260 | 308 | 308 | 379 | 355
0.000128 | 237 | 284 | 308 | 379 | 331 |  | 0.000128 | 260 | 284 | 308 | 379 | 355
0.0001284 | 237 | 260 | 331 | 379 | 331 |  | 0.0001284 | 237 | 284 | 331 | 379 | 331
0.0001288 | 213 | 260 | 331 | 379 | 308 |  | 0.0001288 | 237 | 260 | 331 | 379 | 331
0.0001292 | 213 | 237 | 331 | 379 | 308 |  | 0.0001292 | 213 | 260 | 355 | 379 | 308
0.0001296 | 189 | 237 | 355 | 379 | 284 |  | 0.0001296 | 213 | 237 | 355 | 379 | 308
0.00013 | 189 | 213 | 355 | 379 | 284 |  | 0.00013 | 189 | 237 | 355 | 379 | 284
0.0001304 | 166 | 213 | 379 | 379 | 260 |  | 0.0001304 | 189 | 213 | 379 | 379 | 260
0.0001308 | 166 | 189 | 379 | 379 | 260 |  | 0.0001308 | 166 | 189 | 379 | 379 | 260
0.0001312 | 142 | 166 | 379 | 379 | 237 |  | 0.0001312 | 166 | 189 | 379 | 379 | 260
0.0001316 | 142 | 166 | 379 | 379 | 237 |  | 0.0001316 | 142 | 166 | 402 | 379 | 237
0.000132 | 118 | 142 | 402 | 379 | 237 |  | 0.000132 | 118 | 142 | 379 | 379 | 237
0.0001324 | 95 | 118 | 402 | 379 | 213 |  | 0.0001324 | 118 | 142 | 402 | 379 | 213
0.0001328 | 95 | 118 | 402 | 355 | 189 |  | 0.0001328 | 95 | 118 | 402 | 379 | 189
0.0001332 | 71 | 95 | 402 | 355 | 166 |  | 0.0001332 | 71 | 118 | 402 | 379 | 189
0.0001336 | 47 | 71 | 402 | 355 | 166 |  | 0.0001336 | 71 | 95 | 402 | 355 | 166
0.000134 | 47 | 71 | 402 | 355 | 142 |  | 0.000134 | 47 | 71 | 402 | 355 | 142
0.0001344 | 24 | 47 | 402 | 331 | 142 |  | 0.0001344 | 24 | 71 | 402 | 355 | 142
0.0001348 | 0 | 24 | 402 | 331 | 118 |  | 0.0001348 | 24 | 47 | 402 | 331 | 118
0.0001352 | 0 | 0 | 402 | 331 | 95 |  | 0.0001352 | 0 | 24 | 402 | 331 | 118
0.0001356 | -23 | 0 | 402 | 308 | 95 |  | 0.0001356 | 0 | 24 | 402 | 331 | 95
0.000136 | -23 | -23 | 402 | 308 | 71 |  | 0.000136 | -23 | 0 | 402 | 308 | 71
0.0001364 | -46 | -23 | 379 | 308 | 47 |  | 0.0001364 | -23 | -23 | 402 | 308 | 47
0.0001368 | -46 | -46 | 379 | 284 | 24 |  | 0.0001368 | -46 | -46 | 379 | 284 | 47
0.0001372 | -70 | -46 | 379 | 284 | 24 |  | 0.0001372 | -70 | -46 | 379 | 284 | 24
0.0001376 | -70 | -70 | 379 | 260 | 0 |  | 0.0001376 | -70 | -70 | 379 | 260 | 0
0.000138 | -94 | -94 | 379 | 260 | -23 |  | 0.000138 | -94 | -70 | 355 | 260 | 0
0.0001384 | -117 | -117 | 355 | 237 | -23 |  | 0.0001384 | -94 | -94 | 355 | 260 | -23
0.0001388 | -117 | -117 | 355 | 237 | -46 |  | 0.0001388 | -117 | -117 | 355 | 237 | -46
0.0001392 | -141 | -141 | 355 | 213 | -46 |  | 0.0001392 | -117 | -117 | 355 | 237 | -46
0.0001396 | -141 | -141 | 331 | 213 | -70 |  | 0.0001396 | -141 | -141 | 331 | 213 | -70
0.00014 | -165 | -165 | 331 | 189 | -94 |  | 0.00014 | -141 | -141 | 331 | 213 | -70
0.0001404 | -165 | -165 | 331 | 189 | -94 |  | 0.0001404 | -165 | -165 | 331 | 189 | -94
0.0001408 | -165 | -188 | 308 | 166 | -117 |  | 0.0001408 | -165 | -165 | 308 | 189 | -117
0.0001412 | -188 | -188 | 308 | 166 | -141 |  | 0.0001412 | -188 | -188 | 308 | 166 | -117
0.0001416 | -188 | -188 | 284 | 142 | -141 |  | 0.0001416 | -188 | -188 | 284 | 142 | -141
0.000142 | -188 | -212 | 284 | 118 | -165 |  | 0.000142 | -188 | -212 | 284 | 142 | -141
0.0001424 | -212 | -212 | 260 | 118 | -165 |  | 0.0001424 | -188 | -212 | 260 | 118 | -165
0.0001428 | -212 | -236 | 260 | 95 | -165 |  | 0.0001428 | -188 | -212 | 260 | 118 | -165
0.0001432 | -212 | -236 | 260 | 71 | -188 |  | 0.0001432 | -212 | -236 | 237 | 95 | -188
0.0001436 | -212 | -236 | 237 | 71 | -212 |  | 0.0001436 | -212 | -236 | 237 | 95 | -188
0.000144 | -212 | -259 | 237 | 47 | -212 |  | 0.000144 | -212 | -236 | 237 | 71 | -212
0.0001444 | -236 | -259 | 213 | 47 | -212 |  | 0.0001444 | -236 | -259 | 213 | 47 | -212
0.0001448 | -236 | -259 | 213 | 24 | -236 |  | 0.0001448 | -236 | -259 | 189 | 47 | -236
0.0001452 | -236 | -259 | 189 | 0 | -236 |  | 0.0001452 | -236 | -259 | 189 | 24 | -236
0.0001456 | -236 | -259 | 166 | 0 | -236 |  | 0.0001456 | -236 | -259 | 189 | 0 | -236
0.000146 | -236 | -283 | 166 | -23 | -236 |  | 0.000146 | -236 | -259 | 166 | 0 | -259
0.0001464 | -236 | -283 | 142 | -23 | -259 |  | 0.0001464 | -259 | -283 | 142 | -23 | -259
0.0001468 | -236 | -283 | 142 | -46 | -259 |  | 0.0001468 | -259 | -283 | 142 | -23 | -259
0.0001472 | -236 | -283 | 118 | -46 | -259 |  | 0.0001472 | -236 | -283 | 118 | -46 | -259
0.0001476 | -236 | -283 | 118 | -70 | -259 |  | 0.0001476 | -259 | -283 | 118 | -46 | -259
0.000148 | -236 | -283 | 95 | -70 | -283 |  | 0.000148 | -236 | -283 | 95 | -70 | -283
0.0001484 | -236 | -283 | 71 | -94 | -283 |  | 0.0001484 | -259 | -283 | 71 | -94 | -283
0.0001488 | -236 | -283 | 71 | -117 | -283 |  | 0.0001488 | -236 | -283 | 71 | -94 | -283
0.0001492 | -236 | -283 | 47 | -117 | -283 |  | 0.0001492 | -236 | -283 | 47 | -117 | -283
0.0001496 | -236 | -283 | 47 | -141 | -283 |  | 0.0001496 | -236 | -283 | 24 | -117 | -283
0.00015 | -236 | -259 | 24 | -141 | -283 |  | 0.00015 | -236 | -259 | 24 | -141 | -283
0.0001504 | -236 | -259 | 0 | -141 | -283 |  | 0.0001504 | -236 | -283 | 0 | -141 | -283
0.0001508 | -236 | -259 | 0 | -165 | -283 |  | 0.0001508 | -236 | -259 | 0 | -141 | -283
0.0001512 | -212 | -259 | -23 | -165 | -283 |  | 0.0001512 | -236 | -259 | -23 | -165 | -283
0.0001516 | -212 | -259 | -23 | -188 | -283 |  | 0.0001516 | -236 | -259 | -46 | -165 | -283
0.000152 | -212 | -259 | -46 | -188 | -283 |  | 0.000152 | -212 | -259 | -46 | -188 | -283
0.0001524 | -212 | -236 | -46 | -188 | -283 |  | 0.0001524 | -212 | -259 | -70 | -188 | -283
0.0001528 | -188 | -236 | -70 | -212 | -283 |  | 0.0001528 | -212 | -236 | -70 | -188 | -283
0.0001532 | -188 | -236 | -94 | -212 | -283 |  | 0.0001532 | -188 | -236 | -94 | -212 | -283
0.0001536 | -188 | -212 | -94 | -236 | -259 |  | 0.0001536 | -188 | -236 | -94 | -212 | -283
0.000154 | -188 | -212 | -117 | -236 | -259 |  | 0.000154 | -188 | -236 | -117 | -212 | -259
0.0001544 | -165 | -212 | -117 | -236 | -259 |  | 0.0001544 | -188 | -212 | -117 | -236 | -259
0.0001548 | -165 | -188 | -141 | -236 | -259 |  | 0.0001548 | -165 | -212 | -141 | -236 | -259
0.0001552 | -165 | -188 | -141 | -236 | -236 |  | 0.0001552 | -165 | -188 | -141 | -236 | -259
0.0001556 | -141 | -188 | -165 | -236 | -236 |  | 0.0001556 | -165 | -188 | -165 | -236 | -236
0.000156 | -141 | -188 | -165 | -259 | -236 |  | 0.000156 | -141 | -188 | -165 | -236 | -236
0.0001564 | -141 | -165 | -188 | -259 | -236 |  | 0.0001564 | -141 | -188 | -188 | -259 | -236
0.0001568 | -117 | -165 | -188 | -259 | -212 |  | 0.0001568 | -141 | -165 | -188 | -259 | -212
0.0001572 | -117 | -141 | -188 | -259 | -212 |  | 0.0001572 | -117 | -165 | -212 | -259 | -212
0.0001576 | -94 | -141 | -212 | -259 | -212 |  | 0.0001576 | -117 | -141 | -212 | -259 | -212
0.000158 | -94 | -117 | -212 | -259 | -188 |  | 0.000158 | -117 | -141 | -212 | -259 | -188
0.0001584 | -94 | -117 | -212 | -259 | -188 |  | 0.0001584 | -94 | -117 | -236 | -259 | -188
0.0001588 | -70 | -117 | -236 | -259 | -188 |  | 0.0001588 | -94 | -117 | -236 | -259 | -165
0.0001592 | -70 | -94 | -236 | -259 | -165 |  | 0.0001592 | -70 | -117 | -236 | -283 | -165
0.0001596 | -46 | -94 | -236 | -283 | -165 |  | 0.0001596 | -70 | -94 | -236 | -259 | -165
0.00016 | -46 | -70 | -259 | -259 | -141 |  | 0.00016 | -70 | -94 | -259 | -259 | -141
0.0001604 | -46 | -70 | -259 | -259 | -141 |  | 0.0001604 | -46 | -70 | -259 | -259 | -141
0.0001608 | -23 | -46 | -259 | -259 | -117 |  | 0.0001608 | -46 | -70 | -259 | -259 | -117
0.0001612 | -23 | -46 | -259 | -259 | -117 |  | 0.0001612 | -23 | -46 | -259 | -259 | -117
0.0001616 | -23 | -23 | -259 | -259 | -117 |  | 0.0001616 | -23 | -46 | -259 | -259 | -117
0.000162 | 0 | -23 | -259 | -259 | -94 |  | 0.000162 | -23 | -23 | -283 | -259 | -94
0.0001624 | 0 | 0 | -283 | -259 | -70 |  | 0.0001624 | 0 | -23 | -283 | -259 | -94
0.0001628 | 24 | 0 | -283 | -259 | -70 |  | 0.0001628 | 0 | 0 | -283 | -259 | -70
0.0001632 | 24 | 0 | -283 | -236 | -46 |  | 0.0001632 | 0 | 0 | -283 | -259 | -70
0.0001636 | 24 | 24 | -283 | -236 | -46 |  | 0.0001636 | 24 | 24 | -283 | -259 | -70
0.000164 | 47 | 24 | -283 | -236 | -46 |  | 0.000164 | 24 | 24 | -283 | -236 | -46
0.0001644 | 47 | 47 | -283 | -236 | -23 |  | 0.0001644 | 24 | 24 | -283 | -236 | -23
0.0001648 | 47 | 47 | -283 | -236 | -23 |  | 0.0001648 | 47 | 47 | -283 | -236 | -23
0.0001652 | 47 | 47 | -283 | -212 | 0 |  | 0.0001652 | 47 | 47 | -283 | -236 | 0
0.0001656 | 71 | 71 | -283 | -212 | 0 |  | 0.0001656 | 71 | 71 | -283 | -212 | 0
0.000166 | 71 | 71 | -283 | -212 | 24 |  | 0.000166 | 71 | 71 | -283 | -212 | 0
0.0001664 | 71 | 71 | -283 | -188 | 24 |  | 0.0001664 | 71 | 71 | -283 | -212 | 24
0.0001668 | 95 | 95 | -283 | -188 | 24 |  | 0.0001668 | 71 | 95 | -283 | -188 | 24
0.0001672 | 95 | 95 | -283 | -188 | 47 |  | 0.0001672 | 95 | 95 | -283 | -188 | 47
0.0001676 | 95 | 118 | -283 | -165 | 47 |  | 0.0001676 | 95 | 95 | -283 | -188 | 47
0.000168 | 95 | 118 | -283 | -165 | 47 |  | 0.000168 | 95 | 118 | -283 | -165 | 71
0.0001684 | 118 | 118 | -259 | -165 | 71 |  | 0.0001684 | 118 | 118 | -259 | -165 | 71
0.0001688 | 118 | 118 | -259 | -141 | 71 |  | 0.0001688 | 95 | 118 | -259 | -165 | 71
0.0001692 | 118 | 142 | -259 | -141 | 95 |  | 0.0001692 | 118 | 118 | -259 | -141 | 71
0.0001696 | 118 | 142 | -259 | -141 | 95 |  | 0.0001696 | 118 | 142 | -259 | -141 | 95
0.00017 | 118 | 142 | -259 | -117 | 95 |  | 0.00017 | 118 | 142 | -259 | -117 | 95
0.0001704 | 118 | 142 | -236 | -117 | 118 |  | 0.0001704 | 118 | 142 | -236 | -117 | 118
0.0001708 | 142 | 142 | -236 | -117 | 118 |  | 0.0001708 | 118 | 142 | -236 | -117 | 118
0.0001712 | 142 | 166 | -236 | -94 | 118 |  | 0.0001712 | 142 | 166 | -236 | -94 | 118
0.0001716 | 142 | 166 | -212 | -94 | 118 |  | 0.0001716 | 142 | 166 | -236 | -94 | 118
0.000172 | 142 | 166 | -212 | -70 | 142 |  | 0.000172 | 142 | 166 | -212 | -94 | 142
0.0001724 | 142 | 166 | -212 | -70 | 142 |  | 0.0001724 | 142 | 166 | -212 | -70 | 142
0.0001728 | 142 | 166 | -212 | -46 | 166 |  | 0.0001728 | 142 | 166 | -212 | -70 | 142
0.0001732 | 142 | 166 | -188 | -46 | 166 |  | 0.0001732 | 142 | 166 | -188 | -46 | 142
0.0001736 | 142 | 166 | -188 | -46 | 166 |  | 0.0001736 | 142 | 166 | -188 | -46 | 142
0.000174 | 142 | 166 | -188 | -23 | 166 |  | 0.000174 | 142 | 166 | -165 | -46 | 166
0.0001744 | 142 | 166 | -165 | -23 | 166 |  | 0.0001744 | 142 | 189 | -165 | -23 | 166
0.0001748 | 142 | 166 | -165 | -23 | 166 |  | 0.0001748 | 142 | 166 | -165 | -23 | 166
0.0001752 | 142 | 166 | -141 | 0 | 166 |  | 0.0001752 | 142 | 166 | -141 | 0 | 166
0.0001756 | 142 | 166 | -141 | 0 | 189 |  | 0.0001756 | 142 | 189 | -141 | 0 | 166
0.000176 | 142 | 166 | -141 | 24 | 189 |  | 0.000176 | 142 | 166 | -141 | 0 | 189
0.0001764 | 142 | 166 | -117 | 24 | 189 |  | 0.0001764 | 142 | 189 | -117 | 24 | 189
0.0001768 | 142 | 166 | -117 | 24 | 189 |  | 0.0001768 | 142 | 166 | -117 | 24 | 189
0.0001772 | 142 | 166 | -117 | 47 | 189 |  | 0.0001772 | 142 | 166 | -94 | 47 | 189
0.0001776 | 142 | 166 | -94 | 47 | 189 |  | 0.0001776 | 142 | 166 | -94 | 47 | 189
0.000178 | 142 | 166 | -94 | 71 | 189 |  | 0.000178 | 142 | 166 | -94 | 47 | 189
0.0001784 | 118 | 166 | -94 | 71 | 189 |  | 0.0001784 | 118 | 166 | -70 | 71 | 189
0.0001788 | 118 | 166 | -70 | 71 | 189 |  | 0.0001788 | 142 | 166 | -70 | 71 | 189
0.0001792 | 118 | 142 | -46 | 95 | 189 |  | 0.0001792 | 142 | 166 | -46 | 71 | 189
0.0001796 | 118 | 142 | -46 | 95 | 166 |  | 0.0001796 | 118 | 166 | -46 | 71 | 166
0.00018 | 118 | 142 | -46 | 95 | 166 |  | 0.00018 | 118 | 142 | -23 | 95 | 166
0.0001804 | 118 | 142 | -23 | 95 | 166 |  | 0.0001804 | 118 | 142 | -23 | 95 | 166
0.0001808 | 118 | 142 | -23 | 118 | 166 |  | 0.0001808 | 118 | 142 | -23 | 95 | 166
0.0001812 | 95 | 118 | 0 | 118 | 166 |  | 0.0001812 | 118 | 142 | 0 | 118 | 166
0.0001816 | 95 | 118 | 0 | 118 | 166 |  | 0.0001816 | 118 | 142 | 0 | 118 | 166
0.000182 | 95 | 118 | 0 | 118 | 166 |  | 0.000182 | 95 | 118 | 0 | 118 | 166
0.0001824 | 95 | 118 | 24 | 142 | 166 |  | 0.0001824 | 95 | 118 | 24 | 118 | 166
0.0001828 | 71 | 118 | 24 | 142 | 166 |  | 0.0001828 | 95 | 118 | 24 | 142 | 166
0.0001832 | 71 | 95 | 47 | 142 | 142 |  | 0.0001832 | 95 | 118 | 47 | 142 | 142
0.0001836 | 71 | 95 | 47 | 142 | 142 |  | 0.0001836 | 71 | 118 | 47 | 142 | 142
0.000184 | 71 | 95 | 47 | 166 | 142 |  | 0.000184 | 71 | 95 | 47 | 142 | 142
0.0001844 | 47 | 95 | 71 | 166 | 142 |  | 0.0001844 | 71 | 95 | 71 | 142 | 142
0.0001848 | 47 | 71 | 71 | 166 | 118 |  | 0.0001848 | 71 | 71 | 71 | 142 | 142
0.0001852 | 47 | 71 | 71 | 166 | 118 |  | 0.0001852 | 47 | 71 | 95 | 166 | 118
0.0001856 | 47 | 71 | 95 | 166 | 118 |  | 0.0001856 | 47 | 71 | 95 | 166 | 118
0.000186 | 47 | 47 | 95 | 166 | 118 |  | 0.000186 | 47 | 71 | 95 | 166 | 118
0.0001864 | 47 | 47 | 95 | 166 | 118 |  | 0.0001864 | 47 | 47 | 118 | 166 | 95
0.0001868 | 24 | 47 | 118 | 166 | 95 |  | 0.0001868 | 24 | 47 | 118 | 166 | 95
0.0001872 | 24 | 24 | 118 | 166 | 95 |  | 0.0001872 | 24 | 47 | 118 | 166 | 95
0.0001876 | 24 | 24 | 118 | 166 | 71 |  | 0.0001876 | 24 | 24 | 118 | 166 | 95
0.000188 | 0 | 24 | 142 | 166 | 71 |  | 0.000188 | 24 | 24 | 118 | 166 | 71
0.0001884 | 0 | 24 | 142 | 166 | 71 |  | 0.0001884 | 0 | 24 | 142 | 166 | 71
0.0001888 | 0 | 0 | 142 | 166 | 71 |  | 0.0001888 | 0 | 0 | 142 | 166 | 71
0.0001892 | 0 | 0 | 142 | 166 | 47 |  | 0.0001892 | 0 | 0 | 142 | 166 | 47
0.0001896 | 0 | 0 | 142 | 166 | 47 |  | 0.0001896 | 0 | 0 | 166 | 166 | 47
0.00019 | -23 | -23 | 166 | 166 | 47 |  | 0.00019 | -23 | 0 | 166 | 166 | 47
0.0001904 | -23 | -23 | 166 | 166 | 47 |  | 0.0001904 | -23 | 0 | 166 | 166 | 47
0.0001908 | -23 | -23 | 166 | 166 | 24 |  | 0.0001908 | -23 | -23 | 166 | 166 | 24
0.0001912 | -23 | -23 | 166 | 166 | 24 |  | 0.0001912 | -23 | -23 | 166 | 166 | 24
0.0001916 | -23 | -46 | 166 | 166 | 0 |  | 0.0001916 | -23 | -23 | 166 | 166 | 0
0.000192 | -46 | -46 | 166 | 166 | 0 |  | 0.000192 | -46 | -46 | 166 | 166 | 24
0.0001924 | -46 | -46 | 189 | 166 | 0 |  | 0.0001924 | -46 | -46 | 189 | 166 | 0
0.0001928 | -46 | -46 | 189 | 142 | 0 |  | 0.0001928 | -46 | -46 | 189 | 166 | 0
0.0001932 | -46 | -70 | 189 | 142 | -23 |  | 0.0001932 | -46 | -46 | 189 | 142 | -23
0.0001936 | -46 | -70 | 189 | 142 | -23 |  | 0.0001936 | -46 | -46 | 189 | 142 | -23
0.000194 | -70 | -70 | 189 | 142 | -23 |  | 0.000194 | -70 | -70 | 189 | 142 | -23
0.0001944 | -70 | -70 | 189 | 142 | -46 |  | 0.0001944 | -70 | -70 | 189 | 142 | -23
0.0001948 | -70 | -94 | 189 | 142 | -46 |  | 0.0001948 | -70 | -70 | 189 | 142 | -23
0.0001952 | -70 | -94 | 189 | 142 | -46 |  | 0.0001952 | -70 | -94 | 189 | 142 | -46
0.0001956 | -70 | -94 | 189 | 118 | -46 |  | 0.0001956 | -70 | -94 | 189 | 118 | -46
0.000196 | -70 | -94 | 189 | 118 | -70 |  | 0.000196 | -70 | -94 | 189 | 118 | -46
0.0001964 | -94 | -94 | 189 | 118 | -70 |  | 0.0001964 | -70 | -94 | 189 | 118 | -70
0.0001968 | -94 | -117 | 189 | 118 | -70 |  | 0.0001968 | -94 | -117 | 189 | 118 | -70
0.0001972 | -94 | -117 | 189 | 95 | -70 |  | 0.0001972 | -94 | -117 | 189 | 118 | -70
0.0001976 | -94 | -117 | 189 | 95 | -94 |  | 0.0001976 | -94 | -117 | 189 | 95 | -94
0.000198 | -94 | -117 | 189 | 95 | -94 |  | 0.000198 | -94 | -117 | 189 | 95 | -94
0.0001984 | -94 | -117 | 189 | 95 | -94 |  | 0.0001984 | -94 | -117 | 189 | 95 | -94
0.0001988 | -94 | -117 | 189 | 71 | -94 |  | 0.0001988 | -94 | -117 | 189 | 95 | -94
0.0001992 | -94 | -117 | 166 | 71 | -94 |  | 0.0001992 | -94 | -117 | 166 | 71 | -94
0.0001996 | -94 | -117 | 166 | 71 | -117 |  | 0.0001996 | -94 | -117 | 166 | 71 | -117
0.0002 | -94 | -117 | 166 | 71 | -117 |  | 0.0002 | -94 | -117 | 166 | 71 | -117
0.0002004 | -94 | -117 | 166 | 47 | -117 |  | 0.0002004 | -94 | -117 | 166 | 47 | -117
0.0002008 | -94 | -117 | 166 | 47 | -117 |  | 0.0002008 | -94 | -117 | 166 | 47 | -117
0.0002012 | -94 | -117 | 166 | 47 | -117 |  | 0.0002012 | -94 | -141 | 166 | 47 | -117
0.0002016 | -94 | -141 | 142 | 47 | -117 |  | 0.0002016 | -94 | -117 | 142 | 47 | -117
0.000202 | -94 | -117 | 142 | 24 | -117 |  | 0.000202 | -94 | -117 | 142 | 47 | -117
0.0002024 | -94 | -117 | 142 | 24 | -117 |  | 0.0002024 | -94 | -117 | 142 | 24 | -117
0.0002028 | -94 | -117 | 142 | 24 | -117 |  | 0.0002028 | -94 | -117 | 142 | 24 | -117
0.0002032 | -94 | -117 | 142 | 0 | -117 |  | 0.0002032 | -94 | -141 | 142 | 24 | -117
0.0002036 | -94 | -117 | 118 | 0 | -117 |  | 0.0002036 | -94 | -117 | 142 | 24 | -141
0.000204 | -94 | -117 | 118 | 0 | -141 |  | 0.000204 | -94 | -117 | 118 | 0 | -141
0.0002044 | -94 | -117 | 118 | 0 | -141 |  | 0.0002044 | -94 | -117 | 118 | 0 | -141
| 0 | 0 | 0 | 0 | 0 | #REF! |  | 0 | 0 | 0 | 0 | 0 | #REF!
| 2 | 2 | 2 | 2 | 2 | #REF! |  | 2 | 2 | 2 | 2 | 2 | #REF!
| 2000 | 2000 | 2000 | 2000 | 2000 | #REF! |  | 2000 | 2000 | 2000 | 2000 | 2000 | #REF!
| 2 | 2 | 2 | 2 | 2 | #REF! |  | 2 | 2 | 2 | 2 | 2 | #REF!