## Slide 1: MQXFS01 Lead Repair
Steve Krave, Dan Cheng
12/9/2015

## Slide 2: Last Time
Preparing to splice leads
Modifying G10 to accommodate Splices

## Slide 3: Pre-tinning of Leads
Leads and extensions were tinned using 70/30 Pb/Sn solder and Kester 1544 Rosin Activated Flux
  Tevatron Solder/flux
  VMTF Solder/flux
Good wetting observed
Consistent quality

## Slide 4: Soldering of Leads
Joints were staggered with 3 cable-splice interfaces
Splices appear to be very good.

## Slide 5: Adding Voltage Taps
Voltage taps have been added and the connectors will be rung out shortly.
Pizza box will then be re-assembled.

## Slide 8
Older Slides

## Slide 9: Positive Lead from VMTF Splice to Support Plate
At the beginning of the splice (side closer to magnet) the Kapton looks “crispy”
Burned popsicle stick standoff shows signs of overheating
Near support plate, burned Kapton and discolored staybrite are visible

## Slide 10: Positive Lead from Support Plate to Pizza Box
Below ~15cm above pizza box there was no more discoloration or signs of melted adhesive (less than ~260°C)

## Slide 11: End of Carbon on Warm Finger
Carbon dust was observed on Warm finger that ended at ~145cm from Top Plate
Roughly at the level of the pizza box
Observation that may or may not be indication of helium level…

## Slide 12: Debris Inside Dewar
On raising the magnet out of the dewar, melted lead was visible on bottom plate.
Small piece of metal had also stuck to the glass tape at the center mark of the magnet

## Slide 13: On Opening the Pizza Box
Strong smell of burned G-10
Charred Kapton around voltage tap on Negative Lead (>400°C)
Voltage tap solder on negative lead melted and flowed
Voltage tap solder on positive lead appears OK (>221°C)
  It broke when the plate was removed

## Slide 14: Remainder of Pizza Box
Other leads appear to be fine
Inside bore appears OK on first glance
Positive lead damage appears to end ~1” before exiting pizza box

## Slide 15: Pizza Box Layer 2
Splice in layer 1 was removed without issue without de-soldering
1st layer of pizza box, cover plates, and shims removed.
No visible change
  No change in staybrite
  No charred or discolored G10
  No additional smell
Tape on leads to coils on ID removed
  No visible discoloration or damage of tape, G10, or staybrite

## Slide 16: Moving Forward
We are moving forward with disassembly of the Pizza box
Dan C. Is here to work through the process
Splice will be made in straight section between pizza box and magnet

## Slide 17: Pizza Box Disassembly
Layer 2 of the pizza box was easily separated by lifting the coil to coil splice out of the G10 and slipping the G10 layer off over the leads
Layer 1 was easily removed in 4 pieces by lifting the splices out of the G10
The Pizza Box attachment plate was removed by sawing in half with hacksaw and removing attachment screws

## Slide 18: Lead Preparation
The lead supports on the positive and Negative lead have been removed.
  No signs of any heating
Kapton has been removed back to the voltage taps near the magnet end plate
Voltage taps have been cut to splice into updated electrical plan

## Slide 19: Lead Preparation
We are preparing to solder the extension leads
Standard splice fixture Fits in space with a little persuasion
Tinning of Extensions is in progress
  Practice run before working on the magnet
  Using Kester 1544 flux and 70/30 solder
Magnet leads will be cut at 220mm from end plate after tinned to prevent fraying

## Slide 20: Modifications
Additional cooling holes are being added to Pizza box in both layer 1 and layer 2 (~double the current amount)
Lead support to pizza box is having cable channel increased by 4.5mm to accommodate splice
New lead extensions will use the pizza box routing as designed

## Slide 21: Negative Lead With Scale
Scale starts at end plate.
Lead will be cut at 220 mm from end plate.
Splice will be from 100 mm to 220 mm.
WC Voltage Tap located at ~580mm from end plate.

## Slide 22: Wires for analysis
After tinning, leads will be cut, freeing up cable for analysis.
  2 strands will be removed from each cable of negative lead for RRR, Magnetization etc
  Strands will be labeled as to which cable they came from and which is the magnet side
End of lead nearest magnet will be prepared for microscopy