## Slide 1
APA Prototype Tests at FNAL
O.Prokofiev
APA design and technology
Mechanical tests
Electrical tests
Cold tests in liquid nitrogen
Summary

## Slide 2
64 wires
pc-board
Epoxy
Soldering
35 mm
Module
APA Prototype
Idea of alternate APA design is based on construction of preassembled wire modules which will be stretched and mounted on the frame pins.
The wires in the modules have a specific length so that when wires are being stretched on the frame the wire tension will reach a nominal value.
The epoxy strip applied above FR-4 board of for keeping of wire tension.
Wire soldering is served just for electrical connection.
FR-4 ( 5 mils)
Frame pin

## Slide 3
APA Assembly with Wire Modules
The APA assembly with using prefabricated modules is simple, clean and labor effective procedure. Also, the modules could be easily replaced.

## Slide 4
APA Prototype Construction
First planes assembly
Prototype features:
- 1.2 x 2.5 m prototype dimensions
- 200 lbs weight
- 4 wire planes (mesh, X, U & V planes)
- 450 U & V planes
- 150 mm diameter CuBe wire
- 500 gr wire tension
- 64 wires for X and Mesh module, 4 modules/plane
- 48 wires for U,V module, 11 modules/plane
- 60 wire modules
- 3136 wires
- 1280 readout channels (~ 50% of full scale)
Details of prototype design, technology,
fabrication and test results are available
at LBNE docdb 6044, 6335, 6531
Completed prototype

## Slide 5
    Frame flatness
    Tension
    Wire spacing
Mechanical tests

## Slide 6
Frame Flatness
The frame (1.25m x 2.5m) was fabricated by vendor.
Measured frame flatness (1.25m x 2.5m) was in range of 20 mils
The frame was straightened to meet these tolerances.

## Slide 7
Wire Pitch Measurements
Optical digital scale of 1 m long (10 mm resolution).
Scale has USB output that allowed to record data
automatically.
Wire pitch accuracy was provided by
winding machine (no comb used).
The measured pitch (RMS) was ~0.11mm

## Slide 8
Wire Tension Measurements
Tension measurements was performed with using magnet, switch box, pulse
generator and USB camera. It is required a wire connection with both side.
A setup is simple but measurements are labor intensive.
Automatic multichannel tension measuring device is needed.

## Slide 9
Chart: U&V wire length
Tension Measurements
Chart: X-Plane (B-side)
Chart: A-side, V-plane
Visible wire tension structure demonstrates a tension variation for different wire modules.
A reason for that was connected with insufficient stiffness of transfer frame and stability
of wire tensioner of the winding machine.

## Slide 10
    Capacitance measurements
    HV tests
Electrical tests

## Slide 11
Capacitance Measurements
All other planes grounded
Switch box
LCR meter
1
2
3
4
X-plane
Wire group 1
Wire group 2
Wire group 3
Wire group 4
1
2
3
4
U &V planes
Wire group 4
Wire group 3
Wire group 2
Wire group 1
Wire group 1
Wire group 2
Wire group 3
Wire group 4
Wire length:
  Mesh & X-wires - 2.1 m
  U & V wires (zig-zag) - 4.7 m
Automatic multichannel capacitance measuring device will be very useful.

## Slide 12
Chart: X-Wire capacitance
X-Wire Capacitance Measurements
Average capacitance is about 11 pF  5.2 pF/m

## Slide 13
V-Plane Wire Capacitance Measurements
Chart: V- Wire Capacitance
| 1.0 | 2.0 | 3.0 | 4.0 | 5.0 | 6.0 | 7.0 | 8.0 | 9.0 | 10.0 | 11.0 | 12.0 | 13.0 | 14.0 | 15.0 | 16.0 | 17.0 | 18.0 | 19.0 | 20.0 | 21.0 | 22.0 | 23.0 | 24.0 | 25.0 | 26.0 | 27.0 | 28.0 | 29.0 | 30.0 | 31.0 | 32.0 | 33.0 | 34.0 | 35.0 | 36.0 | 37.0 | 38.0 | 39.0 | 40.0 | 41.0 | 42.0 | 43.0 | 44.0 | 45.0 | 46.0 | 47.0 | 48.0 | 49.0 | 50.0 | 51.0 | 52.0 | 53.0 | 54.0 | 55.0 | 56.0 | 57.0 | 58.0 | 59.0 | 60.0 | 61.0 | 62.0 | 63.0 | 64.0 | 65.0 | 66.0 | 67.0 | 68.0 | 69.0 | 70.0 | 71.0 | 72.0 | 73.0 | 74.0 | 75.0 | 76.0 | 77.0 | 78.0 | 79.0 | 80.0 | 81.0 | 82.0 | 83.0 | 84.0 | 85.0 | 86.0 | 87.0 | 88.0 | 89.0 | 90.0 | 91.0 | 92.0 | 93.0 | 94.0 | 95.0 | 96.0 | 97.0 | 98.0 | 99.0 | 100.0 | 101.0 | 102.0 | 103.0 | 104.0 | 105.0 | 106.0 | 107.0 | 108.0 | 109.0 | 110.0 | 111.0 | 112.0 | 113.0 | 114.0 | 115.0 | 116.0 | 117.0 | 118.0 | 119.0 | 120.0 | 121.0 | 122.0 | 123.0 | 124.0 | 125.0 | 126.0 | 127.0 | 128.0 | 129.0 | 130.0 | 131.0 | 132.0 | 133.0 | 134.0 | 135.0 | 136.0 | 137.0 | 138.0 | 139.0 | 140.0 | 141.0 | 142.0 | 143.0 | 144.0 | 145.0 | 146.0 | 147.0 | 148.0 | 149.0 | 150.0 | 151.0 | 152.0 | 153.0 | 154.0 | 155.0 | 156.0 | 157.0 | 158.0 | 159.0 | 160.0 | 161.0 | 162.0 | 163.0 | 164.0 | 165.0 | 166.0 | 167.0 | 168.0 | 169.0 | 170.0 | 171.0 | 172.0 | 173.0 | 174.0 | 175.0 | 176.0 | 177.0 | 178.0 | 179.0 | 180.0 | 181.0 | 182.0 | 183.0 | 184.0 | 185.0 | 186.0 | 187.0 | 188.0 | 189.0 | 190.0 | 191.0 | 192.0
 | 16.78 | 27.78 | 30.03 | 30.45 | 30.79 | 28.88 | 29.18 | 31.04 | 31.87 | 32.32 | 31.92 | 29.48 | 29.53 | 32.14 | 32.68 | 32.22 | 32.1 | 29.74 | 29.59 | 31.43 | 31.21 | 31.68 | 31.49 | 28.8 | 28.92 | 31.71 | 32.06 | 31.42 | 31.66 | 29.5 | 29.27 | 31.12 | 31.49 | 31.68 | 31.38 | 28.89 | 28.82 | 31.35 | 31.83 | 31.4 | 31.14 | 28.9 | 28.62 | 30.47 | 30.91 | 31.43 | 31.48 | 27.38
 | 31.89 | 38.85 | 39.21 | 39.24 | 39.57 | 35.28 | 35.34 | 39.5 | 40.07 | 40.35 | 39.89 | 34.9 | 34.9 | 39.84 | 40.56 | 40.5 | 40.26 | 35.66 | 35.63 | 39.63 | 39.75 | 40.19 | 39.82 | 34.74 | 34.75 | 39.89 | 40.33 | 40.02 | 39.96 | 35.68 | 35.64 | 39.97 | 40.05 | 40.27 | 39.6 | 34.61 | 34.88 | 39.93 | 40.52 | 40.04 | 39.64 | 34.77 | 34.92 | 38.97 | 39.32 | 39.9 | 39.85 | 33.22
 | 33.23 | 39.39 | 39.82 | 39.81 | 39.81 | 35.82 | 35.86 | 40 | 40.39 | 40.9 | 40.12 | 35.01 | 34.92 | 40.21 | 40.77 | 40.34 | 40.38 | 35.97 | 36.28 | 40.33 | 40.21 | 40.7 | 40.4 | 35.25 | 35.19 | 40.43 | 41.17 | 40.45 | 40.53 | 36.2 | 35.78 | 39.99 | 40.55 | 40.75 | 40.05 | 35.04 | 34.9 | 40.08 | 40.64 | 40 | 40.12 | 35.45 | 35 | 38.77 | 39.15 | 39.52 | 39.49 | 32.78
 | 32.32 | 38.73 | 39.28 | 39.33 | 39.64 | 35.22 | 34.97 | 39.22 | 39.76 | 40.03 | 39.78 | 34.78 | 34.67 | 39.76 | 40.32 | 40.05 | 40.1 | 35.77 | 35.53 | 39.52 | 39.62 | 40.13 | 39.92 | 34.83 | 34.84 | 40 | 40.63 | 40.17 | 40.65 | 36.16 | 35.76 | 39.87 | 39.85 | 39.97 | 39.61 | 34.62 | 34.3 | 39.56 | 40.14 | 39.7 | 39.16 | 34.88 | 34.78 | 38.72 | 37.69 | 37.34 | 35.34 | 20.99
One of the wire group has lower capacitance.
A reason for that is a different number of FR-4- boards
and jumpers.
Wire groups 1-3 have 3 jumpers but group 4 only 2.
Average capacitance (1-3 groups) is about 37 pF  7.9 pF/m
1
2
3
4
Wire group 4
Wire group 3
Wire group 2
Wire group 1
Wire group 1
Wire group 2
Wire group 3
Wire group 4
Wire jumpers

## Slide 14
APA High Voltage Tests
Reasons for HV current leakage:
- from wire boards artwork to pins and frame
- from board to boards
Pins
FR-4 pc-board
Holes for pins
- 355V (U)
0 V (V)
490 ( X)
0 V (Mesh)
Frame and pins are grounded
Bias voltages
Distance from hole edge to
board artwork is 2.0 mm

## Slide 15
Setup for High Voltage Tests

## Slide 16
High Voltage Leak Measurements
Chart: HV leak current to pins (frame and pins grounded only)
| 0.2 | 0.4 | 0.6 | 0.8 | 1.0 | 1.2 | 1.4 | 1.6 | 1.8 | 2.0 | 2.2 | 2.4 | 2.6 | 2.8 | 3.0 | 3.2 | 3.4 | 3.5
Mesh | 2 | 4 | 4 | 4 | 6 | 10 | 18 | 36 | 80 | 150 | 220 | 310 | 410 | 580 | 660 | 850 | 1100 | 1400
X | 2 | 4 | 4 | 6 | 8 | 10 | 12 | 14 | 17 | 20 | 26 | 36 | 56 | 90 | 120 | 160 | 210 | 220
U | 2 | 4 | 4 | 8 | 8 | 10 | 11 | 12 | 15 | 17 | 22 | 30 | 44 | 66 | 105 | 160 | 260 | 300
V | 2 | 5 | 6 | 7 | 9 | 10 | 11 | 12 | 13 | 15 | 18 | 20 | 26 | 34 | 38 | 48 | 100
Chart: HV leak current (nearby planes grounded)
| 0.2 | 0.4 | 0.600000000000001 | 0.8 | 1.0 | 1.2 | 1.4 | 1.6 | 1.8 | 2.0 | 2.2 | 2.4 | 2.6 | 2.8 | 3.0
Mesh | 2 | 3 | 6 | 7 | 9 | 12 | 22 | 52 | 120 | 220 | 380 | 600 | 1200 | 2600 | 3800
X | 2 | 6 | 7 | 9 | 11 | 13 | 19 | 42 | 170 | 290 | 380 | 720 | 1300 | 2400 | 3400
U | 3 | 7 | 7 | 8 | 11 | 14 | 16 | 24 | 42 | 110 | 220 | 350 | 570 | 950 | 3000
V | 2 | 6 | 8 | 9 | 10 | 12 | 14 | 18 | 21 | 36 | 80 | 110 | 220 | 330 | 460
Frame and pins grounded only
Up to 3.5 kV a leak current is less
than 200 nA except Mesh because .
artwork traces are slightly closely
to pin holes.
Frame and pins + nearby planes grounded
A current is higher but still is very
small (~100 nA) at voltages up to 2 kV.
Expectation, that a current in liquid nitrogen will be less.

## Slide 17
    Cryostat construction
    Cooling down procedure
    HV tests
Cold Tests in Liquid Nitrogen

## Slide 18
Cryostat Preparation for APA Cold Tests
Simple cryostat with foam isolation installed into the wooden crate.
Nylon film was placed inside foam box to prevent leak of liquid nitrogen.
A cryostat cost (M&S) ~$200 , labor (fabrication and tests) ~20 hours (tech)

## Slide 19
APA Transportation and Loading
Additional mechanical test on robustness and reliability

## Slide 20
Starting Cold Tests
Liquid nitrogen was supplied from 4 Dewars, simultaneously.
The hoses were mounted at cryostat corners.
. Filling time was about 1 h 10 min (6 Dewars of 160 liter each.}

## Slide 21
Cryostat Opening
After APA inspection no broken wires, epoxy cracks,
damaged connectors or frame deformation were found.
Double click on the picture to watch movie

## Slide 22
HV tests in Liquid Nitrogen
Chart: Voltage, kV
| 0.5 | 1.0 | 1.5 | 2.0 | 2.5
Mesh | 6 | 10 | 16 | 66 | 100
X | 8 | 12 | 260 | 160 | 1800
U | 18 | 30 | 70 | 200
V | 50 | 80 | 160 | 420
Chart: Voltage, kV
| 0.5 | 1.0 | 1.5 | 2.0 | 2.5 | 3.0
Mesh | 6 | 6 | 6 | 8 | 8 | 10
X | 10 | 40 | 7000
Measurements before filling
Measurements 10 min later
Mesh, X, U and V planes for both sides were connected together
After about 10-15 min a current for all planes dramatically increased for all planes.
It is indicated a water appeared and condensed on the FR-4 surface during of fast cooling.
A reason for that may be connected with high level of water absorption of FR-4 boards.
No any reference on HV leakage in FR-4 boards in liquid nitrogen was found on that score.
Need to investigate this problem to understand cool down procedure for TPC.

## Slide 23
Plans
1. Investigation a reason of high current in liquid nitrogen.
- small setup for HV test with boards, stud,….
- different mode of cool down procedure (gas flashing/liquid nitrogen)
- analysis of water contamination in FR-4
- developing a procedure for cooling down
2. Wire tension measurements in liquid nitrogen
- procedure of tension measurements
- setup with X-wire module (64 wires)
3 Integration APA prototype with front-end electronics
- decouple pc-boards, capacitors, connectors
- board connection to APA and readout electronics