## Slide 1: ICEBERG: Status and Plan
Shekhar Mishra
Neutrino Division, Fermilab
06/26/2024

## Slide 2: ICEBERG: Study Plan
Measure Individual Channel Linearity (INL) using pulser data. (Data Taken)
Determine the precision of individual channel gain (LArASIC + ColdADC; e-/LSB) using pulser data. (Data Taken)
Measure and decide the optimal gain setting: 14 mV/fC or 7.8 mV/fC or 25 mV/fC(Data Taken)
Determine optimal baseline for collection plane: 200+ mV (Done)
Study shaping time effect on ENC (1 msec vs 2 msec vs 3 usec) (Data Taken)
Measure cross talks. (14 mV/fC, DAC 15, single and diff pulse) DAC 0 – 63 (Data Taken)
Determine optimal ColdADC Vref settings. (Done)
Learn to use particle interaction in LAr to determine absolute calibration of the TPC + Electronics response (MeV/LSB-tick)
  MIP dE/dX
  Michel electrons
  Ar-39
DUNE-FD2-PD: Integration of DAPHNE-V3 in Ethernet DUNE-DAQ
  Spy-Buffer (Data Taken and Analyzed)
  Starting on FELIX DAQ

## Slide 3: Calibration Run Summary
There are more than 2500 runs to be analyzed
Original Setting (reg24 = 223, reg25 = 51, reg 26 = 137) (Run 12076-12317)
  14 mV/fC, 2 usec, 900/200 mV (Run 12221 – 12253)
Vref 1.8 (reg24 = 240, reg25 = 34, reg 26 = 137) (Run 12471-12967)
  14 mV/fC, 2 usec, 900/200 mV (Run 12741 – 12804)
Vref 1.95 (reg24 = 250, reg25 = 24, reg 26 = 137) (Runs 13077 - 13572)
  14 mV/fC, 2 usec, 900/200 mV (Run 13346 – 13409)
Vref 1.9 (reg24 = 243, reg25 = 27, reg 26 = 135)
  14 mV/fC, 2 usec, 900/200 mV (Run )
  Single Channel Pulse

## Slide 4: Optimal ColdADC Vref settings
Goal: “Expand ColdADC dynamic range so that it does not limit the dynamic range of the channel (make it slightly larger than the LArASIC dynamic range)”
Study: Adjust Vcmi , Vcmo , Vrefp , and Vrefn
Summary: Adjust only Vcmi ,Vrefp , and Vrefn
  coldadc_settings": { "reg_0": 35, "reg_24": 243, (223) "reg_25": 27, (51) "reg_26": 135, (137) "reg_27": 103, "reg_29": 39, "reg_30": 39, "reg_4": 59 },

## Slide 5: Vrefp = 223, Vrefn=51, Vcmo = 137
15206
25 mV/fC, DAC = 63
Run adc stats: 15206 0 0 16383 0 0 0 16383 0 Events: 61

## Slide 6: Vrefp = 223, Vrefn=51, Vcmo = 137
Mean = 954

## Slide 7: Original, 14 mV/fc, DAC = 32

## Slide 8: Original Calibration Scan: 14 mV/fC
DAC = 31
DAC =31

## Slide 9: Vrefp = 245, Vrefn=29, Vcmo = 137
14171
25 mV/fC, DAC = 63

## Slide 10: Vrefp = 241, Vrefn=29, Vcmo = 135
14173
25 mV/fC, DAC = 63

## Slide 11: Vrefp = 243, Vrefn=27, Vcmo = 135
14174
Recommended Setting
25 mV/fC, DAC = 63

## Slide 12: Vrefp = 243, Vrefn=27, Vcmo = 135
14174
25 mV/fC, DAC = 63

## Slide 13: Vref = 1.8, 14 mV/fC
DAC = 31
DAC =18

## Slide 14: Vref = 1.95, 14 mV/fC
DAC = 31
DAC =18

## Slide 15: Vref = 1.8, 25 mV/fC

## Slide 16: Vref = 1.95, 25 mV/fC
DAC = 17
DAC =10

## Slide 17: Vref = ~1.9, 14 mV/fC
DAC = 30
DAC =17
Recommended Setting

## Slide 18: Vref = ~1.9, 25 mV/fC
DAC = 17
DAC =10
Recommended Setting

## Slide 19: Cross Talk

## Slide 20: Cross Talk (in time with Negative)

## Slide 21: Analysis in progress
Cross Talk (single and Differential)
Recommended Vref (Differential)