## Slide 1
Application of Detector in Medicine
Hyejoo Kang, PhD
Associate Professor
Loyola University Chicago
Loyola Medical Center
11-21-2024, EDIT conference, fermilab

## Slide 2
Radiation therapy medical physicist
Formally, high energy experimental physicist
  Stanford linear accelerator center : SLD experiment working on QCD
  Fermi national laboratory: Neutrino oscillation experiment (MINOS)
Re-trained as a medical physicist at Memorial Sloan-Kettering Cancer Center (MSKCC) in NYC
Currently, faculty member at Loyola University of Chicago
Associate director of medical physics at radiation oncology department
Clinical work, teaching/mentoring, research and outreach
Introducing myself

## Slide 3
INTRODUCTION OF MEDICAL PHYSICS
INTRODUCTION OF RADIATION THERAPY
DETECTORS & APPLICATIONS IN RADIATION THERAPY
DETECTORS & APPLICATIONS IN MEDICAL IMAGING
CUTTING-EDGE DEVELOPMENT IN MEDICAL PHYSICS
OUTLINE

## Slide 4
Far detector @ Soudan: 8 meter diameter, 30 meter length; steel & scintillator & PMT
Minos (Main injector neutrino oscillation search) (2005-2016)
The bigger, the better
Near detector @ Fermilab: 3 meter fiducial region steel & solid scintillator & Photomultiplier tube (PMT) detectors
457 mi

## Slide 5
8 meters
Large complex hadron detectors: Fermilab Tevatron
High energy experiments looking for new particles, discovered top quark & searched for Higgs boson
The state-of-the art accelerator & detectors
CDF
D0
(1983-2011)

## Slide 6
Stanford linear accelerator center (slac)
The longer, the better
2 mi (3219 meters)
The Stanford Linear Detector (SLD) experiment to study Z-boson, quark and lepton, Quantum Chromodynamics (QCD) through e- & e+ collision at Z0-pole (1992-1998)

## Slide 7
Medical linear accelerator is 1 m
The shorter, the better (medical linear accelerator)
Chamber detector
1m long linear accelerator
Unit of shaping radiation

## Slide 8
What is medical physics anyway?

## Slide 9
Medical physics: modern branch of physics concerned with the application of the concepts and methods of physics to the diagnosis and treatment of human disease and is allied with health physics (Newhauser et al, JACMP, 2023)
Radiation therapy, imaging, nuclear medicine, radiation protection (health physics)
The mission of a professional organization:
  Advancing medicine through excellence in the science, education and professional practice of medical physics.
Critical thinking scientist
Educator/teacher/mentor
New technology expert
Problem solving champion
Medical physicists: Scientists in patient care

## Slide 10
Radiation treatment uses various forms of radiation to treat cancer and other diseases
  External beam radiation therapy: using photons, electrons, protons, carbons from accelerator to treat diseases
  Internal radiation therapy: using radioactive materials to treat diseases
Introduction of Radiation Therapy/Treatment
Often in conjunction with other treatment types
Approximately, two-third of cancer patients receive radiation as part of treatment
Need medical imaging within the department and from radiology department

## Slide 11
Radiation treatment machines @Loyola Medical Center
Linear accelerator with X-ray imaging: External beam treatment
Linear accelerator with advanced imaging and delivery
High dose rate radioactive material internal beam treatment machine

## Slide 12
When patient comes for Radiation Treatment
Doctor consultation
Learn about radiation treatment. Give permission for treatment
Simulation Imaging
Imaging scans to determine exact area to treat. Crucial step of treatment design
Treatment planning
Design treatment to achieve the treatment goals. Perform highly technical calculation
Quality Assurance (QA)
Treatment Delivery
Physicist assess the quality of plan via verifying & measuring prior to treatment
Make sure each treatment is accurately & safely delivered

## Slide 13
Detectors for medical physics application

## Slide 14
Use to measure radiation amount, a.k.a dose, energy per mass (J/kg)
  Radiation dose in experiment settings & patient (In-vivo) dosimetry
  RT machine calibration and Quality Assurance
  Ensure that machine delivers the correct radiation dose within specifications
  Monitor the characters and properties of radiation beam
  Help maintain the accurate treatment delivery
Imaging detectors
  Flat panel detectors to capture patient image to make sure radiation is delivered to cancer tumors
Radiation protection & safety
  Area monitors: detectors measure the radiation levels in the environment
  Personal badge: monitor radiation levels for radiation works or patient
Radiation therapy (RT) detectors
The smaller, the better. The simpler, the better

## Slide 15
Collected charges is proportional to the radiation dose
Simple and durable; bread&butter detector
Versatile: measure both low & high radiation levels and wide range of radiation types & energy
Real time measurement, but relatively large volume
Cables and other auxiliary devices are needed
Gas- or air-filled ionization chambers
Thimble type ion chamber
Cathode & anode (electrodes) to collect charges or current from electron generated from radiation ionization. Signals are measured using electrometer

## Slide 16
Dump the detector into a big tub of water to measure radiation from the machine and scattering radiation from water
Why? We approximate human to a tub of water and it works for most of the times
What do we do with them?
Luke Laymen, MS

## Slide 17
Medical linear accelerator is 1 m
Dose, symmetry, flatness….
Monitor radiation dose and characteristics
Chamber detector

## Slide 18
Types:
  Thermoluminescent dosimeter (TLD): LiF:Mg,Ti
  Optically stimulated luminescent dosimeter (OSLD): Al2O3:C
    TLD & OSLD: Photomultiplier tube (PMT) to collect luminescent signals
  Diode: Semiconductor detectors (Silicon or Germanium)
  Diamond detectors (insulator)
  Scintillator detectors
Small size, high sensitivity, fast response time, no cables or electrometer
Wait time for radiation reading, not real-time
Solid State Detectors
Notes: LiF:Mg, Ti: Lithium Fluoride doped with Magnesium (Mg) and Titanium (Ti)
AlxO3:C: Carbon-doped aluminum oxide

## Slide 19
Solid state physics:
  Conduction band
  Valence band
  Band gap
Electrons & holes; electron can gain energy to move up to conduction band
Solid state detector principle (conductor, semi-conductor & insulator)
Silicon, Germanium
Diamond, Glass.…

## Slide 20
  Radiation ionization creates electron-hole pairs
  Electrons move to conduction band and gets trapped in band gap (defects) and holes get trapped in defects
  Heat/light stimulation: When stimulated by heat/light, trapped electrons gain sufficient energy to move back to the conduction band, then they can recombine with trapped holes.
  This recombination process emits light, a phenomenon known as luminescence.
  The intensity of the emitted luminescence is proportional to the amount of radiation absorbed by the detector
Semiconductor to emit luminescence light (TLD & OSLD)
Energy level diagram of LD
Energy

## Slide 21
Useful for light detection of very weak signals, is a photoemissive device in which the absorption of a photon results in the emission of an electron. These detectors work by amplifying the electrons generated by a photocathode exposed to a photon flux.
PMTs acquire light through a glass or quartz window that covers a photosensitive surface, called a photocathode, which then releases electrons that are multiplied by electrodes known as metal channel dynodes. At the end of the dynode chain is an anode or collection electrode. Over a very large range, the current flowing from the anode to ground is directly proportional to the photoelectron flux generated by the photocathode.
Photomultiplier Tube (PMT)

## Slide 22
TLD/OSLD: readers receive signals using PMT
HEATING
LIGHT

## Slide 23
Convenient: no cable or electrometer
Small to measure radiation on localized areas
Used to measure entering radiation dose on patient (in-vivo dosimetry)
Characterize radiation treatment beams
Larger measurement uncertainty
Need reading devices
Why TLD & OSLD?
T
Total skin electron beam therapy for cutaneous T-cell lymphoma
Skin radiation dose for breast cancer patients

## Slide 24
Basic physics of diode detectors
A p-n junction diode is formed when a p-type semiconductor is fused to an n-type semiconductor creating a potential barrier voltage across the diode junction
Schematic of a silicon p-n junction diode as a radiation detector:
The excess minority carriers  electron and ○ hole generated by radiation within one diffusion length, Lp & Ln are able to diffuse to the depleted region
They are swept across the junction by the built-in potential 0 and are collected by the electrometer.
Built-in potential 0.3 ~ 1V

## Slide 25
Versatile detectors:
  Can be used to measure radiation dose and characteristics
  Smaller sized detector makes it useful in radiation therapy
Need to consider many dependence factors for radiation dose measurement
Bigger uncertainties compared to ion chamber detector
Why diode detectors?

## Slide 26
RT application: flat panel x-ray detector
  Advanced development: non x-ray images
Radiology & diagnostic imaging
  Computer Tomography (CT)
  Positron Emission Tomography (PET)
  Magnetic Resonance Image (MRI)
  Ultrasound (US)
Medical imaging detectors

## Slide 27
Advances in RT techniques and the use of imaging over time
G
García-Figueiras R. Published Online: January 11, 2024
https://doi.org/10.1148/rg.230152

## Slide 28
Measure the x-ray intensity transmitted through a patient from a radiation source
Radiation signal is converted electronically into a 2D digital radiographic image to verify the correct beam placement w.r.t the patient’s anatomy
Electronic Portal Imaging Detector (EPID):flat-panel detector
EPID
EPID

## Slide 29
Photodiodes (a-Si) attached to TFT as switch
> 1 MeV photons
EPID (flat panel detector): patient positioning & dose distribution model
amorphous silicon photodiodes
2 or 3cm thickness
Cesium iodide (CsI) or Gadolinium Oxysulfide (GOS)
Notes: X-ray photons enter the EPID and interact with the scintillator layer (e.g., cesium iodide or gadolinium oxysulfide). The scintillator converts X-rays into visible light photons. These visible light photons then interact with the amorphous silicon (a-Si) photodiode array, generating electron-hole pairs in the silicon. The liberated electrons are collected in the pixelized detector array, with each pixel accumulating charge proportional to the intensity of the light received. The collected charge is read out one row at a time using thin-film transistor (TFT) switches integrated into the detector. The output signal is amplified, digitized, and sent to the processing unit for image reconstruction.

## Slide 30
Breast cancer patients: showing both soft-tissues and bony anatomy
EPID images
Low metal imaging artifacts
Can be real-time using treatment beams
Poor soft tissue contrast
Increased imaging dose
MeV EPID image
KeV image
Blurry bony anatomy
Breast tissues

## Slide 31
See through internal anatomy by detecting x-rays passing through the body
Create detailed cross-sectional images by reconstructing x-ray maps from multiple angles
Siemens SOMATOM Force
:Dual x-ray sources, AI
Fan beam
Computer tomography (CT) imaging: gold-standard imaging for RT
Notes: CT scans work by detecting X-rays that pass through the body from multiple angles. The X-ray source and detectors are positioned on opposite sides of the patient, and the system rotates around the patient while they move through the scanner. These X-rays create detailed cross-sectional images of the body, allowing doctors to see the internal structures of organs and tissues in 3D.

## Slide 32
Radiation Therapy (RT):
  Treatment designing/planning
  Tumor localization
  Patient positioning verification
  Adaptive radiation therapy
  Monitoring tumor response to RT
Diagnostic Imaging:
  Diagnosing conditions like tumors, trauma, and internal bleeding, especially in the brain, chest, and abdomen
  CT angiography for vascular imaging, virtual colonoscopy for cancer screening, and cardiac and spinal imaging for disease assessment.
CT imaging
Essentially, almost everything we do!
Notes: Diagnostic imaging is a medical technique used to look inside the body without surgery, helping doctors see what’s going on with organs, bones, and tissues.
Additionally, special CT scans, such as CT angiography can be used for vascular imaging, which visualize and evaluate blood vessels and their functionality

## Slide 33
Maintaining detector alignment, absorbing heats and ensuring electrical isolation
The detector is responsible for capturing the X-rays produced by the CT tube and converting them into an electrical signal
Detectors for CT imaging
< 120 kev photons

## Slide 34
Solid-state crystal detectors, consists of photodiodes that detect gamma rays and convert them into digital signal. Image data is then sent to Data Acquisition Subsystem (DAS) for collection and processing
CT scanner detector

## Slide 35
Data stored as raw detector signal output, sinograms, angle specific histrograms of detected events.
Images can be reconstructed using analytic or iterative reconstruction (e.g. back projection)
Reconstructed 3D images can be rendered and displayed in many useful ways including locating cancerous cells
CT scans are used for treatment designing for radiation treatment
CT images
https://humanhealth.iaea.org/HHW/MedicalPhysics/NuclearMedicine/ImageAnalysis/3Dimagereconstruction/index.html

## Slide 36
Example of (A) cancerous cell volumes and (B) treatment design/plan with radiation dose levels for a patient treated with post-operative cavity RT
CT image examples
Diagnostic cerebral CT-angiogram (A) showing a large tangle of abnormal vessels (B) 3D volumetric CT scan of the vessel location
Radiology CT scans: diagnosing & locating disease
Radiation therapy CT scans: defining disease and designing RT
brainstem
optical nerves
eye

## Slide 37
Energy integrating detector (EID): Xray photons are absorbed in a scintillator ceramic (GOS) & generate light which is then collected by photodiodes. The collected signal is proportional to the incoming X-ray intensity
Energy integrated vs. photon counting detector
Photon counting detector (PCD): X-ray photons are absorbed in a semiconductor crystal, generating electron charge clouds directly, inducing electrical pulses at the pixel electrodes. These electrical pulses are analyzed to detect individual photon interactions.
Notes: For example, Cadmium Telluride (CdTe) or Cadmium Zinc Telluride (CZT)

## Slide 38
Why photon counting detector (PCD)?
With “color” information, can have better visualization between bone/soft tissues/calcium..
Adding “color” information to provide more details of images

## Slide 39
PCD and EID CT of vascular reconstruction kernels in photon-counting CT angiographies of the leg
  More detailed substructures are shown
More applications
PCD images
Energy integrated detector (EID)
Photon counting detector (PCD)
Notes: These are CT angiographic images of the leg, obtained using different thresholds in blood volume (BV) measurement.

## Slide 40
Shows location/strength of metabolism vs. anatomy
  Disorders or cancer cells with high metabolic activity have elevated glucose metabolism and increased uptake of radioactivity from the glucose analogue
  Improve disease detection: brain disorders, heart or cancer
  Provides metabolic or biochemical function information
  Positron Emission Tomography (PET) & Single-photon Emission Computed Tomography (SPECT)
Positron Emission Tomography (PET)
  18F-FDG used for glucose metabolism detection
Functional and molecular imaging
Notes: F-FDG: Fluorine-18 Fluorodeoxyglucose

## Slide 41
Full ring of scintillating detector surrounding patient
Positron Emission Tomography (PET)
Notes: A Positron Emission Tomography (PET) scan works by detecting a process called positron-electron annihilation. In this process, radioactive atoms attached to a glucose molecule are injected into the body. When the radioactive atoms decay, they emit positrons, which interact with electrons in the body, resulting in the release of two photons traveling in opposite directions. These photons, each with an energy of 511 keV, are detected by the PET scanner. By measuring the coincident detection of these photons, PET scans can pinpoint the location where the annihilation took place, allowing doctors to visualize areas of high radioactivity, such as tumors or abnormal tissue

## Slide 42
Positron Emission Tomography (PET)
Basic physics
  Radiotracers undergo β+ decay, emitting e+ (positron) for neutron-deficient or proton-rich
  e+ travels in tissues and loses energy & comes to rest
  e+ annihilates with atomic e-
  Emits two back-to-back photons (γ-rays)
  Detect the photons utilizing time coincidence
Tomographic techniques that uses radiotracers administered to the patient: emission imaging

## Slide 43
Annihilation coincidence: two photons detected by opposite detectors within coincidence window of 5-10 ns along the line of the two detectors
Annihilation Coincidence Detection


Coincidence ?
Register
Event
Yes
5-10 ns

## Slide 44
Positron Emission Tomography (PET) images
For a PET acquisition, a patient is placed within a ring of detectors and sinograms are created.
Emission profiles per angle can be generated by mapping and sorting the detector pair events

## Slide 45
PET Image
Hounsfield Unit/CT number (HU)
Standardized Uptake Values (SUV)

## Slide 46
PET detectors
Gamma ray detector blocks
Thermal barrier

## Slide 47
A 511 keV photons enters scintillator to interact with a dense scintillator material (e.g.BGO) generated optical photons (around 1 eV)
Optical photons are detected by an array of photomultiplier tubes (PMTs). PMTs amplify the signal significantly
Measure current pulse for each detected photon
Scintillation detectors for PET
PET scintillator:
Sodium Iodide (NaI)
Lutetium oxyorthosilicate (LSO)
Lutetium Yttrium Orthosilicate (LYSO)
Bismuth Germanate Oxide (BGO)
Gadolinium Gallium Garnet (GaGG)
Cerium-doped Bromide (CeBr)

## Slide 48
Photomultiplier tube (PMT)
PET photodetectors
Silicon photomultiplier (SiPM)
SiPM
Array of photosensitive single-photon avalanche photodiodes (SPAD)
SPAD: operated in Geiger mode. Single photon causes charge avalanche
Siemens PET (Biograph Vision PET/CT) with SiPM and a TimeOfFlight: 214 ps
PET detector block (Siemens, Optiso UDR): 3.2mm LSO crystal arrays with SiPM sensors in a detector mini-block
SiPM
LSO crystal
Detector mini-block
Block
https://www.siemens-healthineers.com/en-us/molecular-imaging/pet-ct/biograph-vision-quadra/udr-detector

## Slide 49
CT+PET
Two gantries combined in same housing
Multislice CT scanner mated to a PET scanner
Improved attenuation correction for PET
Measure anatomy+functionality at once
Possibly three scans acquired
  Attenuation correction CT
  PET
  Treatment planning CT, with contrast if necessary

## Slide 50
Superior soft tissue distinguishability
Non-invasive and painless
Functional Imaging: Provide both anatomy and metabolism information
Can be long scan time
Magnetic Resonance Image (MRI)
CT
T1 MRI
T2 MRI

## Slide 51
https://mriquestions.com/receive-only-coils.html
MRI RF detector
https://www.nature.com/articles/s41598-019-54200-3
Signals are detected by the MRI receiver coils
The frequency and timing of the signals depend on the tissues

## Slide 52
Estimate blood flow
Ultrasound: piezoelectric crystal transducer

## Slide 53
Diagnosing FIGO Stage IVB high-grade serous carcinoma of the ovary using multiple image types
We need multiple imaging to identify disease more accurately
Multi-modality brain imaging includes different MRI and PET types

## Slide 54
Radiation therapy machine + non x-ray imaging unit
Monitoring cancer target during the treatment
Fine-tune radiation treatment plan to daily changes
Goal: Increase therapeutic outcome
State of the art in radiation therapy
Machine+PET
Machine+MRI

## Slide 55
Biologically Guided RT: PET signal, metabolic response or tumor guides RT, “functional imaging based RT”
Tumor-adaptive RT planning
Real-time tumor tracking
Reflexion: RT machine + PET detector

## Slide 56
Limited-time-sample (LTS) PET image to track tumor: 100ms per image
FULL: Sum of the derived partial sample of each LTS, ~ 500 ms in 60 RPM (half rotation)
Signals for PET imaging
BTZ: biology-tracking zone.
Shirvani et al, 2021, BJR

## Slide 57
Accelerator is mounted on a ring gantry
Superconduction coils are repositioned to create a low field toroid and to create a beam portal
Gradient coil is split to allow beam passage
MRI image guided RT
Accelerator is mounted on a ring gantry.
It is sandwiched between two separate MRI units

## Slide 58
RT machine with MRI attached
ViewRay MRIdian: MRI scanner + Treatment machine
MRI
3D CT

## Slide 59
Blue: tumor
Orange: range of tumor excursion
Automatically Radiation beam on/off
Showing tumor motions
Loyola University chicago
MRI real-time tracking of tumors in the abdomen

## Slide 60
Detectors are essential in medical field
  Radiation dose measurement
  Quality assurance & safety
  Radiation dose monitoring
  Essential components in medical images
  Radiation protection & safety
Many of these detectors had been developed in other physics fields, high-energy particle physics
Simple & robust
Adopt newest and latest detector technologies
Clinical validation is must
Summary

## Slide 61
Acknowledgement
Luke Layman, MS
Yussuf Abdelal, MS
Sebastien Gros, PhD
Alexander R. Podgorsak, PhD (University of Rochester)
Andrew Godley (UT Southwestern)
Loyola

## Slide 62
THANK you

## Slide 63
Backup slides

## Slide 64
TLD and OSLD

## Slide 65
A PN-junction diode is formed when a p-type semiconductor is fused to an n-type semiconductor creating a potential barrier voltage across the diode junction
Diode Principle
Zero bias

## Slide 66
Diodes in RT
A PN-junction diode is formed when a p-type semiconductor is fused to an n-type semiconductor creating a potential barrier voltage across the diode junction

## Slide 67
Basic comparison of energy-integrated detector (EID) (a) and photon-counting detector computed tomography (PCT) (b). EID systems use solid-state scintillation detectors where X-rays are first converted to light, then to an electrical signal, but this method is limited by electronic noise and has a finite spatial resolution. PCT systems, in contrast, directly convert X-rays into electrical signals in a semiconductor, offering improved dose efficiency and spatial resolution by eliminating the need for septa between pixels and allowing for sub-pixel readouts [30]
Energy-integrating detectors
  convert incoming photons to visible light, directly followed by a conversion to an electrical signal by photodiodes
  all measured photons in one projection are integrated to create the signal
  detector’s spatial resolution depends on the use of reflecting septa to prevent cross-talk
Photon-counting detectors
  create an electron-hole pair when a photon strikes the semiconductor material of the detector
  the anode attracts the electrons from the electron-hole-pair
  electric pulse signal is proportional to the photon’s energy
Photon counting detector (PCD)
Energy-integrating detectors: commonly used
Photon-counting detectors: advanced

## Slide 68
PCCT-scan of a Patient with endovascular aortic repair. A/B. Virtual-mono-energetic images (VMIs) can be created without additional scanning from the spectral dataset to generate images for virtual photons at a single kiloelectronvolt (keV) energy level. This can be helpful to improve differentiation between tissues, visualization of contrast-enhanced tissues, or reducing artifacts. C. Virtual-non-contrast (VNC) reconstructions remove the signal of iodinated contrast agents from contrast-enhanced scans, with the idea to potentially make non-enhanced scans obsolete. D. Iodine maps are useful to measure perfusion but hinge on their quantification accuracy [2]. E. It is possible to fuse VNC and iodine maps for additional information.
Photon counting detector (PCD)

## Slide 69
Figure 1: Diagrams show detector types. A, In conventional energy-integrating detector, an incident x-ray photon is converted into a shower of visible light photons in a scintillator. Visible light hits an underlying light sensor, where it generates positive and negative electrical charges. B, In photon-counting detector, the x-ray photon is absorbed in a semiconductor material, where it generates positive and negative charges. Under the influence of a strong electric field, the positive and negative charges are pulled in opposite directions, generating an electrical signal.

## Slide 70
Optiso UDR detector: LSO crystal arrays with SiPM
Siemens Biograph Vision
PET/CT GE Omni Legend

## Slide 71
Typical PET scanner detector ring

## Slide 73
PET detector design

## Slide 75
Time of Flight

## Slide 76
Transaxial slices of an image quality phantom, reconstructed with high-resolution iterative reconstruction with resolution recovery. The scans were performed on a Siemens Biograph mCT PET/CT scanner (527 ps intrinsic time resolution) and the new Siemens Biograph Vision (210 ps intrinsic time resolution). From left to right: a non TOF image from mCT, b TOF image from Biograph mCT, and c TOF image from Biograph Vision. The phantom had six hot spheres (5, 8, 10, 13, 17, 22 mm diameter) with a contrast of about 4:1, and a total activity in the background of about 50 MBq of 18F  reference: Conti 2019, the new op

## Slide 77
Time resolution for LSO: 300 ps
1st TOF PET in 2006

## Slide 79
PET/CT Scanner: combined unit
Two gantries combined in same housing
Multislice CT scanner mated to a PET scanner
Improved attenuation correction
Possibly three scans acquired during procedure
  Attenuation correction CT
  PET
  Treatment planning CT, with contrast if necessary
Phillips Veros Digital PET/CT
GE Omni Legend

## Slide 80
Principle:
  MRI utilizes a powerful magnetic field (typically 1.5 to 3 Tesla) to align the protons in the hydrogen atoms
  Send a short radiofrequncy (RF) pulse to temporarily disturb the alignment of the protons to excites them and move out of alignment with the magnetic field.
  When the RF pulse is turned off, the protons relax back to their original alignment with the magnetic field. During this relaxation process, protons release energy in the form of RF signals.
  Signals are detected by the MRI receiver coils. The frequency and timing of the signals depend on the tissue type and water content, etc
  Small magnetic field gradients are applied in different directions (x, y, z axes) to spatially encode the signals. This allows the MRI system to determine the origin of the signals in 3D space.
Different tissues (e.g., muscle, fat, water) have different relaxation properties, leading to distinct signal intensities.
MRI images
T1 relaxation time
T2 relaxation time

## Slide 81
Ultrasound
Gallstones causing “acoustic shadowing” for US imaging
transverse
longitudinal

## Slide 84
Toward a planning scheme for emission guided radiation therapy (EGRT): FDG based tumor tracking in a metastatic breast cancer patient
EGRT treatment scheme consisting of three major components: simulation, pretreatment, and treatment.

## Slide 85
MRI image guided RT

## Slide 86
Biologically guided RT (BgRT)

## Slide 88
Fig. 1. Schematic design of the MR-linac concept. The accelerator is mounted on a ring gantry. The superconduction coils (in orange) are repositioned in order to create a low field toroid (light blue) as well as to create a beam portal. Also the gradient coil (in yellow) is split to allow beam passage