## Slide 1: 302.4.04 – Cryo-assemblies Horizontal Test Weekly status report
G. Chlachidze, S. Stoynev
Nov. 18th, 2024

## Slide 2: Next Steps
LQXFA/B02 Test Traveler is closed
LQXFA/B02 Test Report – in progress, will be ready for the internal review this week
LQXFA/B03 test preparations:
  Run plan drafted
  Test checklist drafted
  Checkout forms – work in progress

## Slide 3: Next Steps
Short term steps to improve/expedite the controlled cooldown and warmup at Stand 4
Complete insulation of the helium lines & relief valves upstream to the HX
  Indoor work in progress
  Waiting equipment for the outdoor work
Update/Improve the controlled cooldown and warmup procedures for cryo-operators – Planned for next week
Replace faulty GHe flow-meter – Work starts this week

## Slide 4: Next Steps
Nitrogen vent

## Slide 5: Next Steps
Long term improvements may include replacement of the heat exchanger at Stand 4
Two main goals for the HX replacement: increased reliability of the cold operation and faster cooldown and warmup
Cost estimate based on the installation experience at CMTF in 2023

## Slide 6: HX Replacement
Powerful immersion heaters already in use at various test facilities, including in IB1
Design for 16.5 kW heater is available, rough cost estimate 50-60 k$
With this improvement we could safely run with Ghe flow up to 18 g/s
  Need to replace the sub-cooler if we want to exceed 20 g/s
Immersion Heater with Conflat Flange
Helium flow