## Slide 1: 302.4.01 Magnets Vertical Test at BNL Weekly Status
Febin Kurian
16-Dec-2024

## Slide 2: Major Power Failure in the Building
Power in most parts of the building is down since Sunday evening 5.30 pm and is still not restored.
Power is fully restored yesterday (12/15/24) night
Assessing any damage, extend of helium loss etc.. starting today morning

## Slide 3: MQXFA12b-Status
Cooldown– started on 19th Nov.
Magnet reached 4.5K on 22nd Nov.
Spontaneous quench test started: Nov-26
Trip#1: IQ =14732 A, QI= 26.62 Miits (trips on the coil difference, as observed in few previous magnets)
Quench#1: IQ =15283 A on Q3, QI= 26.03 Miits, Quench Location: 0.06 m from center towards LE
Quench#2: IQ =15783 A on Q3, QI= 26.46 Miits, Quench Location: 0.91 m from center towards LE
Quench#3: IQ =15880 A on Q2, QI= 26.10 Miits, Quench Location: 1.66 m from center towards LE
Quench#4: IQ =15964 A on Q3, QI= 25.52 Miits, Quench Location: -1.09 m from center towards NLE
Quench#5: IQ =16097 A on Q3, QI= 26.14 Miits, Quench Location: 0.06 m from center towards LE
Quench#6: IQ =16334 A on Q3, QI= 26.25 Miits, Quench Location: -1.34 m from center towards NLE
Quench#7: IQ =16341 A on Q3, QI= 25.49 Miits, Quench Location: -1.24 m from center towards NLE
Quench#8: IQ =16157 A on Q3, QI= 25.17 Miits, Quench Location: -1.34 m from center towards NLE
Quench#9: IQ =16419 A on Q1, QI= 26.34 Miits, Quench Location: 1.31 m from center towards LE
Quench#10: IQ =16273 A on Q1, QI= 25.73 Miits, Quench Location: -1.99 m from center towards NLE
Quench#11: IQ =16332 A on Q3, QI= 25.96 Miits, Quench Location: -1.24 m from center towards NLE
Quench#12: IQ =16419 A on Q3, QI= 26.42 Miits, Quench Location~ -1.29 m from center towards NLE (Likely, not very clear on this location)
Ramp#14: No Quench, 16530 A at a ramp rate of +/-20A/s, 30 min hold.

## Slide 4: MQXFA12b-Quench Summary
Chart: MQXFA 12b Vertical Test All tests at 1.9 K, 20A/s unless specified otherwise
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35
Trip | 14.732
Quench(Q1) | 16.419 | 16.273
Quench(Q2) | 15.88
Quench(Q3) | 15.283 | 15.783 | 15.964 | 16.097 | 16.334 | 16.341 | 16.157 | 16.332 | 16.419
Quench(Q4)
Inominal(kA) | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23 | 16.23
I_Acceptance (kA) | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53 | 16.53
No Quench | 16.53 | 16.53 | 16.23 | 16.53 | 16.23 | 16.23 | 16.23 | 16.23

## Slide 5: MQXFA12b-Quench Locations
Chart: MQXF12b Longitudinal quench location from Quench Antenna
| 1311 | -1989
Q1 (224) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18
Q2(159) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18
Q3 (225) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18
Q4(148) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18
Quench No. | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16

## Slide 6: Second Top Hat Repair
  Heaters: All Passed Hi Pot
  V-taps:
    Passed: VT1, VT2, VT3, VT4, VT6
    Failed: VT5
  VT7 is being replaced to VT5
    To be “hipotted” sometime this week

## Slide 7: CRYOGENIC STATUS
CTI 4000 system status
  shut down
Linde 1610 system status
  Shut down