## Slide 1
Laser Based H- Beam Diagnostics
Yun Liu
for Beam Instrumentation Team
Research Accelerator Division
Spallation Neutron Source
Notes: 1

## Slide 2: Outline
Overview of laser based H- beam diagnostics
SCL laser wire profile monitor
HEBT laser emittance scanner
MEBT laser bunch shape monitor
Commissioning experience
Conclusion
Notes: 2

## Slide 3: Laser Based Beam Instrumentation at the SNS Accelerator Complex
Ion Source
Liquid Hg Target
CCL
SRF, b=0.61
SRF, b=0.81
DTL
MEBT
RTBT
HEBT
Injection
Extraction
Ring
Mode-lock Laser
1
1
MEBT Laser Bunch Shape Monitor
Q-Switch Laser
2
2
SCL Laser Wire Profile Monitor
2.5 MeV
1 GeV
87 MeV
186 MeV
387 MeV
3
HEBT Laser Emittance Scanner
3
Laser Assisted H- Stripping
4
4
Macro-pulse Laser
Notes: 3

## Slide 4: Photo-ionization – Physics behind Laser Based Ion Beam Diagnostics
hn
H-
N
S
Electron Collector
H0
e
H0
Photo-ionization
Notes: 4
4

## Slide 5
Faraday Cup
Laser
H-
H0
electron
Deflector
X-scan
Y-scan
x
y
z
H- beam
Laser Wire Profile Monitor
Notes: 5

## Slide 6
1
2
3
4
32
12
13
14
15
250 m
25 m
160 m
17
5
Camera
27
Laser
Telescope
PZT Mirror
Laser Room
Pick-up BS
M
FM
V-scan
H-scan
Ion Beam
Laser light
M
VW
BD
SCL Laser Wire Profile Measurement System
Liu et al, NIMA 612 (2010) 241–253;
Appl. Opt. 49 (2011) 6816-6823.
Notes: 6
6

## Slide 7
Laser Transport Line
Pick-up Mirror # | Individual Mirror Reflectivity (%) | Receiving Power Ratio (calculated) | Receiving Power Ratio (measured)
CM01 | 100 | 11.7 | 8.9
CM02 | 50 | 11.7 | 11.4
CM03 | 33 | 11.7 | 9.3
CM04 | 25 | 11.7 | 11.6
CM12 | 20 | 11.7 | 12.7
CM13 | 20 | 14.6 | 13.8
CM14 | 10 | 8.1 | 10.2
CM15 | 10 | 9.0 | 8.7
CM32 | 10 | 10.0 | 8.2
a
c
b
1
2
3
4
32
12
13
14
15
250 m
25 m
160 m
17
5
b
27
Laser Beam
1cm
a
c
Notes: 7
7

## Slide 8: Laser Beam Pointing Stabilization
Position (220 m from laser)
Transported laser power
Hardin et al, Opt. Express 19 (2011) 2874-2885.
Feedback off
Feedback on
1.25 mm @ 250 m
laser
piezo-driven mirror
beam sampler
image sensor
error signal generation
controller
set point
driver
L
Notes: 8
8

## Slide 9
From EPICS, user can select one, multiple, or all scanners
EDM Screens for Laser Wire System
From EPICS, user can select scan range, step size, average number. Fitting is automatically conducted.

## Slide 10
Simultaneous Profile Scan

## Slide 11
Simultaneous Profile Scan

## Slide 12
1
2
3
4
12
13
14
15
32
X
Y
Z
5 mm
Chart: 
SCL H- Profiles (1150 KW, Sept. 20, 2013)
 | 1 | 2 | 3 | 4 | 12 | 13 | 14 | 15 | 32
sx (mm) | 3.62 | 2.19 | 1.68 | 4.18 | 3.22 | 3.13 | 2.63 | 2.08 | 2.61
sy (mm) | 1.86 | 1.92 | 1.88 | 1.75 | 3.99 | 4.04 | 3.75 | 2.02 | 1.45

## Slide 13
1
2
3
4
12
13
14
15
32
X
Y
Z
5 mm
Chart: 
SCL H- Profiles (850 KW, April 15, 2013)
 | 1 | 2 | 3 | 4 | 12 | 13 | 14 | 15 | 32
sx (mm) | 2.32 | 2.45 | 3.79 | 2.3 | 2.23 | 2.58 | 2.47 | 2.73 | 2.8
sy (mm) | 1.97 | 1.72 | 1.99 | 1.85 | 2.42 | 2.44 | 2.32 | 2.27 | 1.8

## Slide 14
1
2
3
4
12
13
14
15
32
X
Y
Z
5 mm
Chart: 
 | 1 | 2 | 3 | 4 | 12 | 13 | 14 | 15 | 32
sx (mm) | 3.02 | 1.65 | 3.34 | 4.39 | 1.70 | 2.75 | 3.12 | 3.10 | 3.89
sy (mm) | 2.05 | 2.48 | 2.40 | 1.78 | 2.79 | 3.40 | 3.56 | 2.78 | 3.32
SCL H- Profiles (950 KW, Sept. 13, 2012)
Liu et al, PRST-AB 16 (2013) 012801

## Slide 15
Measurement time window
Beam Status during LW Measurement

## Slide 16
from SCL
to Linac dump
to Ring
Laser Slit
Wire Scanner
Laser Transport Line
~11.6 m
HEBT
Dipole
H-
H-
H0
V-scan
H-scan
Ion Beam
Telescope
HEBT Laser Emittance Scanner
Laser wire scanner converts a narrow channel of H- beam into H0 beam
Titanium wire scanner measures divergence of the H0 beam released from laser slit
Measurement is nonintrusive.
Laser Slit
Wire Scanner
Notes: 16
16

## Slide 17
Performance Evaluation of Wires
100-um Tungsten Wire
Fraction of electrons passing through the wire
50-um Titanium Wire

## Slide 18
Multi-wire Ti Scanner
Wire thickness: 50 mm
Wire spacing: 25 mm
1
2
10
Ti Wire Bundle
Notes: 18
18

## Slide 19
Raw Signal from Improved System

## Slide 20: Emittance Measurement with Improved System
Horizontal
Vertical
Liu et al, NIMA 675 (2012) 97–102
x (mm)
y (mm)
-4
-2
0
2
4
-4
-2
0
2
4
x (mrad)
-1
1
-0.5
0.5
0
y (mrad)
-1
1
-0.5
0.5
0
3
2
1
0
2
1
0
e ~ 0.45 mm mrad
Notes: 20

## Slide 21: Laser Based Longitudinal Profile Measurement
1
2
3
11
~100 ps
Pico-second laser pulses
H- bunches
Measurement
Laser source: Ti:Sapphire mode-locked laser
Externally locked to accelerator clock
Pulse width: 2.5 ps
Repetition rate: 80.5 MHz (5th subharmonic of RF frequency)
MCP
Laser
H-
H0
Frequency shift
H- beam
Magnet
SR570
Output
402.5 MHz
80.5 MHz

Beam dump

Mirror
Notes: 21

## Slide 22
Prism
L
M
PBS
Prism
BBO
Filter
PD
HWP
B
Ti:Sapphire Laser
Isolator
HWP
Pol
L
L
LMA Fiber
Pol
BS
M
BS
PD
Camera
Power Meter
BS
Pulse width measurement
A
B
L1
q0
qin
Fiber
A
R&D: Fiber Transmission of ps Laser Pulses

## Slide 23
Transmission of ps Laser Pulses through LMA Fiber
Near-field
Beam profiles after fiber transmission
Pulse width broadening
Over 85% of overall transmission efficiency through a 30 m fiber
Nearly diffraction-limited output beam
A beam diameter of less than 400 µm at a working distance of 600 mm
At 3 KW transmitted (peak) power, pulse width broadens to 11.6 ps

## Slide 24: Setup Installed at SNS MEBT
PC
USB Juno
Power meter
Picomotor Actuator
Camera
Optical Fiber
Vacuum Chamber
PBS
Fiber probe
H- beam
A

## Slide 25: Measurement Results – Frequency Offset Mode
Chart: MEBT Laser Measurement
Instantaneous measurement
H- beam RF frequency: 402.5 MHz
Laser repetition rate: 80.501 MHz
Measured waveform is a magnified (in time domain) picture of the H- beam microbunch. The magnification factor is f0/Df ~ 80,500.
2.5 ns
~125 ps
Huang et al., Appl. Opt. 52 (2013) 4462-4467

## Slide 26: Measurement Results – Phase Scan Mode
Measurement time depends on H- beam frequency: 20 seconds for 60 Hz beam and 5 minutes for 1 Hz beam.

## Slide 27
Item | Findings | Solution
Laser Transport Line | Drift and vibration | Beam stabilization using active feedback Optical fiber based transport line (for low power)
Laser fluence | Over focusing of laser beam caused vacuum window breakdown | Avoid beam collimation optics close to measurement station. Ensure laser fluence below 1 J/cm2.
Influence on beam | Electron collection magnets can cause tiny beam deflection | Correction magnet installed Orbit correction
Radiation hardness of laser | Laser driver (> 6 m from beam line) damaged in 1-2 days Unclear about laser head | Laser should be located outside the beamline for hadron machine
Image sensors | Gigabit Ethernet cameras (> 1.5 m from beamline) | Have to replace every 1-2 years
Motion control | Stepper motor (~ 30 cm from beam line); Picomotor actuators (1.5 m from beamline) | Stepper motors are very robust Open-loop picomotors have to be used
Commissioning Experience

## Slide 28: SUMMARY
World-first demonstration of simultaneous H- beam profile scan using a single laser source. The system has been brought to operation level – a single push-button initiates profile scan at 9 locations of SCL (corresponding to energy levels of 200 MeV -1 GeV).
Laser emittance scanner has been commissioned at SNS HEBT.
Longitudinal profile measurement system has been developed using optical fiber transmission of picosecond laser pulses.
Laser based beam diagnostics at accelerator facilities is reliable and realistic and provides a useful tool for beam tuning and physics study.
Notes: 28