## Slide 1: Emittance Measurement via H- Laser Stripping at LINAC4
Project Status
September 2013
Thomas Hofmann
CERN – Beam Instrumentation (BE/BI/PM)
thomas.hofmann@cern.ch
T. Hofmann, F. Roncarolo, CERN
S. Gibson, G. Boorman, A. Bosco, Royal Holloway University of London
C. Gabor, Rutherford Appleton Laboratory

## Slide 2: Contents
Project Overview
Simulations
Components (Laser & Detector)
3 MeV Test Setup
Summary - Outlook

## Slide 3: CERN Accelerator overview
LINAC 4
LHC luminosity upgrade!

## Slide 4: LINAC 4 overview
160 MeV
100 MeV
50 MeV
Transfer line to PSB
Laser Emittance Meter
Particle Type | H- ions
Top Energy | 160 MeV
Current | 40 mA
Pulse freq. | 1 Hz
Pulse length | 400 µs
Vacuum pressure | 10-7 mbar
Estimated Completion Date | End 2015
3 MeV Testbench

## Slide 5: Concept of Laser Emittance Meter for 160 MeV
Detector
H- beam
H0 beamlet
beamlet
after drift
Principle:
H-
H0
Bending Magnet
x-scan
y-scan
Laser
H- beam
Notes: ICT: new LHC transformer to replace the FBCT in order to solve the beam position dependency issue

## Slide 6: Current Situation at end of LINAC4
Laser
Detector
PS-Booster
LINAC4 PIMS

## Slide 7: Setup at 3 MeV Diagnostics Testbench
Slit
Dipole
SEM grid

## Slide 8: LINAC4 pulse structure
400 µs - Macropulse
1 s
2.84 ns
full- & empty
buckets
RF-system:
352 MHz
622 ns
372 ns
Micropulse
After Chopper:

## Slide 9: Contents
Project Overview
Simulations
Components (Laser & Detector)
3 MeV Test Setup
Summary - Outlook

## Slide 10: Laser-Particle Interaction
  Gaussian Laser Beam with M²(Beam quality)
Lorentz-Shift of Photon Energy
  At 160 MeV: 1064 nm -> 909 nm
Photon flux varies along particle path
Calculation of pstrip at n positions
Laser requirements:
Laser wavelength: 1064 nm
M2 < 3 for diameter < 200 µm
PPeak > 10 kW
Ref: J.T. Broad and W.P. Reinhardt, Phys. Rev. A14 (6) (1976) 2159

## Slide 11: Particle Distribution @ 160 MeV
Size (1 sigma): x = 1.7 mm; y = 2.3 mm
Divergence (1 sigma): x‘ = 0.6 mrad; y’ = 0.7 mrad
Laser
Drift
~ 4 m
Detector
Laser
200 µm

## Slide 12: H0 distribution at detector plane
Resolution: 500 µm
Detector requirements:
~ 105 H+ / mm2 -> Radiation Hardness
Size: ~ 30 x 30 mm
Resolution: 200 µm -> ~5 strips/sigma

## Slide 13: Contents
Project Overview
Simulations - Background
Components (Laser & Detector)
3 MeV Test Setup
Summary - Outlook

## Slide 14: Sources for Background - Overview
Black-body Stripping
Magnetic field Stripping
  Both negligible
  Ref: Development of beam transverse profile and emittance monitors for the CERN LINAC4, B. Cheymol
Residual Gas Stripping
  Simulations done
Intrabeam Stripping
  Simulations to be done
  Ref: Intrabeam Stripping in H- LINACS, V. Lebedev
Dipole
H-
H0
e-
H-
H2
H0
e-
H-
H0
e-
H-
H-

## Slide 15: Model for residual gas stripping at 160 MeV
Diameter = 10 mm
RMS values:
x’ = 0.8 mrad
y’ = 1 mrad
 | LEBT | RFQ | Chopper | DTL | CCDTL | PIMS
Vacuum Pressure [mbar] | 10-5 | 10-7 | 10-7 | 10-7 | 10-7 | 10-7
PStripping | 2.6E-02 | 2.3E-04 | 1.4E-05 | 8.6E-06 | 4.3E-06 | 2.4E-06
H- arriving [1/ns] | 6.5E+01 | 1.3E+01 | 6.1E-01 | 4.6E+00 | 4.8E+01 | 6.1E+02
40 m
Accumulated background level at detector = 7.4E+02 [1/ns]
Residual gas consits mainly of H2
Detector requirement:
Fast! (τDetector < tLaserPulse)

## Slide 16: Measurement Setup at 3 MeV Teststand
SEM grid
H0 background
H- beam
200 µm slit

## Slide 17: Measurement Results from 3 MeV Teststand
Slit-and-grid (slitsize 200µm)
Scaling of background
  6% laser efficiency
  H0 instead of H-
  Overall Scaling factor: 39
Time behaviour not measured yet!

## Slide 18: Contents
Project Overview
Simulations
Components (Laser & Detector)
3 MeV Test Setup
Summary - Outlook

## Slide 19: Lasers
Laser | Fiber-laser Yb-doped | Solid-State Nd-YAG
Model | Manlight ML-30-PL-R-TKS | Spectron Mini-Q SLMQ1S
Wavelength | 1080 nm | 1064 nm
Energy | 0.9 mJ | 50 mJ
tpulse | 110 ns | 9 ns
Beam quality – M2 | 1.7 | 1.3
Repetition rate | 30kHz | 20 Hz
Beam transport | Fiber | Free Space
Pstrip @ 3 MeV | 6 % | 100 %
Pstrip @ 160 MeV | 1 % | 100 %
Ref: Description of laser transport and delivery system for the FETS laserwire emittance scanner, A. Bosco, IBIC13

## Slide 20: Laser impact on system performance
 | Laser Parameters | Fiber-laser | Nd-YAG
Spatial resolution | M2 | + | ++
Laser pulses / LINAC4 pulse | fRep | 12 | 1
Detector requirements (Bandwidth, Rad hardness) | tpulse; EPulse | ++ | -
Handling; Maintenance | Laser transport | +++ | --
SNR / laserpulse | Fiber-laser | Nd-YAG | Background
3 MeV Particles/bunch | 1.4e6 | 4.3e7 | 7.8e3
SNR | 154 | 5513 | 
160 MeV Particles/bunch | 2.7e5 | 4.3e7 | 2.1E+03
SNR | 129 | 20476 | 
Detector:
  Size: 30x30 mm
  Time resolution in ns-range
Laser in center of beam
Detector 3 m behind the slit
SNR averaged over detector size
Beam shapes as simulated
Signal - Noise Ratio:
Notes: Spatial resolution (w0 –reighleigh)
  Nd-YAG:
    Breiter weil 100% stripping
    Slit-breite einstellbar durch focussierung -> Besseres SNR leicht möglich!!!!
Detector requirements
  Fiberlaser:
    Pulse length higher -> Bandwidth detector can be lower
    Pulse energy lower -> Rad – hardness!
Repetition rate
  Better SNR through avg possible

## Slide 21: Detector alternatives: Overview
Type | SEM grid | Scintillator | Semiconductor e.g. Silicon | Poly crystal Diamond | Gaseous Detector (GEM)
Size | +++ | ++ | ++ | ++ | +++
Resolution | - | +++ | ++ | ++ | ++
Sensitivity | --- | + | +++ | +++ | +++
Bandwidth* | +++ | --- | ++ | ++ | +
Homogeneity** | +++ | ++ | +++ | - | ++
Rad. Hardness | ++ | --- | -- | - | ++
* Just charge creation. System depends of course on readout electronics!
**without irradiation.

## Slide 22: Detector for 3 MeV Tests
 | Detector CIVIDEC
Type | Polycrystalline
Size | 19 x 19 mm
Resolution | 5 channels
Thickness | 500 µm

## Slide 23: Contents
Project Overview
Simulations
Components (Laser & Detector)
3 MeV Test Setup
Summary - Outlook

## Slide 24: 3 MeV Testbench
160 MeV
100 MeV
50 MeV
Transfer line to PSB
Testbench

## Slide 25: Setup at 3 MeV Diagnostics Testbench
Laser
viewport
Diamond Detector

## Slide 26: Integration Laser Setup
Diagnostics rack
Linear stage with 500 mm lens
Class 4
fiber laser
Coupling
tube
PXI-System
Ref: Fiber Laser Transport for the Photodetachment Emittance Measurements at FETS, Alessio Bosco

## Slide 27: Integration Laser Setup

## Slide 28: Integration diamond detector
Installation of detector 3.3m downstream the laser
Vertical movement driven by stepping motor
Interlocked to protect detector in case of magnet failure
Readout: 1 GHz Scope

## Slide 29: DAQ and Control Concept
PXI
Digitizer
Time/Cnt
Triggers
LASER
From ion
source
SCR
Camera
Beam
Expander
Movers
Serial
Filter
Wheel
Firewire/USB
1GHz Scope
ethernet
External PCs

## Slide 30: Contents
Project Overview
Simulations
Components (Laser & Detector)
3 MeV Test Setup
Summary - Outlook

## Slide 31: Summary
Simulations done regarding the following:
  Laser – Particle interaction
  Background level
Review and comparison of suitable components
Preparations towards tests at 3 MeV done (together with RHUL)
  Development of fiber-laser setup and fiber-based transport to beampipe
  Implementation of diamond-detector in testbench
  Design of system for DAQ and controls

## Slide 32: Outlook
Tests at 3 MeV teststand (Okt 2013 – March 2014)
  Background characterization
  Test of diamond detector
  Test of laser stripping with low power fiber-laser
Developments for 160 MeV prototype
  Detector
    Simulations: Gaseous detectors
    Readout concept for high number of channels (~128)
  Laser
    Decision for lasertype (Nd-YAG / fiber-laser)
    Development of optical transport line for 160 MeV setup
Test of system at 160 MeV (Hopefully Mid 2015)

## Slide 33: That’s it.
Thanks for your attention!

## Slide 34
SPARE

## Slide 35: Alternative for 160 MeV Setup: Gaseous detector
High sensitivity (> 100 e-/H0)
Sufficient spacial resolution (100 µm)
Big sizes possible (> 10 x 10 cm)
High bandwidth (ns)
No radiation damage reported!
Similar application already realized
  Ref: GEM-based beam profile monitors for the antiproton decelerator
But: Futher simulations needed

## Slide 36: Comparison with AD setup (Duarte Pinto, Spanggaard)
 | Antiproton Deccelerator | LINAC4
Particle | Antiproton | H+
Beam Energy | 5.3 ... 126 MeV | 160 MeV
Spill length | 150 ... 300 ns | 10 ... 150 ns
Rate | 7*1012 Particles/s*mm2 | 1012 ... 1014 H+/s*mm2
Size | 100 x 100 mm | ~ 20 x 20 mm
Resolution | 1.6 mm (x & y) | ~ 100 µm
Time resolution | 20 ms | ~ 5 ... 20 ns
Radiation hardness | ? | ~ 1015...1016 H+/cm-2
Homogeinity | ? | ~ +/- 1 % over lifetime

## Slide 37: Chopper Quad settings
Quads
Quads
Laser Viewport
Aim: Maximal amount of particles stripped
Optimisation of Quadrupol gradients to focus beam
Q1: 8.15 T/m
Q2: -7.15 T/m
Q3: 11.9 T/m
Q4: -11.5 T/m

## Slide 38: Beam shape @ 3 MeV at laser

## Slide 39: Beam after PIMS

## Slide 40: Result: Beam dynamics @ 3 MeV
Size (1 sigma): x = 4.5 mm; y = 1.6 mm
Divergence (1 sigma): x‘ = 0.7 mrad; y’ = 2.2 mrad
Laser
Drift
~ 3 m
Q1
Q2
Laser

## Slide 41: Stripping & propagation to detector
Divergence (1 sigma): x‘ = 0.7 mrad (conv)
y’ = 2.2 mrad (div)
Laser
Drift
3.0 m
Profile measurement not possible
BUT: No Signal / Background ratio is good

## Slide 42: Residual Gas Stripping
Stripping probability rises with gas pressure
Stripping probability decreases with particle energy
H-
H2
H0
e-
Beam pipe
    l = length
;

## Slide 43: Residual gas composition
Ref: Chiara Pasquino
Residual gas consits mainly of H2
For background simulations just H2 is taken

## Slide 44: 160 MeV Emittance Simulation

## Slide 45: 3 MeV Emittance

## Slide 46: Radiation hardness
3 MeV
160 MeV
  Model: NIEL (Non ionizing energy loss)
  Displacement damage
  Damage equivalent to
  1 MeV neutron
Effects (for Si)
  Leakage current
  Doping change
  -> Change of sensitivity
Issues
  Sensitivity inhomogeneities
  Transverse drift of e- and
  -> Inhomogenous detector behavior

## Slide 47: Fluence Scenario @ 3 MeV
Scenario for 40 mA beam current
Can be scaled with beam current
Fluence scenario for LINAC4 at | 3 MeV
Position | after RFQ
Fluence in central mm^2 of detector per pulse | 
Fiber laser created H0 per laserpulse | 2.50E+05
Hardness parameter (NIEL) | 10
Background per 400us pulse | 1.40E+07
Hardness parameter (NIEL) Background | 10
Equivalent fluence in center per laser position [1/cm^2] | 1.43E+10

## Slide 48: To be done: Detector – Particle interaction
H0
500 µm
48 µm
p
e-
Detector thickness: 500 µm
Range of 3 MeV Proton in Diamond: 48 µm
Gain (e-/H0) ?