## Slide 1: Fiber laser transport for the photo-detachment emittance measurements at FETS
Alessio Bosco JAI @ Royal Holloway University of London
Main Contributors:
G. Boorman, S. Emery, S. Gibson – RHUL
C. Gabor, A. Letchford - STFC RAL
J. Pozimski, P Savage – Imperial College
T. Hofmann - CERN
3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Fermi National Accelerator Laboratory – Batavia, IL, USA – 26-27 September 2013
1

## Slide 2: Outline
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
2
Introduction
Fiber laser
Fiber transport
Final focus assembly
Laser measurements before and after transport
Tests at Linac4
Summary and outlook

## Slide 3: Introduction
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
3
Laser based diagnostic at the Front End Test Stand (FETS) at Rutherford Labs
Photo-neutralization of H- ions for size and emittance measurements @ 3MeV
Please refer to Christoph Gabor’s talk for details on the Photo-detachment Experiment
Summary of Accelerator Parameters

## Slide 4: Introduction
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Experimental setup layout:
Laser System
Fiber Transport (10 – 100 m)
Laser delivery to interaction area
Control and DAQ (Labview based)
Particle detection (not discussed here)
4

## Slide 5: Fiber Laser Description
All Fiber (Yb:doped) MOPA
2m Fiber Pigtail
TTL driven Pump
l = 1080 nm
Pulse Duration 90 – 110 ns
Max Peak Power = 8 kW
Rep rate = 30 – 100 kHz
Max Av. Power (CW Pump) = 30 W
5
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications

## Slide 6
6
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Fiber Transport Description
Fiber Model LMA-GDF-20/130-M Length of fiber = 100 m (under test)
Coupling Efficiency in fiber = 70%
Numerical Aperture = 0.08
Core Diameter = 20 mm
Equivalent M2 = 2.35
Fiber coupling:
3 axes stage with 100 nm resolution
Optics will be aligned in a 19” rack box

## Slide 7
7
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Longitudinal and transversal motorized translation stages (1 mm resolution)
Fixed plate contains all the optical element for virtually alignment-free operation
Motorized zoom beam expander 1 – 8 times magnification factor (collimated beam)
Possibility to add a lens after the BE for spot sizes variable from 1/8 – 1 times (focused beam)
Laser Delivery

## Slide 8
8
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Laser Delivery
Beam delivery system enclosed in aluminum box, on 450x450mm breadboard.
The top and two side panels are removable for access.
Opening the box breaks the interlock and dumps the laser beam.
Key to override interlock for exceptional access / alignment.

## Slide 9
9
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
CAD model of the FETS diagnostic chamber.
Laser Delivery

## Slide 10
10
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Pre-Transport Laser Measurements
Transverse Mode
M2x = 1.81
M2y = 1.61
Shows some astigmatism
Power Vs Current

## Slide 11
11
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Amplification Rise Time = 430 ms
Pulse Duration ~ 92 ns
Pre-Transport Laser Measurements
2.5 ms Pulses @ 50 Hz Rep. Rate (12.5% duty cycle)
28 W within the macropulse, 3.5 W average power

## Slide 12: M2 measurement Setup
12
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements
Camera scanned through +/- 75 mm
to measure beam variation around focus.
Beam waist
Rayleigh range
Beam transverse size (@ 1/e2)

## Slide 13: M2 measurement: Z = -75mm
13
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements

## Slide 14: M2 measurement: Z = -45mm
14
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements

## Slide 15: M2 measurement: Z = -15mm
15
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements

## Slide 16: M2 measurement: Z = +15mm
16
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements

## Slide 17: M2 measurement: Z = +45mm
17
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements

## Slide 18: M2 measurement: Z = +75mm
18
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements

## Slide 19
19
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements
M2x = 1.756 +/- 0.012
w0x (2s) = 176 +/- 1.4 mm
  Transverse mode very near Gaussian at focus.
  M2 better than expected from fiber specs.
  Fiber transport corrects the laser astigmatism and improves laser mode.
M2y = 1.756 +/- 0.018
w0y (2s) = 174 +/- 1.7 mm

## Slide 20: Longitudinal Pulse Profile
20
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements
Peak Power = 0.62 kW
Measured with a 10 W laser (Max Peak Power = 2.5 kW) that we borrowed
from the laser company whilst our 30W system was under maintenance.

## Slide 21: Longitudinal Pulse Profile
21
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements
Peak Power = 1.25 kW
Measured with a 10 W laser (Max Peak Power = 2.5 kW) that we borrowed
from the laser company whilst our 30W system was under maintenance.

## Slide 22: Longitudinal Pulse Profile
22
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements
Peak Power = 1.87 kW
Measured with a 10 W laser (Max Peak Power = 2.5 kW) that we borrowed
from the laser company whilst our 30W system was under maintenance.

## Slide 23: Longitudinal Pulse Profile
23
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Post Transport Measurements
Peak Power = 2.50 kW
Measured with a 10 W laser (Max Peak Power = 2.5 kW) that we borrowed
from the laser company whilst our 30W system was under maintenance.

## Slide 24
24
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Tests At Linac4
Please refer to Thomas Hofmann’s talk for details on the detection and simulations
Diagnostics rack
Class 4 fiber laser
Coupling
PXI-System
Linear stage with 500 mm lens
tube
Laser and fiber installation expected for the first half of November 2013

## Slide 25
25
Fiber laser transport for the photo-detachment emittance measurements at FETS 3rd mini-Workshop on H- Laser Stripping and Accelerator Applications
Summary and Outlook
ADVANTAGES of FIBER TRANSPORT
Extremely stable and virtually alignment-free
Excellent transverse mode quality
Close to zero position jitter
LIMITS of FIBER TRANSPORT
Peak power handling capabilities (limit on transport length?)
Coupling efficiency (70 – 80%)