## Slide 1: Design for Radiation Shielding of PAL X-ray Free Electron Laser
Nam-Suk Jung, Hee-Seock Lee, Joo-Hee Oh
Radiation Protection Team
Pohang Accelerator Laboratory/POSTECH

## Slide 2: Outline
Introduction of PAL-XFEL and Safety Control
▷ Safety Policy & Area Classification
Beam Loss Scenarios and Radiation Shielding Design
▷ Normal Loss & Accidental Loss Scenario
▷ Bulk Shielding
▷ Entrance, Trench, Sleeve, Duct Design at Tunnel with Target Assumption
▷ Tune-up Dump Design (Preliminary)
▷ Main Beam Dump Design
▷ Safety Shutter at Accidental Loss Scenario (Preliminary)
▷ Neutron Skyshine Evaluation
Summary

## Slide 3: Bird-Eye-View of PALXFEL
PLS-II
(3rd SR)
PALXFEL
(4th SR)

## Slide 4: What is the PAL-XFEL ?
Pohang Accelerator Laboratory (PAL)
Hard X-ray FEL using 10 GeV electron
Soft X-ray FEL using 3.15 GeV electron
Linac | Beam energy, GeV | 10
 | Beam charge, nC | 0.2
 | Slice emittance, mm-mrad | 0.4 (0.6+)
 | Injector gun | PC RF-gun
 | Peak current at undulator, kA | 3
 | Repetition rate, Hz | 60
 | Number of bunch | Single or Two
 | Linac structure | S-band
Undulator | Undulator type | Out-vacuum
 | HX undulator gap (minimum), mm | 7.2
FEL | Hard X-ray wavelength, nm Soft X-ray wavelength , nm | 1 ~ 0.06 10 ~ 1.0
 | FEL radiation power @ 0.1 nm wavelength and 60 fs photon beam length, GW | 30++
 | Photon flux @ 0.1 nm, photons/pulse | > 1.0 E+12
Electron beam power : 120 or 240 W
Notes: 4

## Slide 5
General Layout of PALXFEL
Out vacuum Undulator
(EURO-XFEL undulator design )

## Slide 6
□ Radiation Control Policy based on Korean Regulation
- Annual dose limit for radiation worker: 20 mSv/year
- Annual dose limit for publics: 1 mSv/year
- Environmental limit (Site boundary): 0.25 mSv/year
- 6 mSv/yr for frequent visitors, 10 μSv during 1 hour for temporary visitors
□ Shielding Criteria based on ALARA
- 10 mSv/year on Surface (2000 h ⇒ 5 μSv/hour)
- 1 mSv during 1 hour for accidental event
- Area requirements (e.g. 0.4 mSv/week)
□ Area Classification
- Restricted Area: 0.25 mSv/y < Dose < 1 mSv/y
- Generally-Controlled Area: 1 mSv/y < Dose < 20 mSv/y
(a dosimeter is required)
- Radiologically-Controlled Area: 20 mSv/y < Dose < 1 mSv/h
- High Radiation Area: 1 mSv/h < Dose
(No Access)
Safety Policy of PAL

## Slide 7
Area Classification of PAL-XFEL
High Radiation Area (Radiologically-Controlled Area after machine OFF)
Radiologically-Controlled Area
Generally-Controlled Area
1F
Safety Shutter closing

## Slide 8
Normal beam loss scenario of the PAL-XFEL
1) Full power beam loss at main beam dump
: 120 or 240 W for HX linac (10 GeV, 0.2 nC, 60 Hz, 1 or 2 bunch)
2) Low power beam loss at linac and undulator section
  < Limit of loss level from beam physic group >
  - Linac : 1.2 W, 1% (2.4 W for 2 bunches)
  - Undulator : 0.12 W, 0.1% (0.24 W for 2 bunches)
  3) Low power beam loss at several dumps such as
  tune-up dump (10 Hz) and spectrometer dumps
  (60 Hz at 330 MeV, less than 10 Hz at 3.15, 10 GeV)
  4) Low power beam loss upstream the main dump magnet : 0.1% loss
  5) Beam halo interaction with collimator : not considered yet
  6) Beam interaction with diagnostics like screen, wire scanner
Normal Beam Loss Scenario

## Slide 9
Accidental beam loss scenario of the PAL-XFEL
1) High power beam loss above loss limit at linac and undulator section
⇒ Interlocked by MIS and PSIS
2) Failure or malfunction of main dump magnet
3) High power beam loss upstream the main dump magnet
⇒ High energy radiation at forward direction can reach experimental area.
Accidental Beam Loss Scenario
<Redundancy>
- Beam Off Detector
- Charge Comparator
- Safety magnet
(permanent)
- Collimators
- Fixed beam stopper
- Shutters

## Slide 10
Bulk tunnel shielding using fundamental semi-empirical formula
: Methods different to one for PLSI & II, SHIELD11 (SLAC)
: using Monte Carlo code, FLUKA
Analyzing for detailed geometries such as entrance, trench,
sleeve and duct at the tunnel
: using Monte Carlo code, FLUKA
Analyzing for complicated area after undulator to hutch
including main beam dump : using Monte Carlo code, FLUKA
Skyshine : using Monte Carlo code PHITS(with known eq.& SHINE-III)
Muon shielding : using Monte Carlo code, FLUKA
Gas bremsstrahlung : Ferrari & Tromba’s equation (Negligible!)
Shielding Analysis Methods

## Slide 11
Required Thickness at Different Energy
Using SHIELD11

## Slide 12
Tunnel Thickness
Section | Electron energy [GeV] | Thickness of the concrete [cm]
Injector | 0.139 | 100 (*L), 150 (R)
Linac | < 3 | 150 (L, R)
 | 3 ~ 10 | 200 (L, R)
SX undulator | 3.15 | 150 (L, R)
HX undulator | 10 | 200 (L, R)
* L : left side toward beam direction, R : right side
Determine the different thickness according to the electron energy
Main beam dump (100x100x130 cm3 iron) installed at the pit
: changing beam direction using main dump magnet

## Slide 13
Door, Maze, Trench, Sleeve, Duct
< Target assumption for FLUKA calculation >
• Location of door, maze, … : Sideward direction (90o)
⇒ Assumption of the target which produces intense radiation fields
toward the sideward direction is required to evaluate
conservative shielding design.
• 10 GeV, 3.15 GeV electrons on 1X0, 10X0, 20X0 thickness iron target
(1X0 of iron = 1.76 cm)
1) Calculation of maze entrance at the HX and SX undulator tunnel
⇒ Decide the thickness of target
2) Calculation of upper trench(square hole) and sleeve at the SX
undulator tunnel
⇒ Decide the relative position between target and structures

## Slide 14
Door, Maze, Trench, Sleeve, Duct
< 10 GeV electrons on 1X0 iron target >
⇒ Not effective for sideward structure but reasonable assumption for accidental loss
scenario such as ‘high power beam loss upstream the main dump magnet’

## Slide 15
Door, Maze, Trench, Sleeve, Duct
< 10 GeV electrons on 10X0 iron target >
< 1 μSv/h

## Slide 16
Door, Maze, Trench, Sleeve, Duct
< 10 GeV electrons on 20X0 iron target >
< 1 μSv/h
⇒ leakage dose eq. rate (maze exit or tunnel wall) with 10X0 and 20X0 Fe < 1 μSv/h (similar)

## Slide 17
Door, Maze, Trench, Sleeve, Duct
< 3.15 GeV electrons on 10X0, 20X0 iron target >
< 1 μSv/h
< 1 μSv/h
10X0
20X0
⇒ leakage dose eq. rate
(maze exit or tunnel wall)
with 3.15 GeV electrons
on 10X0 and 20X0 Iron
< 1 μSv/h (similar)
∴Select 20X0 thick iron
for sideward shielding
calculation

## Slide 18
Door, Maze, Trench, Sleeve, Duct
< Position assumption of thick target >
Target, 20X0
2.5 m 2.5 m
1
2
3
2
3
Dose Equivalent, μSv/h, XY plane, 0.1%
1
Dose Equivalent @ ZX plane, square hole height
⇒ position (90o from thick target) : reasonable assumption w/o position definition
2

## Slide 19
Tunnel Entrance : Door
- Material & thickness : O.C., 2 m
(wheel location) : Iron, 1.26 m
- Gap bet. door and floor : 10 mm
- Gap bet. door and tunnel wall
: 15 mm
- Overlap length : 60 cm (all sides)
- Thick iron target
- 2.4 W loss (1%, 2 bunches)
2 μSv/h
3 μSv/h
3 μSv/h
2 μSv/h
3 μSv/h
3 μSv/h
3 μSv/h
Dose Equivalent, μSv/h, ZX plane
Dose Equivalent, μSv/h, XY plane
< Example of sliding door >

## Slide 20
Trench, Sleeve, Duct at Tunnel
Exit of waveguide sleeve : 7.2 ± 1.9 μSv/h
Exit of LCW sleeve (upper) : 6.8 ± 0.9 μSv/h
Trench exit, 1 m away from wall : 5.5 ± 0.7 μSv/h
Duct exit : 40.3 ± 1.5 μSv/h
1 m away from duct exit : 6.8 ± 1.6 μSv/h
2.7 m away from duct exit : 3.0 ± 2.0 μSv/h
Opposite side of duct exit : 5.7 ± 2.4 μSv/h
Roof of duct : 2.2 ± 1.1 μSv/h
1
2
3
4
5
7
8
< 5 μSv/h
6
Trench, sleeve
: Measure dose and
install additional
shield if required
Duct : Do the access
control to the tunnel roof
(fence installation)
- Thick iron target
- 2.4 W loss (1%, 2 bunches)
Dose Equivalent, μSv/h, XY plane

## Slide 21
Tune-up dump
▷Purpose of tune-up dump
: Protection of undulator magnet through the mis-aligned electron beam
in the undulator during tune-up period
▷Concept of PAL-XFEL tune-up dump
: Deflecting beam to beam stopper (collimator) by inserting the dipole magnet
▷Location and initial condition of PAL-XFEL tune-up dump
Preliminary !
10 or 4 GeV, 0.2 nC
10 Hz operation (Goal)
: 20 W or 8 W

## Slide 22
Tune-up dump
▷ Electron stopper : ( 2.1 cm(6X0) W + 25 cm Cu ) x Ф20 cm
▷ Additional Shielding : Iron
Preliminary!!
- Vacuum pipe : Ф13 mm (outer),
1 mm thickness Al
- Cut the bottom of vacuum pipe &
install 50 μm Al foil for wake problem
- Neutron fluence limit @ magnet
: 1011 n/cm2
(from result of T-493 experiment
using 13.6 GeV electron @ SLAC)
- Calculated neutron flux @ magnet
: 7 x 107 n/cm2hr
⇒ Lifetime : ~ 1500 hr
We will perform design correction
including 17 cm thickness
quadrupole@z = -590
70 cm
80 cm
40 cm

## Slide 23
Tune-up dump
Preliminary!!
Dose Equivalent @ ZX plane (plan view)
Dose Equivalent @ ZY plane (vertical view)
Tunnel outside
: Effective!!
Lower than
shielding criteria

## Slide 24
  120 W beam, 20 m from beam loss point
  Dose from muon (0 degree) = 100 ~ 240 uSv/h
  Range of 1 GeV Muon = 0.8 m Iron
At least 80 cm-thick Iron beam stopper at optical hutch
Muon Yield Estimation
Using FLUKA code

## Slide 25: Major Points of Shielding Issues

## Slide 26: Prompt Dose around Main Beam Dump
Concrete cover
Main beam dump
(Iron)
Size : 100(H)x100(W)
x130(D) cm3
Hole : 15(H)x15(W)
x60(D) cm3
Main dump magnets

## Slide 27: Safety Shutter at Accidents
50 cm thickness Tungsten ⇒ Effective!!
- Detail design including separating and reducing the thickness
are performing.
Units :
x 2.7E+8 μSv/hr
<Assumption>
: 100% beam loss on
thin target near
main dump magnet
Dose Equivalent, μSv/h, ZY plane

## Slide 28
Operation Condition
- Electron Beam
- 10 GeV, 0.2 nC/pulse, 60 Hz, 2 bunches : 240 W
- Major Shielding of Linac and Undulator Tunnel
2 m-thick Ordinary Concrete
- Operation Time : 6000 hour/year
- Beam loss: < 1% at Linac tunnel (60 hours)
< 0.1% at undulator tunnel (6 hours)
Dose Control Level (Site boundary) : 0.25 mSv/year
- On-site & Off-site background level : < 0.1 μSv/hour
Environmental Radiation
Practical Condition
- On the top of mountain, at 130 m from residents apartment
Levels For the Public

## Slide 29: Concept for Skyshine Evaluation
Using PHITS code (photoneutron calculated by FLUKA)
① sky → detectors
② accelerator tunnel → detectors
③ ground → detectors
④ sky → ground → detectors
⑤ ground → sky → detectors
Skyshine term
Direct term
Groundshine term
Multipleshine term
Multipleshine term

## Slide 30
Skyshine effective dose (0.1% loss)
Dose at 100 m = 0.0548 mSv/y < 0.25 mSv/y
Dose by pure skyshine at 100 m
= 0.0062 mSv/y < 0.25 mSv/y

## Slide 31: Summary
Radiation shielding of the PAL-XFEL was designed.
- Beam loss scenario was established.
- Bulk tunnel thickness was determined using semi-empirical formula,
SHIELD11. : Ee > 3.15 GeV, 2 m tunnel thickness
- Complicated tunnel structures like entrance and duct were analyzed using
Monte Carlo code, FLUKA.
- Main beam dump and preliminary tune-up dump was designed.
- Thickness requirement for safety shutter was evaluated : 50 cm Tungsten
- Neutron skyshine was evaluated. : less than the limit at residence apartment
Thank you for your attention!!

## Slide 32
Backup slides

## Slide 33
High Radiation Area
1) Tunnel inside during accelerator operation
2) OH inside with opening of the safety shutter(SS)* during accelerator operation
Radiologically-Controlled Area
1) Tunnel inside after the accelerator operation
2) OH inside with closing of the SS during accelerator operation
3) Tunnel roof during accelerator operation
Generally-Controlled Area
1) Experimental Area
2) Klystron Gallery
3) Loading Area, Control room and Office
* Safety Shutter(SS) : the shutter installed bet. frontend and OH, and located at undulator tunnel end
Area Classification of PAL-XFEL
Notes: 4세대 구역 설정안 발표 하겠습니다.
고방사선구역으로

## Slide 34: BL Shielding Effectiveness at Accidents
Dose Equivalent, μSv/h, ZX plane
Dose Equivalent, μSv/h, ZY plane
※ Assumption
- 100% beam loss on thin target
near main dump magnet
- Tungsten safety shutter is not operated yet.
- 80 cm iron fixed stopper at optical hutch with 3 cm offset
Fixed stopper is effective but
high dose eq. at uncovered area
⇒ Detailed consideration about
offset and stopper should be
required.
< Example of accidental loss >
150 μSv/h
54 mSv/h
Preliminary!!

## Slide 35
Cylindrical shell (t-track) at 0, 5 10, 50 100, 150, 200, 250, 300, 350, 400, 500, 600, 700, 800, 900, 1000 m from outer surface of tunnel
E-beam
Geometry at PHITS Simulation
Geometry of accelerator building