## Slide 1: Health & Safety
Handling hazardous and greenhouse gases at ALBA Synchrotron facility

## Slide 2: INDEX
ALBA LINAC
BL22 - CLӔSS

## Slide 3: ALBA LINAC
1
LINAC Function
3
Sulfur hexafluoride (SF6)
2
Klystrons
4
Maintenance tasks
5
Conclusions
6
Acknowledgement
Greenhouse gas case

## Slide 4: 1. LINAC FUNCTION
EGun
Bunching
Main Acceleration
v ~ c
90 keV
16 MeV
70 MeV
100 MeV
LINAC accelerates the electrons and increases the energy

## Slide 5: 2. KLYSTRONS
Function:
  Supply RF power to LINAC through waveguides
Problems:
  Electric arcs
SF6 is used as electrical isolation
waveguides
waveguides

## Slide 6: 3. SULFUR HEXAFLUORIDE (SF6)
Disadvantages of SF6 usage:
  Greenhouse gas
  Global-warming Potential (GWP):
    1 GWP SF6  23900 GWP CO2
  Thermal decomposition generates toxic substances(produced during repeated sparking inside
  the waveguide)
  (GWP): is a relative measure of how much heat a greenhouse gas traps in the atmosphere.

## Slide 7: 3. SULFUR HEXAFLUORIDE (SF6)
HF(g): Hydrogen fluoride
SO2 (g): Sulfur dioxide
SF4(g): Sulfur tetra fluoride
CuF2(s): Copper fluoride
More common SF6 Decomposition By-products during repeated sparking
SF6 + H2O + Cu + ε  HF, SO2, SF4, CuF2, …

## Slide 8: 3. SULFUR HEXAFLUORIDE (SF6)
It means that during our maintenance tasks (specially after sparkling events) we have three problems to solve:
  Presence of:
  - Used gas (SF6).
  - Toxic & Corrosive decomposition gaseous & powder products.

## Slide 9: 3. SULFUR HEXAFLUORIDE (SF6)
Actions to be taken when the waveguides are opened:
  To recover the used & contaminated gas
  To protect our staff during the operation
waveguide

## Slide 10: 4. MAINTENANCE TASKS
Step 1: Gas recovery
SF6 Supply Bottle
Outdoor
Recovery Bottle
Outdoor
Compressor
Indoor

## Slide 11: 4. MAINTENANCE TASKS
Step 2: Open waveguides
  Working area delimited
  Personal Protective Equipment:
    Filter mask: FFP3 EN type
    Disposable coverall: Tyvek C type 5 EN
    Latex gloves
  Waveguides cleaning:
    Check if there are white powder
    Use of a vacuum cleaner with HEPA filters

## Slide 12: 5. CONCLUSIONS
Maintenance performed properly:
From a safety point of view
From an environmental point of view

## Slide 13: 6. ACKNOWLEGMENT
LINAC section staff:
David Carles
Cristina Orozco
Núria Martí
Infrastructure section staff:
Francis Pérez
Raquel Muñoz
Davide Lanaia

## Slide 14: BL22 - CLӔSS
1
CLӔSS Function
3
New set-up
2
Experiment set-up
4
System reaction
5
Improvements
6
Acknowledgement
Hazardous gases case

## Slide 15: 1. CLӔSS FUNCTION
The CLÆSS beamline develops “in situ” chemical and catalytical research
X-ray Beam
X-ray Beam

## Slide 16: 2. EXPERIMENT SET-UP
Pressurized Gas Bottle location:
1b. Inside EH
3. Mass Flow Controller (MFC)
2. Dedicated pressure regulator
Inside EH
P up to 10 bar
High P up to 50 bar
1a. Outside EH
Kr
He
N
Ar
He
Soft Mixture pipeline

## Slide 17: 2. EXPERIMENT SET-UP
5. Exhaust system
Air flow = 280 m3 / h
Dedicated line
4. Reactor Cell
T = 700ºC
P = 6 bars
V = 100 ml
3. Mass Flow Controller (MFC)
Allows remotely control: up to 6 gases
1 x 1l / min flow max.
5 x 100ml / min flow max.
Can be adjusted to different reactor cells

## Slide 18: 3. NEW SET-UP
Fixed gas detectors:
Safety Cabinets:
Flammable gases:
  H2
  HxCx
Hazardous gases:
  CO
  H2S
  NO

## Slide 19: 3. NEW SET-UP
Electrovalves:
Alarm gas system:

## Slide 20: 3. NEW SET-UP
Alarm exhaust system:

## Slide 21: 4. SYSTEM REACTION
In case a leak is detected:
Fixed detectors read values on line (inside the hutch and the safety cabinet)
Electrovalve closes the bottle immediately

## Slide 22: 4. SYSTEM REACTION
Alarm gas system warns the leak to user
  Sound
  Light

## Slide 23: 5. IMPROVEMENTS
The exposition & gas volume lost are reduced in case of leak
  Automatic electrovalves Vs Manual intervention
Time reduction in taking actions

## Slide 24: 5. IMPROVEMENTS
Reliability
  Permanent Vs Temporal

## Slide 25: 6. ACKNOWLEGMENT
Infrastructure section staff:
BL22-CLӔSS staff:
David Carles
Cristina Orozco
Núria Martí
Ignasi Serra
Edmundo Fraga
Laura Simonelli
Carlo Marini
Marta Ávila
Wojciech Olszewski

## Slide 26
THANK YOU
ANY QUESTION?