## Slide 1
High beam tube temperature stability (<3 mK/h) using
two phase Neon cooling
The windowless gaseous tritium source (WGTS) of the KATRIN experiment
Laura Kuckert for the KATRIN Collaboration, KCETA, Karlsruhe Institute of Technology
4 of 5 KATRIN main systematic uncertainties related to source
One systematic uncertainty allowed to cause shift in neutrino
mass squared of ∆mν2 < 7.5∙10-3 eV2/c4
Direct and model independent determination of the neutrino mass with sensitivity of 200 meV/c2 (90% CL)
WGTS
Temperature and gas dynamics
Entire spectrum
End point region
Injection orifices
4-fold pump port with connection to TMP‘s
∆TS =3mK, ∆TSensor =4mK
KATRIN Experiment
Source parameter | Requirement | Max. syst. ∆mν2 / eV2
Column Density → Temperature → Injection pressure | ∆ρd/ρd < 2∙10-3 ∆T/T < 2∙10-3 ∆pin/pin < 0.06 | < 1.5∙10-3
WGTS potential → Space charges | ∆U < 10mV | < 1.2∙10-3 3∙10-4
Final state distribution (FSD) → Ortho-Para ratio | ∆FSD/FSD < 0.01 | < 6.2∙10-3 < 3∙10-3
Unfolding of energy loss function → Inel. scattering cross section | ∆σin/σin < 2∙10-3 | 2 - 6.2∙10-3 < 4∙10-3
Source sytematic uncertainties
Raman
spectrum
of injected
gas
Source and transport section
Spectrometer and detector section
Calibration
and
monitoring
system
DPS1-F
DPS1-R
CPS
DPS2-F
WGTS
Example: syst. uncertainty
in WGTS temperature
Shift in fitted
neutrino mass
Difference in integrated β-spectrum
for ∆T=30mK
Systematic uncertainty in parameter
for ∆T=30mK
Measuring integrated electron spectrum from beta decay
of molecular tritium with large electrostatic spectrometer
(MAC-E filter)
Temperature measurement
+
Monitoring of injection pressure, column density
Simulation of
gas dynamics
1D density profile
Precise temperature measurement with PT500 sensors
+ vapour pressure sensors (∆T < 4mK)
Gas density profile from numerical simulation
Neutrino mass shifts due to source systematics are in agreement with sensitivity of of 200 meV/c2
2-phase LNe, T0=30K
0.4
0
-0.4
Pseudo-3D density profile
rel. density deviation n(x,y)
Beam tube: length 10m,
diameter 90 mm
Magnetic guiding field 3.6 T
Beam tube temperature 30K
Pressure reduction by turbo molecular pumps (TMP) at both ends
Tritium injection 5∙1019 T2-molecules/s (column density 5∙1021m-2 ) →
→ luminosity 1.7∙1011 Bq
Closed Tritium cycle with continuous purification →
purity of injected tritium > 95% (monitoring with Raman spectroscopy)