## Slide 1: Scintillators: Setup, performance and lessons learned
Ran Hong
CENPA, University of Washington

## Slide 2
Scintillator/Light-guide Setup
PMT Setup
Signal reading

## Slide 3: Scintillator Setup
μ Beam
μ Veto
ScL
ScR

## Slide 4: Scintillator Setup
Easy to change Scintillator angle, but hard to change it’s height, impossible to change its horizontal position
Tiny shift between ScL and SiL
Very hard to mount the light guides after the scintillators are glued on

## Slide 5: Light guide configuration
Plastic Light Guide
Torr Seal: Cause severe leak!
Aluminum/Stainless steel sleeve
KF40 flange
Diff. Pumping port
Sealed by pressure/friction: plastic to metal

## Slide 6: Light guide leak
LEAK!

## Slide 7: Fix the leak
Newest Torr-seal, leaks most
UHV sealant does not help
Easy Fix: Covering with excessive
UV cured optical glue

## Slide 8: PMT
PMT Module: XP2982 (datasheet elog:490)
Bias: -1800V (see elog:507)
2 PMT bases broke, 1 person shocked!
  ScR first, then Mu-Veto

## Slide 9: PMT Base Problem

## Slide 10: Why shocked?
HV supply
Cathode
Dynodes
Anode
HV cable GND
PMT Base GND
Becomes HV!

## Slide 11: Signal of ScL/ScR/MuVeto
ScL
ScR
MuVeto
ScL may be broken (elog:440)
Pictures are from elog:475

## Slide 12: Signal of ScL/ScR/MuVeto
Pulse shapes using FADC, elog:513,535

## Slide 13: Signal of ScL/ScR/MuVeto
Using BU digitizer to read the Scintillators (elog:539)
  Ch0: muSC
  Ch1: ScL
  Ch2: ScR
  Ch3: muVeto
For analysis: see whether we have good energy resolutions / good pulse shapes.

## Slide 14: Lessons learned
The seal between the light guide and its sleeve is flaky, the KF feed-throughs are OK.
PMT bases are fragile, need more careful handling. Check connections before biasing. Need more spare ones (or more robust ones).
FADC sample rate is not high enough for PMTs, need shaping amplifiers or faster digitizers. Analysis group should pay attention to the pulse shapes of the PMTs.

## Slide 15: Bonus: Si16

## Slide 16: Bonus: Si16
Cover the bunch of cables with copper tape
Braids of the cables are isolated from the connector
Tricky to get rid of noise/ground loop (elog:599)
Achieved resolution using 1MeV pulser: 100 keV
Not be able to use Am241 to calibrate: The air attenuate the alpha particles too much (elog:607)