## Slide 1: ASTA Laser system overviw and Gun Commisioning Results
Jinhao Ruan

## Slide 2: Outline
Status of Subsystems
  Photoinjector Gun
  Laser
Gun commissioning
Near future schedules
Summary

## Slide 3: FY14 - 50 MeV Injector + 1 Cryomodule
Goal: installation complete and beam commissioning started by end of FY14
  RF gun + RF system and photocathode laser system (Done)
  2 SRF booster cavities (CC1 and CC2) + RF systems (CC2 is ready , CC1 need more attention)
  50 MeV beam line elements and instrumentation to the low energy dump (2 girder is in and the third one is close)
  Low energy beam dump. (Dump installed)
  SRF cryomodule (RFCA002/CM2) (7 cavity has been commissioned)
Installation of 1st AARD experiment (high brightness X-ray channeling source) (Will be installed in girder 3)
electron gun
booster cavities
flat beam transform
chicane
beam dump
1st cryomodule
test beamlines
beam dump
spectrometer magnet
50 MeV Beam Energy
ASTA Photoinjector

## Slide 4
Status – Photoinjector Gun

## Slide 5: Status – Laser Diagram
Major change since year’s user meeting: 3 single pass amplifier replaced MP structure
  Get rid of the longitudinal laser structure because of the MP cavity
  Enable the possibility of even high repetition rates train (>3MHz)
Preamp ~3 single pass amp

## Slide 6: User Interface

## Slide 7: System Performance: Seed Laser

## Slide 8: System Performance: Seed Laser

## Slide 9: System Performance: Seed Laser

## Slide 10: System Performance: Amplifier

## Slide 11: System Performance: UV section

## Slide 12: System Performance: UV section

## Slide 13: System Performance

## Slide 14: 1ms IR laser train

## Slide 15: Current Laser system Status
Laser is up running with 3MHz pulse train up to 1ms and 5Hz repetition frequency
User interface is set up
Laser system is designed to be upgradable for up to 1.3GHz in principle, which will be essential to the ICS proposal.

## Slide 16: Gun Commissioning Milestone so far
Conditioning of the gun with 3MW Klystron is finished at 2013
Photoelectrons first produced at ASTA on June 20, 2013
  Molybdenum (uncoated) cathode
  First visible on Loss Monitor, Faraday cup
  Charge is very limited due to the low QE
CsTe cathode is installed in Feb 2014.
Photoelectron from CsTe was observed at Mar. 18 2014
Photoelectron up to 1ms long establish in June 2014

## Slide 17
First Electrons June 20 2013

## Slide 18: First electrons from CsTe Cathode
First Gun Phase scan done with the CsTe cathode

## Slide 19: First electrons from CsTe Cathode
100 us long electron bunch train on the faraday cup. Later calibration resolved the current is more than 1.5nC/bunch

## Slide 20: First electrons from CsTe Cathode
Signal from the wall current monitor

## Slide 21: First electrons at 9-way YAG screen

## Slide 22: QE map of the new Cathode
The average QE from the cathode is about 1~1.5% according to the calculation.
Courtesy of Dean. Edstrom

## Slide 23: Detailed gun Phase scan
Blue curve is the ASTRA simulation curve for charge vs gun phase. Green curve was also from ASTRA. The black circle is the experiment data.
The initial try fit the rising edge relatively well, however still need better understanding at falling edge

## Slide 24: Recent result

## Slide 25: Recent result

## Slide 26: Summary
Photoinjector has been commissioned successfully.
  electron pulse train up to 1ms long with about nC electron (3mA).
  9mA is observed with 0.3ms long bunch train
CC1 and CC2 will be commissioned at the end of this FY.
First user experiment will be carried out very soon.

## Slide 27: Thank you for your attention !

## Slide 28: First electrons from CsTe Cathode