## Slide 1: Mu2e Tracker Straws
Chiho Wang
Straws L3 Manager
7/9/2014
U.S. DEPARTMENT OF
ENEGY
Office of
Science

## Slide 2: Outline
Straw requirements
Straw property measurements
Straw assembly components and procedures
Cost

## Slide 3: Straw requirements
5mm ID metalized Mylar tube cathode.
Minimal material
  6 μm Mylar + 3 μm adhesive + 6 μm Mylar double helical wrap
  Inner wall coating: 500Å Al + 200Å Au
  Outer wall coating: 500Å Al
Operates in vacuum
  Sustains > 1 atm pressure difference
  Leak rate < 7 ccm / detector volume
Stability
  Straw straightness: max. transverse deviation/sagging < 300 μm for HV stability.
    Longitudinal tension is applied to keep straw straight.
    Initial tension need to be higher to counter for material relaxation over time (creep).
  Sustains radiation over the life time of operation.

## Slide 4: Straw properties
Destructive tests:
  Pressure: sustained 60 psi (10 min.). Destroyed ≥ 70 psi
  Stretch: sustained 1.6 kg (2 yrs.). Destroyed ≥ 2.9kg
Mechanical:
  Linear density: 0.34±0.01 g/100cm
  Wall thickness: 15μm
    Derived from linear density, assuming Mylar density 1.39g/cm3, and same for the polyester based adhesive.
  Spring constant: 0.891cm/kg/100cm
Electrical:
  Cathode resistance: 120Ω/100cm

## Slide 5: Straw leak measurement 1
Pressurize to ~ 1 atm and measure the pressure drop

## Slide 6: Straw leak measurement 1
CO2
N2
N2 leak rate =0.095±0.026 psi/day
CO2 leak rate =0.36±0.06 psi/day
Assume Ar leak rate similar to N2 , and straw volume ~ 0.3 m3, estimated Ar(80%)CO(20%) leak rate ~ 2ccm / tracker volume

## Slide 7: Straw leak measurement 2
Fill straw with ArCO2 and place in vacuum,
Measure pressure rise
Tested 10x 129cm straws
Averaged leak rate = 3.6x10-4ccm/straw
~ 4 ccm / tracker volume
Gas inlet
To straw
Vacuum pump
Vacuum
gauge
1 psi
Pressure relief
Straw
Feed through

## Slide 8: Straw straightness measurement To determine the required straw tension
Horizontal Traveling Microscope (X)
Vertical Traveling Microscope (Y)
Straw
Straw deviation measured for different straw length and tension.

## Slide 9: Straw straightness measurement
Mid point deviation, relative to the highest tension (500g)
Chart: 
| 100 | 150 | 200 | 300 | 400 | 500
ΔR 0psi 119cm | 0.10242 | 0.029017 | 0.03015 | 0.019799 | 0.0036056 | 0
ΔR 14psi 119cm | 0.14426 | 0.061294 | 0.04111 | 0.022627 | 0.005 | 0
ΔR 0psi 91cm | 0.023854 | 0.016643 | 0.013454 | 0.0078102 | 0.0084853 | 0
ΔR 14psi 91cm | 0.03265 | 0.020809 | 0.013 | 0.0092195 | 0.0092195 | 0
ΔR 0psi 71cm | 0.027459 | 0.0036056 | 0.0022361 | 0.0036056 | 0 | 0
ΔR 14psi 71cm | 0.014765 | 0.01063 | 0.0070711 | 0.0070711 | 0 | 0
ΔR (cm)
gm
ΔR=Sqrt(Δx2+Δy2)

## Slide 10: Straw creep measurement
Glue straws on a support frame (120cm) with tensions:
  300gm, 400gm, 500gm, 600gm.
Measure straw tension by resonant frequency as a function of time.
Chiho Wang, Seog Oh, Wes Johnson
5/10/12
10
Precision
Function
generator
Frequency
Counter
Weight
Speaker
Magnet
Straw
Support Frame

## Slide 11: Straw creep measurement
(corrected for temperature & humidity variation)
300gm
400gm
500gm
600gm

## Slide 12: Straw creep measurement
Fractional tension shows scaling effect, giving prediction power to other initial tensions.
Assume T(t)/T(0)=(c0*e-c1*t+c2), with c2=0
6-year extrapolation  T(t)/T(0) ~ 43%
700 gm initial tension is sufficient to keep tension > 250 gm after 6 years.

## Slide 13: Radiation aging study
A single straw detector, constructed with relevant components, was operated under Sr-90 irradiation.
Expected total dose (including beam flash) of 0.9 C/cm was irradiated.
Gain change monitored by Fe55 peak amplitude.
Observed no measurable degradation in gain or cathode resistivity
Period | Charge (mC/cm) | Current (nA/cm) | Gain(x)/Gain(110cm) |  | 
 |  |  | p1 (7cm) | p2 (37cm) | p3 (67cm)
1 | 120 | 18 | 0.979 | 0.982 | 0.994
2 | 120 | 36 | 1.010 | 1.009 | 1.002
3 | 670 | 70 | 0.994 | 0.990 | 0.983
Total | 910 |  | ±2% |  | 
Gas flow rate = 2 vol/hr/m
20±0.02% CO2 with balanced Ar
Airgas “primary standard” grade. Purity 99.99% with individual impurities below:
Fractional gain change over 3 irradiation periods. Measured at 3 locations from downstream. Referenced to un-irradiated point at 110cm:

## Slide 14: Straw property summary
Straw is robust and has no problem operating in vacuum. The burst pressure exceeds 60 psi and is well above 14psi.
Straw leak rate of ~<4ccm is within acceptable range of 7ccm
700 gm initial tension is needed to keep straw tension above 250 gm after 6 years, which is needed to keep straw straightness within ±0.3mm for HV stability.
Straw can sustain tension > 1.6Kg and is well above the needed 700 gm initial tension.
No measurable degradation in gain after expected dosage of radiation.

## Slide 15: Straw assembly
After a straw is received from manufacturer it needs to go through QC procedures and get assembled into a straw sub-assembly. These procedures include:
  Visual inspection
  Continuity check (measure resistance)
  Leak test
  Cut to length
  Assemble/Attach end pieces
A QA/QC data base is being developed to track/link component data.

## Slide 16: Straw cutter

## Slide 17: Straw sub-assembly
Sense wire is aligned by external precision fixture (PAAS), not by straw components)
Wire fixation

## Slide 18: Assembly jigs for assembling end pieces

## Slide 19: Assembly jigs for gluing end pieces to straw
Orient end pieces on both end of the straw to the same direction
and glue to straw with conductive glue.

## Slide 20: A 2x4 prototype
Front end electronics attached

## Slide 21: QA/QC data base
RDBMS
  Open source
  Multimaster replication
  Input programming language independence
  MariaDB (formerly MySQL)
GUI

## Slide 22: Straw assembly summary
Straw end components had been manufactured and tested.
Assembly tools and test equipment for straw assembly procedures were developed and functional.
Using the material and tools developed, a 2x4 prototype is completed with front end electronics attached.
A QA/QC data base is being implemented for production process.

## Slide 23: Cost Distribution by L4
Base Cost by L4 (AY $k)
Chart: 
| 475.06.02 Straws Actuals | 475.06.02.01 Straw Tubes | 475.06.02.03 Wire Stringing
 | 124329.44 | 1277429.9406 | 17626.6563

## Slide 24: Cost Distribution by Resource Type
Base Cost (AY $k)
Chart: 
| L Labor | M Material | N Non-Fermi Labor
Total | 64930.0019 | 938258.7938 | 416197.2412

## Slide 25: Quality of Estimate
Base Cost by Estimate Type (AY$k)
Chart: 
| L1 Actual / M1 Existing P.O. | L3 / M3 Advanced | L4 / M4 Preliminary | L5 / M5 Conceptual
Total | 124329.44 | 262348.6549 | 602393.7494 | 430314.1928

## Slide 26: Labor Resources
FTEs by Discipline
Chart: Annual Cost $K
| 42277 | 42643 | 43008
L Labor | 7751.3099 | 0 | 0
M Material | 162471.6778 | 663375.3177 | 33421.9225
Non-Fermi Labor | 4679.6305 | 141187.6733 | 270329.9374
Cumulative Total | 311071.186 | 1115634.177 | 1419386.0369

## Slide 27: Cost Table
Costs are fully burdened in AY $k
 | Base Cost (AY $) |  |  |  |  | 
 | M&S | Labor | Total | Estimate Uncertainty (on remaining costs) | % Contingency on ETC | Total Cost
475.06 Tracker |  |  |  |  |  | 
475.06.02 Straws |  |  |  |  |  | 
475.06.02 Straws | 75 | 49 | 124 |  |  | 124
475.06.02.01 Straw Tubes | 1,262 | 16 | 1,277 | 486 | 38% | 1,763
475.06.02.03 Wire Stringing | 18 |  | 18 | 3 | 15% | 20
Grand Total | 1,354 | 65 | 1,419 | 488 | 38% | 1,908

## Slide 28: Conclusion
Detecting elements of the tracker: straws and straw-end components are manufactured and tested.
Straw tube is robust and leak tight, and exceeds the requirements to operate in vacuum for 6 years.
Assembly procedures and assembly tooling are realized and exercised.
A QA/QC data base is being implemented.
Cost is better estimated.