## Slide 1: WBS 475.04.08 Field Mapping System
Michael Tartaglia
L3 for Magnetic Field Mapping
DOE CD-2 Review
October 21-24, 2014

## Slide 2: Solenoid Team
Mike Tartaglia – L3 Manager for Field Mapping System
  18 years in TD magnet test operations, magnetic measurement and analysis
  TD Test & Instrumentation Department Head since May 1, 2014
Team
  Marc Buehler, Associate Scientist in Magnet Systems Dept.
    And students
  Mau Lopes, Associate Scientist in Magnet Systems Dept.
    And students
  Charles Orozco, Co-op Mechanical Engineer
  Konstantinos Vellidis, Guest Engineer

## Slide 3: Requirements
Field Mapping System Requirements – Docdb#1275
In Summary:
  Verify that the solenoid system meets all magnetic field requirements (Docdb #1266)
  Monitor fields in critical regions of the muon beamline during solenoid operation
  Obtain precision field maps of the DS spectrometer and calorimeter regions at nominal & calibration operating currents

## Slide 4: Requirements
In Greater Detail:
  Map {Bx,By,Bz} along DS bore (0<R<70cm)
    DS-only during commissioning – QA check of DS magnetic field
    Final magnetic configuration; DS at reduced currents (calibration settings)
    Most demanding – precise tracker region requires dB/B~10-4
  Map {Bz,Br} and {dBz/dz} in PS along HRS bore (limited to R<17cm)
    PS-only during commissioning – QA check of PS magnetic field, HRS permeability
    Final magnetic configuration (all solenoids powered)
  Map {Bz} into PS/TS1 and DS/TS5 collimators (limited to R<15cm)
    Final magnetic configuration; TS adjustments and trim power supply settings
  Monitor {B} in DS spectrometer region during Mu2e operation
  Monitor {Bz} and {dBz/dz} along TS1, TS3u, TS3d, TS5 collimators (R~15cm)
    Measure straight section continuous negative gradients, final magnetic configuration
  Perform Electron Source Test (off project)
    Map electron transport through TS, final magnetic configuration

## Slide 5: Design
DS Field Mapper
  Probe positioning mechanism, similar to CMS solenoid mapper
  Axial motion along precise DS detector support rails
  Modern calibrated 3D Hall probes at fixed {Ri} on rotating “propellers”
    Shaft extension to map into TS5 collimator
Z drive: stepper air motor, piezo control valves
toothed drive belt, position encoder
ϑ drive: 270° continuous air motor, geared 360° motion
Angle encoder
Fixed position NMR probes

## Slide 6: Design
DS Field Mapper Status
  3D model, bill of materials ~complete
  Developing preliminary drawings
Field Mapper DAQ and Controls
  New FNAL/TD Magnet Measurement System
    Configurable, Extensible, Modern Software Architecture
    Software framework & component developments are advanced
    Utilize mostly commercial off-the-shelf components
      PXIe, NI, Labview; Metrolab NMR
    Hall and NMR prototype (calibration) system testing started
  Detailed instrumentation design will begin soon
    Readiness for CD-3

## Slide 7: Design
PS Field Mapper
  Hall probes to measure field strength and axial gradient
    Will be calibrated in Tevatron Dipole to 5 T
  Probe positioning mechanism still very preliminary
    Following evolution of PS bore (HRS=heat & radiation shield)
    Utilize linear guide supported by HRS bore and upstream flange
    Cantilevered section downstream of target support ring
      Extends map into TS1 collimator
    External Z-drive system & encoder
    Azimuthal positioning with non-magnetic rotary piezo stage
  Detailed mechanical design to begin soon
  Same Control & Instrumentation as DS
    Tailored interface boxes and separate cabling

## Slide 8: Design
PS Field Mapper
  PS with Heat & Radiation Shield
Space may not be sufficient , ~5-6m long cantilevered arm
[CERN MICE Mapper]
upstream end
flange
d.s. end of TS1 Collimator
PS
HRS
Target support ring
inner bore tube

## Slide 9: Design
In-Situ Sensors
  NMR probes in DS spectrometer region
    Mounted to inner bore (fixed positions to detect field changes)
      In place during field mapping
    Mounted to tracker frame ends (distinguish position changes)
      Re-use probes from DS field mapper
    Same electronics used during DS field mapping
  Hall probes in TS Collimators [Now part of “TS Instrumentation”]
    This is the only viable approach to check TS3 straight regions
    Expect to map TS1 & TS5; these sensors monitor changes over time
    1D Hall probes (& RTDs) embedded in collimators (3 designs)
      after epoxy-impregnated construction in precisely spaced pre-machined slots
    Read out by multiplexed instrumentation system (all magnet sensors)
      Working on wiring and signal feed-through design details

## Slide 10: Design
TS Collimator In-Situ Sensors
TS1 Collimator
TS5 Collimator

## Slide 11: Design
Electron Transport Test
  Off project; design & cost studied under Field Mapping WBS
  Validate transport from primary target region through TS1,3,5 collimators to stopping target in “final magnetic configuration”
    Position source near primary target position
      Few-MeV Beta or Electron Conversion Source
      Re-use PS field mapping device to move the source
      Need components to allow rough evacuation of Muon Beamline
    Detect electrons at downstream positions
      Thin Scintillator tiles with on-board SiPM detectors (& electronics?)
      Study detected position versus start position
        Simulation in progress to define detector segmentation
      Step 1: Detect at TS3 pbar-window gap (test TSu transmission)
      Step 2: remove TS3 detector, detect at DS stopping target
    Detailed studies and simulations in progress

## Slide 12: Changes since CD-1
Production Solenoid/TS1 Collimator
  No longer mapping the PS full aperture
    Heat & Radiation Shield will now be an integral part of PS
    Field mapping only in limited aperture of the HRS bore
    Target support ring ~17 cm radius
  Bkgd Simulations → new pbar window requirement
    upstream of TS1 Col → could impact TS1 mapping
    Looking at timing and method of this material installation
    Electron Source Test could also be affected
In-Situ Hall Probes
  Collimator instrumentation is TS responsibility

## Slide 13: Value Engineering since CD-1
Control and DAQ Readout Same Controls and DAQ Readout System will be utilized for both DS and PS (map one at a time; tailored interface boxes and software configurations)
PS field map Hall Probe positioning device will also be used to position the Electron Source in the e-SourceTest
In-Situ Hall probe readout system will be common with all other magnet sensors
Same NMR probes and instrumentation will be used both for mapping and in-situ monitoring of DS field

## Slide 14: Downselects
Electron Source Test at CD-1 considered using low energy electron gun array, vs radioactive source on positioning stage
  Few MeV Beta or Conversion source has several advantages

## Slide 15: Remaining work before CD-3
Complete preliminary designs
Complete preliminary design drawings
Complete final designs
Complete final design drawings

## Slide 16: Quality Assurance
All mechanical components & materials will be inspected, and tested for magnetic properties (proximity to NMR in calibration magnet) for conformance with specifications
All Hall probes & electronics will be calibrated to the required level over the specified operating range; probe tilt angles and sensitivity to temperature variations will be measured
Multiple probes provide redundancy for cross-checks; NMR probes provide absolute field at reference points
Metrology will be used to determine the precise probe positions: for internal alignment, adjustment, and motion
Entire system will be assembled, bench- and field- tested well in advance of solenoid cryogenic and powered operations
Data analysis tools will be developed and ready to allow rapid evaluation of data quality during the field mapping operations

## Slide 17: Risks
No major risks have been identified with field mapping

## Slide 18: ES&H
ES&H issues associated with Field Mapping System and Operation:
  DS confined space: access to DS bore is required for installation and removal of Field Mapper Z-drive belt
  During cryogenic operations: ODH conditions may apply within regions of the Mu2e Hall
  During magnet powered operations: high magnetic field hazard will apply within some regions of the Mu2e Hall
  Electron Source Test: training in proper handling and storage of radioactive source (beta or electron conversion) will be needed

## Slide 19: Major Milestones
475.04.08 Field Mapping System
  001170 Final Design of Magnetic Field Mapping System Complete
  001310 Vendor for magnetic field mapping system purchased parts selected
  001360 PO issued for magnetic field mapping system purchased parts

## Slide 20: Cost Table
 | Base Cost (AY K$) |  |  |  |  | 
 | M&S | Labor | Total | Estimate Uncertainty (on remaining costs) | % Contingency on ETC | Total Cost
475.04.08 Magnetic Field Mapping System |  |  |  |  |  | 
475.04.08 Magnetic Field Mapping System | 339 | 715 | 1,053 | 448 | 43% | 1,501
Grand Total | 339 | 715 | 1,053 | 448 | 43% | 1,501

## Slide 21: Quality of Estimate
Base Cost by Estimate Type (AY $k)

## Slide 22: Resource Type

## Slide 23: Labor Resources by FY
FTEs by Discipline

## Slide 24: Schedule
Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4 | Q1 | Q2 | Q3 | Q4
FY14 FY15 FY16 FY17 FY18 FY19 FY20 FY21
CD-3a
CD-2/3b
CD-4
Detector Prototypes and Construction
Accelerator and Beamline Construction
Solenoid Infrastructure
Solenoid Installation and Commissioning
KPPs Satisfied
Solenoid Fabrication and QA
Solenoid Design/Prototypes
Accelerator Commissioning (off Project)
CD-3c
Fabricate and QA Superconductor
Detector Hall Construction
Cosmic Ray System Test
FM Final Design
FM Parts Procured
Field Mapping

## Slide 25: Summary
Field Mapping System preliminary design is in progress
  Mapping is not on the critical path
    In-situ sensors needed at end of TS construction
    Other Mapping activities happen during Solenoid commissioning
  DS is the most important, is the most advanced
    Next focus on the daq & controls
  PS Mapper
    mechanical design following DS mechanical completion
Field Mapping System is ready for CD2