## Slide 1: Mu2e Tracker Straws
Chiho Wang
Straws L3 Manager
10/22/2014

## Slide 2: Outline
Straw requirements
Straw property measurements
Straw assembly components and procedures
Cost

## Slide 3: Straw requirements
5 mm ID metalized Mylar tube cathode.
Minimal material
  6 μm Mylar + 3 μm adhesive + 6 μm Mylar double helical wrap
  Inner wall coating: 500Å Al + 200Å Au
  Outer wall coating: 500Å Al
Operates in vacuum
  Sustains > 1 atm pressure difference
  Leak rate < 7 ccm / detector volume
Stability
  Straw straightness: max. transverse deviation/sagging < 300 μm for HV stability.
    Longitudinal tension is applied to keep straw straight.
    Initial tension need to be higher to counter for material relaxation over time (creep).
  Sustains radiation over the life time of operation.

## Slide 4: Straw properties
Destructive tests:
  Pressure: sustained 60 psi (10 min.). Destroyed ≥ 70 psi
  Stretch: sustained 1.6 kg (2 yrs.). Destroyed ≥ 2.9 kg
Mechanical:
  Linear density: 0.34±0.01 g/100cm
  Wall thickness: 15 μm
    Derived from linear density, assuming Mylar density 1.39 g/cm3, and same for the polyester based adhesive.
  Spring constant: 0.891 cm/kg/100cm
Electrical:
  Cathode resistance: 120 Ω/100cm

## Slide 5: Straw leak measurement 1
Pressurize to ~ 1 atm and measure the pressure drop
Duke C. Wang, S. Oh

## Slide 6: Straw leak measurement 1
CO2
N2
N2 leak rate =0.095±0.026 psi/day
CO2 leak rate =0.36±0.06 psi/day
Assume Ar leak rate similar to N2 , and straw volume ~ 0.3 m3, estimated Ar(80%)CO(20%) leak rate ~ 2 ccm / tracker volume
Duke C. Wang, S. Oh

## Slide 7: Straw leak measurement 2
Fill straw with ArCO2 and place in vacuum,
Measure pressure rise
Tested 10×129cm straws
Early results ≲ 4 ccm / tracker volume
Improvements underway
Vacuum pump
Gas inlet
To straw
Vacuum
gauge
1 psi
Pressure relief
Straw
Feed through
York J. Popp, K. Lynch

## Slide 8: Straw leak measurement 3
Fill straw with ArCO2 and place in vacuum,
Measure pressure rise
Tested 119 ×129cm straws
Averaged leak rate = 5·10-5 ccm/straw
~1 ccm for tracker volume
Rice D. Rivera, M. Corcoran
EE891 and EE892 CO2 sensor

## Slide 9: Straw straightness measurement To determine the required straw tension
Horizontal Traveling Microscope (X)
Vertical Traveling Microscope (Y)
Straw
Straw deviation measured for different straw length and tension.
Duke C. Wang, S. Oh

## Slide 10: Straw straightness measurement
Mid point deviation, relative to the highest tension (500g)
Chart: 
| 100 | 150 | 200 | 300 | 400 | 500
ΔR 0 psi 119cm | 0.10242 | 0.029017 | 0.03015 | 0.019799 | 0.0036056 | 0
ΔR 14psi 119cm | 0.14426 | 0.061294 | 0.04111 | 0.022627 | 0.005 | 0
ΔR 0 psi 91cm | 0.023854 | 0.016643 | 0.013454 | 0.0078102 | 0.0084853 | 0
ΔR 14 psi 91cm | 0.03265 | 0.020809 | 0.013 | 0.0092195 | 0.0092195 | 0
ΔR 0 psi 71cm | 0.027459 | 0.0036056 | 0.0022361 | 0.0036056 | 0 | 0
ΔR 14 psi 71cm | 0.014765 | 0.01063 | 0.0070711 | 0.0070711 | 0 | 0
ΔR (cm)
g
Duke C. Wang, S. Oh

## Slide 11: Straw creep measurement
Glue straws on a support frame (120 cm) with tensions:
  300 g, 400 g, 500 g, 600 g.
Measure straw tension by resonant frequency as a function of time.
Chiho Wang, Seog Oh, Wes Johnson
11
Precision
Function
generator
Frequency Counter
Weight
Speaker
Magnet
Straw
Support Frame
Duke C. Wang, S. Oh

## Slide 12: Straw creep measurement
Corrected for temperature & humidity variation
300 g
400 g
500 g
600 g
Duke C. Wang, S. Oh

## Slide 13: Straw creep measurement
Fractional tension shows scaling effect, giving prediction power to other initial tensions.
Assume T(t)/T(0)=(c0*e-c1*t+c2), with c2=0
6-year extrapolation  T(t)/T(0) ~ 43%
700 g initial tension is sufficient to keep tension > 250 g after 6 years.
Duke C. Wang, S. Oh

## Slide 14: Radiation aging study
A single straw detector, constructed with relevant components, was operated under Sr90 irradiation.
Expected total dose (including beam flash) of 0.9 C/cm was irradiated.
Gain change monitored by Fe55 peak amplitude.
Observed no measurable degradation in gain or cathode resistivity
Period | Charge (mC/cm) | Current (nA/cm) | Gain(x)/Gain(110 cm) |  | 
 |  |  | p1 (7cm) | p2 (37cm) | p3 (67cm)
1 | 120 | 18 | 0.979 | 0.982 | 0.994
2 | 120 | 36 | 1.010 | 1.009 | 1.002
3 | 670 | 70 | 0.994 | 0.990 | 0.983
Total | 910 |  | ±2% |  | 
Gas flow rate = 2 vol/hr/m
20±0.02% CO2 with balance Ar
Airgas “primary standard” grade. Purity 99.99% with individual impurities below:
Fractional gain change over 3 irradiation periods. Measured at 3 locations from downstream. Referenced to un-irradiated point at 110 cm:
Fermilab R. Wagner

## Slide 15: Straw property summary
Straw is robust and has no problem operating in vacuum. The burst pressure exceeds 60 psi and is well above 14 psi.
Straw leak rate of ≲4 ccm is within acceptable range of 7 ccm
700 g initial tension is needed to keep straw tension above 250 g after 6 years, which is needed to keep straw straightness within ±0.3 mm for HV stability.
Straw can sustain tension > 1.6 Kg and is well above the needed 700 g initial tension.
No measurable degradation in gain after expected dosage of radiation.

## Slide 16: Straw assembly
After a straw is received from manufacturer it needs to go through QC procedures and get assembled into a straw sub-assembly. These procedures include:
  Visual inspection
  Continuity check (measure resistance)
  Leak test
  Cut to length
  Assemble/Attach end pieces
A QA/QC data base is being developed to track/link component data.

## Slide 17: Straw cutter
Fermilab R. Wagner

## Slide 18: Straw sub-assembly
Sense wire is aligned by external precision fixture (PAAS), not by straw components)
Wire fixation
Fermilab R. Wagner, A. Mukherjee

## Slide 19: Straw Termination
Simulation studies show it is not necessary to deaden straws near gas manifold
Old
4 piece construction
Kapton sleeve trimmed to deaden ~1 cm past manifold
Mostly brass
New
3piece construction
Mostly plastic Lighter & cheaper
Fermilab R. Wagner, A. Mukherjee

## Slide 20: Assembly jigs for assembling end pieces
Fermilab R. Wagner

## Slide 21: Assembly jigs for gluing end pieces to straw
Orient end pieces on two ends
Glue to straw with conductive glue
Fermilab R. Wagner

## Slide 22: A 2x4 prototype
Front end electronics attached
Fermilab R. Wagner, V. Rusu

## Slide 23: QA/QC data base
GUI
Rice M. Corcoran
York K. Lynch
Use a RDBMS FNAL supports PostgreSQL
Barcode users & most parts Straws via their storage tubes
QC and construction steps recorded in DB
Leverage NOvA software experience and barcode readers, but write new software interface

## Slide 24: Straw assembly summary
Straw end components had been manufactured and tested.
Assembly tools and test equipment for straw assembly procedures were developed and functional.
Using the material and tools developed, a 2×4 prototype is completed with front end electronics attached.
A QA/QC data base is being implemented for production process.

## Slide 25: Cost Table
 | Base Cost (AY K$) |  |  | Estimate Uncertainty (on remaining costs) | % Contingency on ETC | Total Cost
 | M&S | Labor | Total |  |  | 
475.06 Tracker |  |  |  |  |  | 
475.06.02 Straws |  |  |  |  |  | 
475.06.02.01 Straw Tubes | 1,170 | 67 | 1,237 | 425 | 38% | 1,662
475.06.02.03 Wire Stringing | 29 | 2 | 31 | 4 | 21% | 35
Grand Total | 1,200 | 69 | 1,268 | 429 | 38% | 1,697

## Slide 26: Cost Breakdown
Base Cost (AY K$)
Chart: 
| 475.06.02.01 Straw Tubes | 475.06.02.03 Wire Stringing
Total | 1237459.5072 | 30961.4503
Chart: 
| L Labor | M Material | N Non-Fermi Labor
Total | 68665.2227 | 807032.8699 | 392722.8649

## Slide 27: Quality of Estimate
Base Cost (AY K$)
Chart: 
| L1 Actual / M1 Existing P.O. | L3 / M3 Advanced | L4 / M4 Preliminary | L5 / M5 Conceptual
Total | 124329.1362 | 208058.8559 | 549697.3513 | 386335.6141

## Slide 28: Labor & Material by FY
Chart: Annual Cost $K
| 42277 | 42643 | 43008 | 43373
L Labor | 6942.9317 | 0 | 0 | 0
M Material | 107246.6147 | 5449.8666 | 579356.0058 | 27498.0297
Non-Fermi Labor | 463.2361 | 25300.2027 | 142056.3619 | 224903.0642
Cumulative Total | 263857.4266 | 294607.4959 | 1016019.8636 | 1268420.9575
Base Cost (AY K$)

## Slide 29: Labor Resources by FY
FTEs by Discipline
Chart: Annual FTE's
| 42277 | 42643 | 43008 | 43373
AD Administrative | 0.0237 | 0 | 0.0836 | 0
EN Engineering | 0 | 0 | 0 | 0
ES Environmental, Safety & Health | 0 | 0 | 0 | 0
FM Facilities Management
IT Information Technology
SC Scientific | 0.1832 | 0.1026 | 3.1402 | 4.6988
TE Technical | 0.0438 | 0.1981 | 0 | 0
Cumulative | 0.2507 | 0.5514 | 3.7752 | 8.474

## Slide 30: Conclusion
Detecting elements of the tracker: straws and straw-end components are manufactured and tested.
Straw tube is robust and leak tight, and exceeds the requirements to operate in vacuum for 6 years.
Assembly procedures and assembly tooling are realized and exercised.
A QA/QC data base is being implemented.
Ready for CD-2