## Slide 1: Mu2e Tracker Infrastructure
A. Mukherjee
Tracker L2 Manger Acting Infrastructure L3 manager
22 Oct 14

## Slide 2: Cooling
10kW tracker + 5kW calorimeter → 15kW SUVA system
  Passive distribution by equalizing line lengths
  Tap points for calorimeter cooling not shown

## Slide 3: Cooling
One ring per plane
VCO disconnect for removing planes
±2ºC at OD

## Slide 4: Cooling
Temperature gradient with full electronics load
+5C

## Slide 5: Power
Reduce magnetic field perturbations and power loss in lines by sending power at 48V
Buck DC-DC convertor with air-core toroid
  Tested to 36V with hand-wound toroid
  Developing 48V version with toroid manufactured on PCB
Chart: Input Voltage (V)
| 11.2 | 13.2 | 14.5 | 15.9 | 17.2 | 19.899999999999999 | 21.5 | 23.2 | 25.2 | 26.8 | 28.5 | 29.8 | 30.8 | 31.1 | 32.700000000000003 | 33.4 | 34.1 | 35.1
Iin (A)
Efficiency | 0.92263 | 0.92099 | 0.8944 | 0.93217 | 0.928 | 0.86893 | 0.8809 | 0.81635 | 0.83003 | 0.78048 | 0.81547 | 0.77989 | 0.75457 | 0.84071 | 0.79638 | 0.78594 | 0.76981 | 0.74788

## Slide 6: Alignment Monitoring: Hall Probe
Not for absolute alignment: monitor variation over time
2-axis Hall probe aligned perpendicular to magnetic field
20 µradian resolution with Bz=1Tesla field
Chart: x -coil Current (A) ~200G/A
| 0.11 | 0.2 | 0.3 | 0.4 | 0.5 | 0.61 | 0.7 | 0.8 | 0.9 | 1
 | 18.226 | 40.062 | 62.958 | 90.306 | 113.626 | 136.946 | 157.51 | 180.194 | 205.21 | 224.502

## Slide 7: Cabling
Horizontal beams support cabling, gas lines on tracker
Cables run past calorimeter and beam stop to IFB
Azimuthal position set to minimize interference with calorimeter electronics
Leave room to sight panel survey monuments
Communication
Power
Gas
High Voltage

## Slide 8: Cabling
Gas
  One pair per plane
  IFB penetration: Standard welded feedthroughs
  Panel penetration: stainless steel compression fitting with epoxy
Low voltage
  One pair per plane
  In vacuum: Nomex® covered ~2.5mm square magnet wire Similar to that used by CDF
  Outside vacuum: THHN (flexible) wire
  IFB penetrations: standard electric vacuum feedthroughs
  Panel penetrations: Copper with Kapton® sleeve

## Slide 9: Cabling
High Voltage
  One line per panel
  Silicone insulated, 0.05” pitch ribbon cable
  IFB penetration: Vacuum rated DB25
    Note all lines are at the same voltage
    Panel penetrations: Copper with ceramic sleeve
Copper communication
  Two coax per panel
  IFB and Panel Penetration: Standard vacuum rated SMA

## Slide 10: Cabling
Optical Communication
  One pair per panel
  Jacket but no fiber reinforcement
    Reinforcement traps air and forms a virtual leak
  IFB penetration: standard vacuum rated fiber feedthrough
  Panel penetration: individual fibers with epoxy seal

## Slide 11: Outgassing from Cables
Optical fibers tested for outgassing
  Reinforcing fibers normally used trap too much air
  CERN’s unreinforced optical cable gives acceptable results
    ~8·10-4 ccm/370mm for 12-fiber ribbon after pumping for 48 hours
    <4·10-4 ccm/370mm after pumping ~1 week
    Worst case extrapolation (linear in length)
      <1.0 ccm in 2 days
      <0.5 ccm in a week

## Slide 12: Electrical Breakdown (Townsend Effect)
Our residual gas is (probably) mostly water absorbed into plastics… not air
Tested for breakdown with just water
  Standard solder tail DB25 (female) connector, being considered for HV penetration at IFB
  Coat solder connections with acrylic (nailpolish)
  Connector side left open
  Withstands 1.8kV up to ~0.1 Torr
No problem operating at nominal DS pressure of 10-4 Torr

## Slide 13: Cost Table
 | Base Cost (AY K$) |  |  | Estimate Uncertainty (on remaining costs) | % Contingency on ETC | Total Cost
 | M&S | Labor | Total |  |  | 
475.06 Tracker |  |  |  |  |  | 
475.06.05 Tracker Infrastructure |  |  |  |  |  | 
475.06.05.01 Gas System | 8 | 85 | 93 | 44 | 51% | 137
475.06.05.02 High Voltage System | 12 | 64 | 76 | 33 | 43% | 109
475.06.05.03 Low Voltage System | 65 | 91 | 156 | 70 | 45% | 226
475.06.05.04 Cooling | 129 | 222 | 351 | 142 | 41% | 493
475.06.05.05 Cabling & Control | 160 | 104 | 265 | 63 | 25% | 327
Grand Total | 374 | 567 | 941 | 352 | 38% | 1,293

## Slide 14: Cost Breakdown
Base Cost (AY K$)
Chart: 
| 475.06.05.01 Gas System | 475.06.05.02 High Voltage System | 475.06.05.03 Low Voltage System | 475.06.05.04 Cooling | 475.06.05.05 Cabling & Control
Total | 92691.4786 | 76374.3246 | 156155.4489 | 350616.5167 | 264838.3254
Chart: 
| L Labor | M Material | N Non-Fermi Labor
Total | 567002.5341 | 366948.1367 | 6725.4234

## Slide 15: Quality of Estimate
Base Cost (AY K$)
Chart: 
| L1 Actual / M1 Existing P.O. | L3 / M3 Advanced | L4 / M4 Preliminary | L5 / M5 Conceptual
Total | 16768.7997 | 35070.328 | 448016.8822 | 440820.0843

## Slide 16: Labor & Material by FY
Base Cost (AY K$)
Chart: Annual Cost $K
| 42277 | 42643 | 43008 | 43738
L Labor | 113704.2022 | 102719.7142 | 32664.4541 | 263918.248
M Material | 10000.0584 | 0 | 353542.1592 | 7924.482
Non-Fermi Labor | 0 | 0 | 0 | 0
Cumulative Total | 179907.0367 | 282626.7509 | 668833.3642 | 940676.0942

## Slide 17: Labor Resources by FY
FTEs by Discipline
Chart: Annual FTE's
| 42277 | 42643 | 43008 | 43738
AD Administrative | 0.0011 | 0 | 0.0495 | 0
EN Engineering | 0.2532 | 0.485 | 0.0622 | 0.5433
ES Environmental, Safety & Health
FM Facilities Management
IT Information Technology
SC Scientific | 0.0226 | 0 | 0.146 | 0.1811
TE Technical | 0.3018 | 0 | 0.1132 | 0.7925
Cumulative | 0.5787 | 1.0637 | 1.4346 | 2.9515

## Slide 18: Summary
Preliminary design exists for key subsystems: cooling, power
Conceptual design exists for all other subsystems
Outgassing studies in progress
Vacuum penetration design will be done at York
  Vacuum equipment and expertise available
Fermilab engineering becomes available as other L3s move into production