## Slide 1: RFD Cold Test
300 K
0 K
~10 days
Cavity Test at 2K:
Cavity sent under vacuum
  Cavity not vented at CERN.
Performance @ CERN
  Q0 = 7.5 x109
  RResidual = 16 nOhm
  Max VT = 6.2 MV
Performance @ ODU/JLAB
  Q0 = 1.25 x 1010
  RResidual = 9.5 nOhm
  Max VT = 7.0 MV
Comments
  Multipacting processable
  No issues
Thermal cycle to 15 K

## Slide 2: DQW Cold Test
300 K
0 K
~10 days
Cavity Test at 2K:
RF surface contaminated
  HPR required before test
Performance @ CERN
  Q0 = 3.5 x109
  RResidual = 24 nOhm
  Max VT = 4.0 MV
Performance @ BNL
  Q0 = 4.5 x 109
  RResidual = 18.3 nOhm
  Max VT = 4.5 MV
Comments
  Results compromised by cable breakdown on input power cable in cryostat
    => retesting obligatory

## Slide 3: UK-4Rod Cold Test
Cavity Tests at 4K, 2K & 1.8K
Issues with the cavity
  Bad Cavity Vacuum at all temperatures
  NbTi Flanges with very poor knife edges
  Niobium porosity on weld at base of rods at 2K
  Insufficient HPR at time of test
Performance @ CERN
  Q0 = ~1.0 x109
  RResidual = 38 nOhm
  Max VT = 3.3 MV
Comments
  Cavity flanges needed repair

## Slide 4: Other observations
Frequency change with temperature
  DQW:
    Δf = 323 kHz from 4.5 το 2.0K
  RFD:
    Δf = 311 kHz from 4.1 το 2.0K
Low field Multipacting
  Easily processable
    RFD @ 0.54- 0.6 MV
    DQW @ 0.01 -0.1 MV
DQW
RFD

## Slide 5: Crab Present Status
SM18 Tests stand:
  SM18_V3 operational & shared between Crab and High Gradient programs
  New insert required for DQW + tuner test: presently under manufacture
  Tests slot availability dictated by compensatory measures for radiation protection, cryogenics availability, and clean room availability/readiness
RFD:
  PoP Cavity returned to ODU without venting.
DQW:
  PoP Cavity vented, welds etc checked with X-rays + ultrasound
  PoP Cavity: Prepared for light BCP -> 1st available slot = week 21
  Repeat test of PoP Cavity: HPR on week 22, cold test on weeks 23-24
  Cold test with tuner: High Priority -> Awaiting new insert -> Aim for testing in July
UK-4Rod
  Beam port flanges replaced. BCP on week 22, then await test slot
SM18 Horizontal Bunker
  Preparations for horizontal test starting: Cryo distribution box for M7 bunker now under design study, with initial planning estimate expected beginning of June.

## Slide 6: SM18_V3 Insert
Upgrade of insert needed for DQW +tuner test
  Required changes: Tuner motor integration, resized rupture disks, insert load balance, improved pumping line connection
Bare Cavity
Cavity + Tuner
Existing Insert