## Slide 1: Status of 650 MHz cavities for PIP-II
Anna Grassellino
PIP-II Machine Advisory Committee
9-11 March 2015

## Slide 2: Outline
Processing/testing infrastructure progress/upgrades
Cavity high Q surface processing results
  120C bake versus N doping
  Flux expulsion/trapping cooling studies
  Also, vendor qualification (PAVAC)
Cavity mechanical design status
Plan forward

## Slide 3: 650 MHz testing/processing infrastructure
120-300C ovens—ready and in use
High temperature furnace – ready and in use
EP, HPR tools, ready for 1-cell, not for 5-cell 650 MHz
Work ongoing on upgrading to multicell processing (considering AES for bulk chemistry) – longer term solution for dedicated HPR infrastructure needed to avoid conflicts PIP-2/LCLS-2

## Slide 4: 650 MHz Processing/testing infrastructure
Fluxgate magnetometers
H
VTS Ti four bar frame
VTS1—ready for 1-cell 650 MHz, fully instrumented
VTS2,3—commissioned, ready for 5-cell 650 MHz

## Slide 5: Summary of accomplishments – cavity surface processing development
Goal: establish surface processing recipe for Q maximization on single cell in FY15 – crucial for making the linac capable of operating CW or high duty factor
Implement by end of FY15 on two multicells
Six high beta 650 MHz single cell cavities processed – 3 PAVAC (also for vendor qualification purposes), 3 AES
Four tested – 1 baseline standard ILC/XFEL recipe (120C bake), 3 high Q recipe development (N doping)
Exceeded specifications a factor of 3.5 with N doping!
Cooling studies for setting mag shielding/cooling specifications in cryomodule ongoing

## Slide 6: N Doping – small variation from standard ILC treatment
Example from FNAL 2/6 doping process:
Bulk EP
800 C anneal for 2 hours in vacuum
2 minutes @ 800C nitrogen diffusion
800 C for 6 minutes in vacuum
Vacuum cooling
5 microns EP
Cavity after Equator Welding
EP 140 um
Ethanol Rinse
External 20 um BCP
Short HPR
800C HT Bake
RF Tuning
EP 40 um
Ethanol Rinse
Long HPR
Final Assembly
Long HPR
Helium Tank Welding Procedure
VT Assembly
HPR
HOM Tuning
Ship to DESY
Leak Check
120C bake
XFEL
X
A. Grassellino et al, 2013 Supercond. Sci. Technol. 26 102001 (Rapid Communication)

## Slide 7: Results – highlights – 120C bake versus N doping Q~ 7e10 at 2K, 17 MV/m – world record at this frequency!
Applying N doping to 650 MHz (beta=0.9) leads to Q x3 exceeding specs
High duty factor operation (30%) may be possible even with the existing (limited) capacity cryoplant!

## Slide 8: Optimal cavity cooling studies
Cooldown rate and thermal gradients around Tc have a drastic impact on the Meissner effect and:
Can drastically degrade residual resistance due to trapped magnetic flux
Can be used to achieve low residual resistances even in high ambient fields
Thermoelectrically induced mag fields could contribute to degrade residual
A. Romanenko, A. Grassellino, O. Melnychuk, D. A. Sergatskov, J. Appl. Phys. 115, 184903 (2014)
A. Romanenko, A. Grassellino, A.Crawford, D. A. Sergatskov, Appl. Phys. Lett. 105, 234103 (2014)
D. Gonnella et al, J. Appl. Phys. 117, 023908 (2015)

## Slide 9: Need to cool down fast through Tc to preserve Q
9
Dressed N doped nine cell cavity vertical test at T=2K

## Slide 10: Magnetic probes reveal the new physics
Full expulsion of the magnetic field should increase the field at equator ~2 times when going superconducting
Fluxgate
magnetometers
H
Efficient flux expulsion
Poor flux expulsion
It turns out the expulsion efficiency can be controlled by the cooldown procedure through Tc=9.2K (fast/slow, uniform or not)
Same Meissner behavior for EP, EP+120C, N doping, fine/single grain, cooling is what matters

## Slide 11: Results on 650MHz – cooling studies – fast vs slow cooling
Cooling details have been shown to play an important role in Q retention
Slow cooling through Tc shown to severely deteriorate Q for 1.3 GHz cavities
At 650 MHz we find a weaker effect, likely to smaller impact of trapped flux at lower frequency
Very promising for Q retention in cryomodule
T=2K
A. Romanenko, A. Grassellino, O. Melnychuk, D. A. Sergatskov, J. Appl. Phys. 115, 184903 (2014)
A. Romanenko, A. Grassellino, A.Crawford, D. A. Sergatskov, Appl. Phys. Lett. 105, 234103 (2014)

## Slide 12
RF design is completed;
Concept mechanical design for CW is finished (2013);
Lorentz Force Detuning (LFD) is an issue in a pulsed mode:
Active resonance control (W. Schappert)
“Passive” compensation - modification of the cavity and the He vessel design in order to reduce LFD: work is in progress
*I. Gonin, T. Khabiboulline
Sensitivity to He pressure, df/dP, Hz/Torr | <25
LFD sensitivity, df/dE2, Hz/(MV/m)2 | -0.5
LFD at nominal voltage (FRS), Hz | -136
Half-bandwidth, Hz | 29
Status of mechanical design of HB 650 dressed cavity*:
Modification of the cavity in order to reduce LFD.
Upper Ring R=125mm
Stiffness ~10 kN/mm
Lower Ring R=85 mm
LFD ~ - 0.31 Hz/(MV/m)^2

## Slide 13
Name | Value | Unit
Density | 5700 | kg/m^3
Young's modulus | 68e9 | Pa
Poisson's ratio | 0.33 | 1
Name | Value | Unit
Density | 4540 | kg/m^3
Young's modulus | 117e9 | Pa
Poisson's ratio | 0.37 | 1
Name | Value | Unit
Density | 8700 | kg/m^3
Young's modulus | 118e9 | Pa
Poisson's ratio | 0.38 | 1
Niobium
Titanium
Niobium-
Titanium
1030mm
450mm
FIXED
Chart: df/dP, Hz/mbar vs Tuner Stiffness
| 0.0 | 5.0 | 10.0 | 20.0 | 40.0 | 60.0 | 80.0 | 100.0
0.92 Rend=Rreg=110 | -8.7 | 3.1 | 6.9 | 9.9 | 11.9 | 12.6 | 13 | 13.3
SPRING CONSTANT
Ivan Gonin
Reducing Sensitivity to Helium Pressure fluctuations, HB cavity
Design goal |df/dP| < 25 Hz/mbar is achieved

## Slide 14
Reducing Lorenz Force Detuning, HB cavity
R2 = 110 mm, R1 = 83.7 mm K Lfd = -0.277 Hz/(MV/M^2)
dL~1.3μm
dL~0.5μm
LFD ~ -0.91 Hz/(MV/m)^2
Tuner Stiffness 100kN/mm
Fixed
Fixed
Ivan Gonin
HV stiffness needs to be increased to reach |LFD| < 0.6 Hz/(MV/m)^2

## Slide 15
Chart: 
| 0.0 | 10.0 | 20.0 | 40.0 | 100.0
R=125mm | -78.4 | 4.6 | 11.5 | 13.8 | 15.6
R=120mm | -155.8 | -5.3 | 0 | 7.9 | 11.9
df/dP in LB Cavity for PIP II
Bellow radius
Ivan Gonin
fixed
fixed
Evgeniy Zaplatin
Design goal |df/dP| < 25 Hz/mbar can be achieved

## Slide 16
LFD in LB Cavity for PIP II
R1
R2
Chart: LFD (Hz/(MV/m^2) vs. St. Ring R2/RCavity
| 0.3 | 0.35 | 0.4 | 0.45 | 0.5
1 Ring | -1.15 | -0.69 | -0.68 | -0.96 | -1.6
2 Rings | -0.52 | -0.47 | -0.46 | -0.51 | -0.7
CAVITY Stiffens ~2.7 kN/mm
fixed
fixed
Evgeniy Zaplatin
Ivan Gonin
|LFD| < 0.6 Hz/(MV/m)^2 can be achieved

## Slide 17: Plan forward
Driving milestone: pCM in FY18
Implement high Q best treatment by end of FY15 on two multicells:
  Adapt ANL EP and HPR tools to process multicell, consider using AES for EP
  Work on long term solution for having dedicated processing facilities for production
  Finalize RF and mechanical designs, procure cavities
In FY16, goal is to process 8 multicells to be lined up for dressing
In FY17 horizontal testing program, specify cooling procedure and magnetic shielding requirements, line up cavities for pCM

## Slide 18: Conclusions
Cavity processing recipe producing 3.5 times higher quality factor than baseline specifications reproducibly demonstrated on single cells
This translates into an increase to 30% duty factor “for free”, and possibility to operate the linac CW
Q preservation via cooling studies ongoing with very promising results
RF designs of both LB650 and HB650 cavities are completed
Concept mechanical design for CW of HB650 is finished (2013)
Sensitivity to Helium Pressure fluctuations satisfy requirements
Lorentz Force Detuning (LFD) is an issue in a pulsed mode
“Passive” compensation - modification of the cavity and the He vessel design in order to reduce LFD: work is in progress