## Slide 1
Joe Ozelis
PIP-II Risk Workshop
Thursday 4/6-4/7/2017
PIP-II Project Risk: SRF Infrastructure

## Slide 2: SRF Infrastructure Risks
Loss of SRF Cavity Processing Equipment
Significant HF Acid Spill or Exposure
HTS-2 Fails to Meet Performance Specifications
IB1 Cryoplant Failure

## Slide 3: Risk: Loss of SRF Cavity Processing Equipment
Cavity processing equipment (BCP, EP, HPR, Furnaces) is critical to preparation of SR cavities for testing and CM Assy.
  Expensive, unique, and complex facilities generally not identically duplicated on site, and where redundancy exists, it is often of limited capacity
Probability
  10%
Technical Impact
  Significantly Substandard
Risk Rank
  2 (Medium)
Risk Cause or Trigger
  A failure of a system or component of the facility/tooling can lead to system inoperability or diminished performance

## Slide 4: Risk Summary : Threat
Risk Owner: Allan Rowe
Summary : If various items of SRF cavity processing equipment (includes both production and testing) fail, then throughput in support of PIP-II program will suffer as work is redirected to back-up systems, except in cases where no substitute is available, such as the largest vacuum oven at MP9 (must repair).
Risk Effect:
  Schedule slippage due to reduced (or zero) throughput of facilities required to process cavities
  Cost associated with replacement/repair of equipment

## Slide 5: Risk: Cost Impact and Estimate
Cost Impact: (Two-Point)
  25k$-300k
Schedule Impact
  1mos - 6mos
Estimate:
  Cost impact varies based upon mode of failure. An isolated component failure may require replacement at a cost of 5-25k$, while failure of a larger element of a facility (EP frame, Furnace hot-zone, HPR pump) which is either a long-lead time procurement or a custom component (not off-the-shelf) would incur replacement costs of 100’s of k$
Quality of Risk Estimate
  Based upon recent operational experience with LCLS-II, other SRF projects, and known system component costs.

## Slide 6: Risk Mitigation
Routine maintenance of systems to reduce probability of failures.
Maintain adequate training of operations personnel to ensure facilities operated in compliance with design capabilities
Generate additional redundancies where possible (constrained by available resources within SRF Program)
Start/End Date:
  Ongoing through the duration of the project.
Monitoring Required:
  Continuous by the CM CAM/Production Managers/Risk Owner

## Slide 7: Risk: Significant HF Acid Spill or Exposure
Cavity surface processing requires the use of hazardous (deadly) materials (HF acid) for which strict process control measures, engineering protections, and comprehensive personnel training have been implemented in order to reduce probabilities of exposure to personnel and facilities.
  Any exposure to HF is potentially deadly and requires immediate emergency medical treatment/intervention.
  If an exposure were to occur, operations will be suspended until an investigation is completed and mitigations established, delaying cavity processing for up to 6 months.
Probability
  0.5%
Technical Impact
  Slightly Substandard
Risk Rank
  2 (Medium)
Risk Cause or Trigger
  A spill or otherwise uncontrolled release of HF acid that either leads to or has the potential to lead to personnel exposure to HF

## Slide 8: Risk Summary : Threat
Risk Owner: Allan Rowe
Summary : ​If a significant hydrofluoric acid spill occurs resulting in a project-wide processing operations shutdown, then SRF cavity processing activities may be delayed up to six months at the location of the spill or exposure while an investigation is pursued and corrective measures are put in place.
Risk Effect:
  Cavity preparation and testing of all FNAL/ANL processed cavities for the R&D and CM Production tasks is impacted and delayed. This impacts SSR1 (2), SSR2 R&D and Production CMs (1-7), LB650 R&D, and HB650 R&D and CMs (1-4). FNAL responsible for 1/2 of all cavities.

## Slide 9: Risk: Cost Impact and Estimate
Cost Impact: (Triangular)
  100k$ - 250k$ - 500k$
Schedule Impact
  3mos - 6mos
Estimate:
  Cost range is based upon potential range of mitigation measures. Minimum cost represents cleanup, decontamination, and disposal of exposed surfaces/components, maximum range represents implementation of additional engineered safety measures or systems.
Quality of Risk Estimate
  Based upon previous minor incidents that have caused processing activities to be suspended at FNAL/ANL and DOE/FNAL ES&H policies and procedures that govern enterprise response to incidents of a more major nature.

## Slide 10: Risk Mitigation
​Maintain trained personnel and updated process documentation that control chemical operations procedures and handling requirements for the duration of the PIP-II project.
Develop redundant chemical processing capability to the extent possible as ANL is the lone entity planned perform all PIP-II chemical processing tasks.
Routine maintenance of systems to reduce probability of failures.
​Develop full chemical processing redundancy at FNAL in advance of the PIP-II production. This includes BCP of SSR cavities as well as completion of the EP tool EP capabilities of B.61 and B.92 E-cell cavities (estimated cost ~1M$).
Start/End Date:
  Now through 6/20/2024
Monitoring Required:
  Continuous by the CM CAM/Production Managers/Risk Owner

## Slide 11: Risk: HTS-2 Fails to Meet Specifications
The HTS-2 Facility consists of a cryostat/feedcan and associated systems to provide for testing of dressed (cavity + coupler + tuner) cavities at 2K. This facility is required to perform Design Verification activities during the R&D stage of 650MHz cavity development, and also for Production Qualification testing of cavities to be assembled into cryomodules.
  Failure of the facility (cryostat) to meet design specifications (e.g., static heat load, stable operating temperature/cryogen level, ability to dissipate cavity power) could prevent successful testing of dressed cavities
  Failure of associated systems (RF amplifier, controls, interlocks, etc.) could prevent operation of the facility and successful testing of dressed cavities.
Probability
  10%
Technical Impact
  Slightly Substandard
Risk Rank
  2 (Medium)
Risk Cause or Trigger
  ​Inability to effectively commission FNAL's HTS-2 cryostat and associated systems at RRCAT or damage or other problems in transit or upon re-integration at FNAL lead to inadequate performance when system is re-commissioned at FNAL.
  Failure of a supplied component or subsystem after delivery and commissioning at FNAL

## Slide 12: Risk Summary : Threat
Risk Owner: Joe Ozelis
Summary : ​If the HTS-2 (integrated facility) as delivered fails to meet performance specifications, then 650MHz dressed cavities can not be tested, delaying 650MHz Design Verification or Production Cavity Qualification activities. The cryostat may require repair/rework or components of the associated systems may require repair/replacement.
Risk Effect:
  ​Testing of dressed 650MHz cavities for 650MHz Design Verification activities or Production Cavity Qualification is delayed. This in turn may delay timely procurement of qualified components for CM fabrication or may delay CM assembly.
  May require procurement of components with untested performance or assembly of cryomodules with unqualified cavities, leading to potential CM performance failure.

## Slide 13: Risk: Cost Impact and Estimate
Cost Impact: (Two-Point)
  50k$ - 1,000k$
Schedule Impact
  3mos -18 mos
Estimate:
  Cost range is based upon potential range of mitigation measures. Minimum cost represents moderate repair or re-work of cryostat performed at FNAL. Maximum represents replacement of vital components provided by external collaborators (e.g., RF Source, controls/interlock hardware, cryostat/feedcan). Schedule range is based upon minor cryostat repairs(3mos) to procurement of RF power components or replacement cryostat (18 mos)
Quality of Risk Estimate
  Cost and schedule estimates are based upon experience with similar repairs or procurements. Probability estimate is based upon general knowledge of collaborator capabilities and collaborator procured or supplied equipment.

## Slide 14: Risk Mitigation
​​Development of modified STC cryostat to accept 650MHz cavities and commissioning of existing 650MHz IOT RF source, for testing - this allows testing activities to proceed at FNAL on time though at a lower rate (since cryostat is shared with SSR1 program).
Careful oversight of commissioning activities at RRCAT
Start/End Date:
  4/30/2019 – 9/30/2024
Monitoring Required:
  Continuous by HTS-2 SPM

## Slide 15: Risk: IB1 Cryoplant Failure
The IB1 cryoplant provides cryogens (primarily LHe) to 3 cavity vertical test stands (VTS1, VTS2, VTS3)and the solenoid test stand (Stand 3). These test stands are used to test SRF cavities as part of the R&D (Design Verification) and Production Qualification activities. The cryoplant provides 4 critical services : cryogen liquefaction, cryogen transfer, cryogen purification, and sub-atmospheric return.
  Loss of any one of these services severely impacts the ability to test SRF cavities and SC solenoids for CMs.
  This impact could manifest itself as a slowdown or suspension of testing activities
Probability
  15%
Technical Impact
  Slightly Substandard
Risk Rank
  2 (Medium)
Risk Cause or Trigger
  ​​A major component or system of the IB1 cryoplant fails (e.g., compressor, Kinney pumps, PLC controls, valve/distribution box), leaving the facility unable to produce, provide, or purify LHe.

## Slide 16: Risk Summary : Threat
Risk Owner: Joe Ozelis
Summary : ​The cryoplant provides 4 critical services to the VTS and VMTF : cyrogen liquefaction, cryogen transfer, cryogen purification, and sub-atmospheric return. Loss of any of these services leads to a suspension or slowdown in cavity and solenoid testing, impacting the schedule of SRF R&D activities (design verification) or production cavity qualification (and hence CM assembly rate).
Risk Effect:
  ​​VTS testing of SSR1, SSR2, LB650 and HB650 R&D cavities and solenoids. This can impact completion of related Design Verification activities.
  Production testing of some fraction of SSR1, SSR2, LB650 and HB650 cavities and solenoids. This can impact cryomodule assembly schedule

## Slide 17: Risk: Cost Impact and Estimate
Cost Impact: (Three-Point)
  25k$ - 600k$ - 1,250k$
Schedule Impact
  1 mos - 9 mos
Estimate:
  The high point of the cost estimate is based upon having to use industrial vendors to supply LHe for cavity tests, at a reduced rate of 2 tests/week (25k$/test for LHe). This assumes venting (loss) of the LHe after a test is completed due to inability to store/purify gas if purification system is down or storage is already at capacity.
Quality of Risk Estimate
  Risk probability has been estimated at 15% - this is based upon past experience where IB1 cryoplant has experienced one unscheduled maintenance/repair shutdown of ~ 4 weeks duration every 6-8 months.
  Costs for vendor-supplied LHe are based on recent actual costs and LHe usage for same scenario

## Slide 18: Risk Mitigation
​​​Routine maintenance of components and replacement of components nearing end of life-cycle.
Operations consistent with all documented procedures
Note : TD has identified specific risk factors and mitigations for IB1 cryoplant failure modes. See risks : RT-OPS-TD-005 thru RT-OPS-TD-009. The facility is “owned” by TD who are responsible for maintaining operability and performing maintenance/repair when identified risks occur.
Start/End Date:
  4/5/2017 – 9/30/2025
Monitoring Required:
  Continuous by CM CAM, Production Manager, Facility Manager, etc.