## Slide 1: Aerogel Beam Line Cherenkov Detector for the LArIAT Experiment
Brandon Soubasis
Texas State University
Advisors: Jennifer Raaf, Will Flanagan, Karol Lang
August 4, 2015
Brandon Soubasis
Tues August 04, 2015

## Slide 2
Brandon Soubasis
Liquid Argon Time Projection Chambers(LArTPCs)
LArIAT is a LArTPC at Fermilab Test Beam Facility that provides beams of known energies, selection of different particle types, and controlled environment in which to tune simulations and to develop tools for particle identification (PID), calorimetry, and event reconstruction without relying solely on simulation.
Tues August 04, 2015

## Slide 3
Brandon Soubasis
Particle ID
We currently have 2 time of flight scintillator counters which give discrimination for particles but not for to close in mass such as p/π, but not π/μ.
Tues August 04, 2015

## Slide 4
Brandon Soubasis
Particle ID
We currently have 2 time of flight scintillator counters which give discrimination for particles but not for to close in mass such as p/π, but not π/μ.
Tues August 04, 2015

## Slide 5
Brandon Soubasis
Motivation
How can we measure stopped muons and pions at LArIAT?
- Given simple assumptions, muons should stop within the LAr TPC with p<280 MeV and pions with p<330 MeV.
Tues August 04, 2015

## Slide 6
Brandon Soubasis
Motivation
How can we seperate stopped muons and pions at LArIAT?
- Can’t use a muon range stack since these muons will be stopping in the LAr TPC. - Time of Flight (ToF) can be difficult since muon and pion masses only differ by 30%
-What about a material with a small index of refraction (Aerogel)?
Tues August 04, 2015

## Slide 7
Brandon Soubasis
Aiming to separate muons and pions in a momentum range where muons emit Cherenkov radiation while pions do not.
Different indices of refraction are sensitive to different momentum ranges.
The combination of the two aerogel Cherenkov counter, pions and muons can be identified for p<400MeV/c
 |  | n = 1.10 | n = 1.05
220<p<285 | μ | ✓ | x
 | π | x | x
300<p<400 | μ | ✓ | ✓
 | π | ✓ | x
Tues August 04, 2015

## Slide 8
Brandon Soubasis
Designs
We have our aerogel tiles horizontal instead of vertical.
This increases our thickness for particles to pass through from 6.5 to 7.4 inches
Currently using the EMI 9954B 2” PMT with another 3” PMT.
Active
Tues August 04, 2015

## Slide 9
Brandon Soubasis
PMTs
Calibrated PMTs using LED pulses.
This allows us to translate from ADC counts to NPE
Tues August 04, 2015

## Slide 10
Brandon Soubasis
Due to the size of the new design(s), our counter is placed DS just after the KEK counter.
Tues August 04, 2015

## Slide 11
Brandon Soubasis
TOF vs Pulse Area – Positive Magnet Polarity Runs - with pile up
TOF vs Pulse Area – Negative Magnet Polarity Runs - with pile up
Can better understand our pile up from the aerogel counters?
Tues August 04, 2015

## Slide 12
Brandon Soubasis
Tues June 09, 2015
TOF vs Pulse Area – Positive Magnet Polarity Runs
Efficiency of detecting pion:
- HMMS PMT: 57%
- PTN PMT: 41%
- KEK1 PMT:67%
- KEK2 PMT: 82%
Efficiency of detecting proton:
- HMMS PMT: 22%
- PTN PMT: 29%
- KEK1 PMT: 5%
- KEK2 PMT: 2%

## Slide 13
Brandon Soubasis
Tues June 09, 2015
TOF vs Pulse Area – Negative Magnet Polarity Runs
Efficiency of detecting pion
- HMMS PMT: 52%
- PTN PMT: 42%
- KEK1 PMT: 69%
- KEK2 PMT: 74%
Efficiency of detecting proton
- HMMS PMT10%
- PTN PMT: 13%
- KEK1 PMT: 9%
- KEK2 PMT: 15%

## Slide 14
Brandon Soubasis
Reference
E.H.Bellamy et al. NIM A339 (1994) 468-476
P.Adamson et al. NuMI-L-661 (2000)
R.Godang et al. MUC-NOT-COOL-EXP-304
Tues June 09, 2015

## Slide 15
Brandon Soubasis
Other Involvement here at Fermilab
US Wall Halo Scintillator Counter(HSC) Installation
Modified Collimator Running and simulations
Studies with Silicon photomultiplier (SiPMs)
Tues June 09, 2015

## Slide 16
Brandon Soubasis
US Wall Halo Scintillator Counter(HSC) Installation
Tues June 09, 2015
 Counters on the upstream wall! - Measure any secondary beam halo. - Help understand radiation field of MC7.

## Slide 17
Brandon Soubasis
 Counters on the upstream wall! - Measure any secondary beam halo. - Help understand radiation field of MC7.
Tues June 30, 2015

## Slide 18
Brandon Soubasis
Tues June 30, 2015
Wall halo counters are connected into the quad scalar and sent into acnet.
Plugged into a spare v1740 board.

## Slide 19
Brandon Soubasis
Tues June 30, 2015
We can see now the Halo Scintillator Counter’s reaction to the beam spill.

## Slide 20
Brandon Soubasis
Tues June 30, 2015
I have made changes to the event display to output all 6 wall HSCs.
Single tracks seem to quiet in the counters

## Slide 21
Brandon Soubasis
Tues June 30, 2015
Notice the wall counter are lighting up when we have a lot of through the LArTPC.
Can we veto on the wall HSCs in the v1495.

## Slide 22
Brandon Soubasis
Tues June 30, 2015
We have noticed that HSC 4, HSC 5, and HSC 6 are seeing higher counts then HSC 1,
HSC 2, and HSC 3.
~6E4 counts
~2E5 counts

## Slide 23
Brandon Soubasis
 As of June 29th the movable rack has also been installed.
Tues June 30, 2015

## Slide 24
Brandon Soubasis
I have made changes to the event display again adding the movable rack Halo Scintillator Counter’s channels.
Tues June 30, 2015

## Slide 25
Brandon Soubasis
Modified Collimator Running + Simulations
Tues June 09, 2015
Prepared steel plates and braces to narrow our downstream
collimator.
Been simulating the change in our observed momentum spectrum that these plates would give.
Observing this change in our momentum spectrum would test both our beamline simulation and momentum reconstruction.

## Slide 26
Brandon Soubasis
Wed Apr 15, 2015
We have found some steel plates that have been tapped so that they can be placed inside the collimator to have the option of narrowing the beam. This will give us four options for the beam width.
Collimator
I am working on the beam line simulation of these plates within the collimator

## Slide 27
Brandon Soubasis
Wed Apr 15, 2015

## Slide 28
Brandon Soubasis
Studies with SensL Silicon photomultiplier (SiPM) – 35 - 50μm microcells
Compact!
Inexpensive!
Insensitive to magnetic field !
Low operating voltage
Tues June 09, 2015
30-33v

## Slide 29
ACKOWLEDGEMENTS
I would like to thank Jennifer Raaff, Will Flanagan, Dung Phan, Flavio Cavanna,Karol Lang, and the LArIAT collaboration for there time and guidance this summer, and the SIST committee for providing me with the opportunity to continue my research at Fermilab.